## Positive
adapted from supplementary Data:
Keller, B.O.; Sui, J.; Young, A.B.; Whittal, R.M. Interferences and contaminants encountered in modern mass spectrometry.
Analytica Chimica Acta (Review/tutorial, Special Issue on Mass Spectrometry), 2008.
List of potential interference- or contaminant ions in modern mass spectrometry, positive mode
Monoisotopic ion mass (singly charged) | Ion type | Formula for M or subunit or sequence | Compound ID or species | Possible origin and other comments |  | ESI | MALDI | References | Waters | NewObjective | Leiden Univ. | MaConDa | Agilent
[9 empty] | http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf | http://www.newobjective.com/downloads/technotes/PV-3.pdf | http://www.lc-ms.nl/ | http://www.maconda.bham.ac.uk/index.php | Common LC/MS Contaminants - CIGS
33.0335 | [M+H]+ | CH3OH | Methanol | Acetonitrile, solvent |  | X |  | A | ü |  |  | ü
42.0338 | [M+H]+ | CH3CN | ACN | Acetonitrile, solvent |  | X |  | A | ü |  | ü | ü
59.0604 | [M+NH4]+ | CH3CN | ACN | Acetonitrile, solvent | Acetonitril + Ammonia | X |  | A, F | ü |  | ü | ü
63.0441 | [A1B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
64.0158 | [M+Na]+ | CH3CN | ACN | Acetonitrile, solvent | Acetonitril + Sodium | X |  | A, F | ü |  | ü | ü
65.0597 | [M2+H]+ | CH3OH | Methanol | Methanol, solvent |  | X |  | A | ü |  | ü | ü
74.06 | [M+H]+ | C3H7NO | Dimethyl formamide | solvent |  | X |  | A, C, F | ü
74.06 | [A1B1+H]+ | (CH3CN)n(CH3OH)m | Acetonitrile/Methanol | ESI solvents |  | X |  | K |  |  |  | ü
77.0597 | [A1B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
79.0212 | [M+H]+ | C2H6OS | DMSO | Dimethylsulfoxide, solvent |  | X |  | A | ü |  | ü | ü
83.0604 | [M2+H]+ | CH3CN | Acetonitrile | ESI solvents |  | X |  | A, C | ü |  | ü | ü
85.026 | [A1B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
85.0589 | [M+H]+ | C2D6OS | d6-DMSO | d6-Dimethylsulfoxide, solvent |  | X |  | A | ü |  |  | ü
88.0393 | [A1B1+H]+ | (CH3CN)n(HCOOH)m | Acetonitrile/Formic Acid | ESI solvents |  | X |  | A, C | ü |  |  | ü
96.042 | [A1B1+Na]+ | (CH3CN)n(CH3OH)m | Acetonitrile/Methanol | ESI solvents |  | X |  | C |  |  |  | ü
99.0417 | [A1B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
100.0757 | [M+H]+ | C5H10NO | NMP | N-methyl 2-pyrrolidone; solvent, floor stripper |  | X |  | K |  |  |  | ü
100.9999 | [A1B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
101.0032 | [M+Na]+ | C2H6OS | DMSO | Dimethylsulfoxide, solvent | DMSO + Sodium | X |  | A, C | ü | ü | ü | ü
101.0808 | [A2B2+H]+ | [MeOH]n[H2O]m | Methanol/Water | ESI solvents |  | X |  | K |  |  |  | ü
102.055 | [A1B1+H]+ | (CH3CN)n(CH3COOH)m | Acetonitrile/Acetic Acid | ESI solvents |  | X |  | K |  |  |  | ü
102.1277 | [M+H]+ | C6H15N | TEA | Triethylamine, buffer | Common LC buffer, very persistent. Tubing, column, etc. Passivate LC with 5-10% acetic acid in mobile phase. Note: If TEA is used regularly, it is better to have PEEK tubing & columns dedicated for TEA use. | X |  | A, C | ü | ü | ü | ü | ü
103.9556 | [M+63Cu]+ | C2H3N | ACN | Acetonitrile, solvent | followed by 106 | X |  | A,C | ü | ü |  | ü
104.9923 | [M+Na]+ | C2H3O2Na | Sodium acetate | ESI solvents |  | X |  | K |  |  |  | ü
105.0423 | [M2+Na]+ | C2H3N | ACN | Acetonitrile, solvent | Acetonitrile + Sodium | X |  | A,C | ü | ü | ü | ü
105.9538 | [M+65Cu]+ | C2H3N | ACN | Acetonitrile, solvent |  | X |  | A, C | ü |  |  | ü
107.0703 | [A2B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
115.0156 | [A1B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
115.0866 | [A1B1+H]+ | (CH3CN)n(C3H7NO)m | Acetonitrile/Dimethylformamide | solvent |  | X |  | A | ü |  |  | ü
120.0478 | [M+CH3CN+H]+ | C2H6OS | DMSO | Dimethylsulfoxide, solvent | DMSO + Acetonitril | X |  | A | ü | ü | ü | ü
122.0812 | [M+H]+ | C4H11NO3 | TRIS | TRIS, buffer | Tris (HOCH2)3CNH2 | X |  | A | ü | ü | ü | ü
123.0628 | [A2B2+Na]+ | [MeOH]n[H2O]m | Methanol/Water | ESI solvents |  | X |  | K |  |  |  | ü
123.0917 | [M+H]+ | C7H10N2 | DMAP | Dimethylaminopyridine, solvent |  | X |  | A, C | ü | ü | ü | ü
124.0369 | [A1B1+Na]+ | (CH3CN)n(CH3COOH)m | Acetonitrile/Acetic Acid | ESI solvents |  | X |  | K |  |  |  | ü
129.0522 | [A2B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
130.159 | [M+H]+ | C8H19N | DIPEA | Diisopropylethylamine, solvent |  | X |  | A | ü | ü | ü | ü
132.9049 | M+ | Cs | Cs-133 | Cesium, from Cesium Iodide used as calibrant |  | X |  | V |  |  |  | ü
133.1071 | [A3B2+H]+ | [MeOH]n[H2O]m | Methanol/Water | ESI solvents |  | X |  | K |  |  |  | ü
135.1016 | [A2B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
137.0743 | [M+CH3CN+NH4]+ | C2H6OS | DMSO | Dimethylsulfoxide, solvent |  | X |  | A, F | ü | ü |  | ü
142.0297 | [M+CH3CN+Na]+ | C2H6OS | DMSO | Dimethylsulfoxide, solvent | DMSO + Acetonitril | X |  | A |  |  | ü | ü
144.1747 | [M+H]+ | C9H21N | TPA | Tripropylamine, solvent |  | X |  | A, C | ü | ü | ü | ü
144.9821 | [M2+63Cu]+ | CH3CN | ACN | Acetonitrile, solvent, together with m/z 147 |  | X |  | A, C | ü | ü |  | ü
145.0262 | [A2B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
146.0689 | [M3+Na]+ | CH3CN | ACN | Acetonitrile, solvent | Acetonitrile + Sodium | X |  | A, C | ü | ü | ü | ü
146.9803 | [M2+65Cu]+ | CH3CN | ACN | Acetonitrile, solvent, together with m/z 145 |  | X |  | A, C | ü | ü |  | ü
147.1128 | [A2B2+H]+ | (CH3CN)n(CH3OH)m | Acetonitrile/Methanol | ESI solvents |  | X |  | K |  |  |  | ü
149.0233 | [f+H]+ | C8H4O3 | Pthalic Anhydride | fragment ion originating from phthalate esters |  | X |  | A, Z | ü |  |  | ü | ü
150.1277 | [M+H]+ | C10H15N | Phenyldiethylamine | solvent |  | X |  | A, C | ü | ü | ü | ü
151.0965 | [A3B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
153.1386 | [M+H]+ | C9H16N2 | DBU | 1,8-Diazabicyclo[5.4.0]undec-7-ene |  | X |  | A, C | ü | ü | ü | ü
155.089 | [A3B2+Na]+ | [MeOH]n[H2O]m | Methanol/Water | ESI solvents |  | X |  | K |  |  |  | ü
157.0351 | [M2+H]+ | C2H6OS | DMSO | Dimethylsulfoxide, solvent |  | X |  | A, C | ü | ü | ü | ü
157.0835 | [A2B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
158.964 | [M+Na]+ | C2F3O2Na | NaTFA | Sodium trifluoroacetate, salt |  | X |  | A, C | ü | ü | ü | ü
163.039 | [M-CH3OH+H]+ | C10H10O4 | Dimethyl phthalate | Phthalate esters, plasticizer |  | X |  | A | ü |  |  | ü
163.1329 | [M+H]+ | C8H18O3 | DGBE | Diethylene glycol monobutyl ether, cpd. In scintillation cocktail |  | X |  | Q |  |  |  | ü
167 |  |  |  | dioctyl phthalate |  | X |  |  | ü
167.0339 | [M+H]+ | C8H6O4 | Phthalic acid |  |  | X |  | Guo, X. H., Bruins, A. P., & Covey, T. R. (2006). Characterization of typical chemical background interferences in atmospheric pressure ionization liquid chromatography-mass spectrometry. Rapid Communications in Mass Spectrometry, 20(20), 3145-3150, doi:10.1002/rcm.2715. |  |  |  | ü
169.0947 | [A2B2+Na]+ | (CH3CN)n(CH3OH)m | Acetonitrile/Methanol | ESI solvents |  | X |  | K |  |  |  | ü
169.1105 | [M2+H]+ | C2D6OS | d6-DMSO | d6-Dimethylsulfoxide, solvent |  | X |  | A | ü |  |  | ü
171.0053 | [f+Na]+ | C8H4O3 | Phthalic anhydride | from phthalate esters, plasticizer |  | X |  | A | ü |  |  | ü
172.0393 | [M-H2O+H]+ | C10H7NO3 | 4-HCCA | matrix compound |  | X | X | H, I, J |  |  |  | ü
173.0575 | [A2B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
173.0784 | [A3B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
179.0171 | [M2+Na]+ | C2H6OS | DMSO | Dimethylsulfoxide, solvent | DMSO + Sodium | X |  | A, C | ü | ü | ü | ü
181.1223 | [M+H]+ | C11H16O2 | BHA | Butylated hydroxyanisole, antioxidant additives |  | X |  | K |  |  |  | ü
183.0804 | [M+H]+ | C13H10O | DPK | Diphenyl ketone |  | X |  | K |  |  |  | ü
183.1438 | [A4B3+H]+ | [MeOH]n[H2O]m | Methanol/Water | ESI solvents |  | X |  | K |  |  |  | ü
185.1148 | [M+Na]+ | C8H18O3 | GE | glycol ether |  | X |  | K |  |  |  | ü
186.2216 | [M+H]+ | C12H27N | TBA | Tributylamine, solvent |  | X |  | A | ü | ü | ü | ü
189.0524 | [A3B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
190.0499 | [M+H]+ | C10H7NO3 | 4-HCCA | matrix compound |  | X | X | H, I, J |  |  |  | ü
193.1434 | [A3B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
195.0652 | [M+H]+ | C6H4[COOCH3]2 | DMP | Dimethyl phthalate, Phthalate esters, plasticizer |  | X |  | A | ü |  | ü | ü | ü
195.1227 | [A4B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
203.1043 | [M+Na]+ | C11H16O2 | BHA | Butylated hydroxyanisole, antioxidant additives |  | X |  | K |  |  |  | ü
204.0781 | [M+C4H8]+ | C8H4O3 | Phthalic anhydride | Plasticiser |  |  |  | Guo, X. H., Bruins, A. P., & Covey, T. R. (2006). Characterization of typical chemical background interferences in atmospheric pressure ionization liquid chromatography-mass spectrometry. Rapid Communications in Mass Spectrometry, 20(20), 3145-3150, doi:10.1002/rcm.2715. |  |  |  | ü
205.1258 | [A4B3+Na]+ | [MeOH]n[H2O]m | Methanol/Water | ESI solvents |  | X |  | K |  |  |  | ü
209 |  |  | Pthalic Anhydride | Phthalates Plastics.(Nalgene bottles, tubing) Rinse plasticware with isopropal alcohol. Minimize use of low-density polyethylene plasticware. Use Nalgene (high-density polyethylene or polypropylene) or teflon plasticware. | [8 empty] | ü
212.0318 | [M+Na]+ | C10H7NO3 | 4-HCCA | matrix compound, sodiated |  | X | X | H, I, J |  |  |  | ü
214 | (M+H)+ |  |  | unidentified contaminant, maybe related to methanol |  | X |  |  |  |  | ü
214.0896 | [M+H]+ | C10H15NO2S | n-BBS | n-butyl benzenesulfonamide, plasticizer |  | X |  | A, K | ü | ü |  | ü
215.1254 | [A3B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
217.1046 | [A4B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  | ü |  | ü
221.19 | [M+H]+ | C15H24O | BTH | Butylated hydroxytoluene, Antioxidant |  | X |  | K |  |  |  | ü
223.0965 | 222.0887 | C6H4[COOC2H5]2 | DEP | Diethyl phthalate | [6 empty] | ü
225.1961 | [M+H]+ | C13H24N2O | DCU | N,N'-Dicyclohexylurea |  | X |  | A | ü | ü | ü | ü
228.0058 | [M+K]+ | C10H7NO3 | 4-HCCA | matrix compound, potassiated |  | X | X | H, I, J |  |  |  | ü
231.0993 | [A3B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
231.1162 | [M+NH4]+ | C10H15NO2S | n-BBS | n-butyl benzenesulfonamide, plasticizer | † Derived from rubber tip from disposable syringe plunger | X |  | A | ü | ü |  | ü
233.0786 | [A4B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
235.0491 | (3M+H)+ |  | DMSO |  |  | X |  |  |  |  | ü
236.0716 | [M+Na]+ | C10H15NO2S | n-BBS | n-butyl benzenesulfonamide, plasticizer |  | X |  | A | ü |  |  | ü
237 |  |  | Pthalic Anhydride | Plasticizers – Tygon tubing, solvents like methanol, and nitrogen generators. | [8 empty] | ü
239.1489 | [A5B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
239.2249 | [(M.H35Cl)2-Cl]+ | C6H15N | TEA.HCl | Triethylamine-hydrochloride, buffer |  | X |  | A | ü | ü | ü | ü
241.2219 | [(M.H37Cl)2-Cl]+ | C6H15N | TEA.HCl | Triethylamine-hydrochloride, buffer |  | X |  | A | ü | ü | ü | ü
242.2842 | M+ | C16H36N | TBA | Tetrabutylammonium, buffer | Tetrabutylammonium (C4H9)4N+ | X |  | A, C | ü | ü | ü | ü
243.1168 | M+ | C19H15 | Trityl cation | Trityl cation, [Ph3C]+ |  | X |  | A | ü | ü | ü | ü
243.1719 | [M+Na]+ | C15H24O | BTH | Butylated hytroxytoluene, Antioxidant additives |  | X |  | K |  |  |  | ü
247.0965 | 246.0887 | C6H4(COOCH2CH=CH2)2 | DAP | Diallyl phthalate | [6 empty] | ü
251.1278 | 250.12 | C6H4(COO(CH2)2CH3)2 | DPP | Di-n-propyl phthalate | [6 empty] | ü
251.1853 | [A4B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
251.2006 | [AB1+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X |  | D, K |  |  |  | ü
255 |  |  | Pthalic Anhydride | [9 empty] | ü
257.031 | [M3+Na]+ | C2H6OS | DMSO | Dimethylsulfoxide, solvent |  | X |  | A | ü | ü | ü | ü
261.1309 | [A5B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
265.2162 | [AB1+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X |  | D, K |  |  |  | ü
267.172 | [M+H]+ | C12H27O4P | TBP | Tributylphosphate |  | X |  | A | ü | ü |  | ü
273.1274 | M+ | C20H17O | MMT | Monomethoxytrityl cation |  | X |  | A | ü | ü | ü | ü
273.1672 | [A4B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
273.1825 | [AB1+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X |  | D, K |  |  |  | ü
277.1048 | [A5B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
279.0933 | [M+H]+ | C18H15OP | TPO | Triphenylphosphine oxide |  | X |  | A | ü |  |  | ü
279.1591 | [M+H]+ | C6H4(COO(CH2)3CH3)2 | DBP | Dibutylphthalate, Plasticiser, phtalate ester |  | X |  | A, C | ü | ü | ü | ü
279.1591 | 278.1513 | C6H4(COOCH2CH(CH3)2)2 | DIBP | Diisobutyl phthalate | [6 empty] | ü
279.2295 | [AB1+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X |  | D, K |  |  |  | ü
282.2791 | [M+H]+ | C18H35NO | Oleamide | Slip agent in polyethylene films |  | X |  | A | ü | ü | ü | ü
283.1751 | [A6B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
284.2948 | [M+H]+ | C18H37NO | Stearamide | Slip agent in polyethylene films |  | X |  | A | ü |  |  | ü
287.1982 | [AB1+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X |  | D, K |  |  |  | ü
288.2533 | [M+H]+ | C16H33NO3 | n,n-DDA | n,n-bis(2-hydroxyethyl) dodecanamide, anti-static agent |  | X |  | A | ü |  |  | ü
289.1412 | [A4B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
293.2451 | [AB1+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X |  | D, K |  |  |  | ü
295.2268 | [AB2+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X |  | D, K |  |  |  | ü
301.141 | [M+Na]+ | C16H22O4 | Dibutylphthalate | Dibutylphthalate, plasticizer | Dibutylphthalate (plasticiser) + Sodium | X |  | A | ü | ü | ü | ü
304.2611 | [M+Na]+ | C18H35NO | Oleamide | Slip agent in polyethylene films |  | X |  | A | ü |  |  | ü
305.1571 | [A6B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
305.1747 | 304.1669 | CH3(CH2)3OOCC6H4COOC6H11 | BCP | Butyl cyclohexyl phthalate | [6 empty] | ü
306.2767 | [M+Na]+ | C18H37NO | Stearamide | Slip agent in polyethylene films |  | X |  | A |  |  |  | ü
307.1904 | 306.1826 | C6H4(COO(CH2)4CH3)2 | DNPP | Di-n-pentyl phthalate | [6 empty] | ü
309.2272 | [A5B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
309.2424 | [AB2+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X |  | D, K |  |  |  | ü
313.1434 | 312.1356 | CH3(CH2)3OOCC6H4COOCH2C6H5 | BBP | Butyl benzyl phthalate | [6 empty] | ü
315.253 | [M+H]+ | C18H34O4 | DBS | Dibutyl sebacate, plasticizer |  | X |  | X |  |  |  | ü
317.115 | [M+K]+ | C16H22O4 | Dibutylphthalate | Dibutylphthalate, plasticizer | Dibutylphthalate (plasticiser) + Potassium | X |  | A | ü | ü | ü | ü
317.2087 | [AB2+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X |  | D, K |  |  |  | ü
321.131 | [A6B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
323.2557 | [AB2+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X |  | D, K |  |  |  | ü
325.2585 | [M2+H]+ | C8H18O3 | DGBE | Diethylene glycol monobutyl ether, cpd. In scintillation cocktail |  | X |  | Q |  |  |  | ü
327.0781 | [M+H]+ | C18H15O4P | TPP | Triphenyl phosphate, flame retardant in plastics |  | X |  | X |  |  |  | ü
327.2013 | [A7B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
331.1904 | 330.1826 | C6H4(COOC6H11)2 | DCP | Dicyclohexyl phthalate | [6 empty] | ü
331.2091 | [A5B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
331.2244 | [AB2+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X |  | D, K |  |  |  | ü
335.2217 | 334.2139 | C6H4(COO(CH2)5CH3)2 | DNHP | Di-n-hexyl phthalate | [6 empty] | ü
335.2217 | 334.2139 | C6H4(COO(CH2)3CH(CH3)2)2 | DIHxP | Diisohexyl phthalate | [6 empty] | ü
337.1184 | [M+H]+; (120Sn) | C13H28O2Sn | Tributyl tin formate | Tributyl tin formate, catalyst |  | X |  | A | ü | ü |  | ü
337.2713 | [AB2+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X |  | D, K |  |  |  | ü
338.3417 | [M+H]+ | C22H43NO | Erucamide | Erucamide, (Cis-13-docosenoic amide) |  | X |  | A | ü | ü | ü | ü
339.253 | [AB3+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X |  | D, K |  |  |  | ü
347.183 | [A5B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
349.1833 | [A7B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
353.2686 | [AB3+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X |  | D, K |  |  |  | ü
355.0699 | [M+H-CH4]+ | [C2H6SiO]5 | Polysiloxane | Polysiloxane, (neutral methane loss from m/z 371) |  | X |  | A, C, L | ü |  |  | ü
355.3683 | [M-Cl]+ | C22H47N2OCl |  | Palmitamidopropyl-trimonium chloride, personal care products additive |  | X |  | O |  |  |  | ü
360.3237 | [M+Na]+ | C22H43NO | Erucamide | Erucamide, (Cis-13-docosenoic amide) | Erucamide + Sodium | X |  | A | ü | ü | ü | ü
361.2349 | [AB3+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X |  | D, K |  |  |  | ü
363.253 | 362.2452 | C6H4(COO(CH2)4CH(CH3)2)2 | DIHpP | Diisoheptyl phthalate | [6 empty] | ü
363.253 | 362.2452 | CH3(CH2)3OOCC6H4COO(CH2)9CH3 | BDP | Butyl decyl phthalate | [6 empty] | ü
365.1572 | [A7B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
367.269 | [A6B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
367.2819 | [AB3+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X |  | D, K |  |  |  | ü
368.4251 | [M-Cl]+ | C25H54NCl | BTAC-228 | Behentrimonium chloride, personal care product additive |  | X |  | O |  |  |  | ü
371.1012 | [M+H]+ | [C2H6SiO]5 | Polysiloxane | Polysiloxane, followed by m/z 388 |  | X |  | A, C., E, L | ü | ü |  | ü
371.2276 | [A8B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
371.3156 | [M+H]+ | C22H42O4 | DEHA | Bis(2-ethylhexyl) adipate, plasticizer |  | X |  | N, X
371.3156 | [M+H]+ | C22H42O4 | DOA | Dioctyl adipate, plasticizer |  | X |  | X |  |  |  | ü
375.2506 | [AB3+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X |  | D, K |  |  |  | ü
379.0925 | [M2+H]+ | C10H7NO3 | 4-HCCA | matrix cluster, dimer |  | X | X | H, I, J |  |  |  | ü
381.2975 | [AB3+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X |  | D, K |  |  |  | ü
383.2792 | [AB4+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X |  | D, K |  |  |  | ü
388.1278 | [M+NH4]+ | [C2H6SiO]5 | Polysiloxane | Polysiloxane, (see m/z 371) |  | X |  | A, C, L | ü |  |  | ü
389.251 | [A6B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
391.2843 | [M+H]+ | C6H4(COO(CH2)5CH(CH3)2)2 | DIOP | Diisooctyl phthalate, plasticiser | usually from contaminated solvents, can concentrate on the column and elute during a gradient | X |  | A, C, E | ü | ü | ü | ü
391.2843 | 390.2765 | C6H4(COOCH2CH(C2H5)(CH2)3CH3)2 | DEHP, DOP | Di(2-ethylhexyl) phthalate | ubiquitous plasticiser bis(2-ethylhexyl) phthalate | [5 empty] | ü |  | ü
391.2843 | 390.2765 | C6H4(COO(CH2)7CH3)2 | DNOP | Di(n-octyl) phthalate | [6 empty] | ü
393.2095 | [A8B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X |  | D, K |  |  |  | ü
397.2949 | [AB4+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X |  | D, K |  |  |  | ü
405.2249 | [A6B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
405.2611 | [AB4+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
409.1834 | [A8B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
411.3081 | [AB4+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
413.2662 | [M+Na]+ | C24H38O4 | Diisooctyl phthalate | Diisooctyl phthalate, plasticiser | usually from contaminated solvents, can concentrate on the column and elute during a gradient | X | X | A, C, E | ü | ü | ü | ü
415.2538 | [A9B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
419.2768 | [AB4+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
419.3156 | 418.3078 | CH3(CH2)7OOCC6H4COO(CH2)9CH3 | ODP | n-Octyl n-decyl phthalate | [6 empty] | ü
419.3156 | 418.3078 | C6H4(COO(CH2)6CH(CH3)2)2 | DINP | Diisononyl phthalate | [6 empty] | ü
425 | (M+Na)+ |  |  | unidentified contaminant (plasticiser) | Unidentified contaminant (plasticiser) + Sodium | X |  |  | ü | ü | ü
425.3109 | [A7B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
425.3237 | [AB4+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
427.3054 | [AB5+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
429.0887 | [M+H-CH4]+ | [C2H6SiO]6 | Polysiloxane | Polysiloxane, (neutral methane loss from m/z 445) |  | X | X | L |  |  |  | ü
429.2402 | [M+K]+ | C24H38O4 | Diisooctyl phthalate | Diisooctyl phthalate, plasticiser |  | X | X | A, C | ü | ü | ü | ü
437.2357 | [A9B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
441.0148 | [M3+63Cu(I)]+ | C10H7NO3 | 4-HCCA | matrix-copper adduct, together with m/z 443 |  | X | X | K |  |  |  | ü
441.3211 | [AB5+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
443.013 | [M3+65Cu(I)]+ | C10H7NO3 | 4-HCCA | matrix-copper adduct, together with m/z 443 |  | X | X | K |  |  |  | ü
445.12 | [M+H]+ | [C2H6SiO]6 | Polysiloxane | Polysiloxane, followed by m/z 462 | (M+H)+ dodecamethylcyclohexasiloxane seen w/ silica capillary NanoSprayer | X | X | A, C, E, L | ü | ü |  | ü
447.2928 | [M+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
447.3469 | 446.3391 | C6H4(COO(CH2)7CH(CH3)2)2 | DIDP | Diisodecyl phthalate | [6 empty] | ü
449.2874 | [AB5+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
449.385 | [M2+H]+ | C13H24N2O | DCU | N,N'-Dicyclohexylurea |  | X | X | A, C | ü | ü | ü | ü
453.2097 | [A9B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
453.3435 | [M+H]+ | C24H44N4O4 | nylon | Cyclic oligomer of polyamide 66, (adipic acid/hexylmethylene diamine condensation) |  | X | X | M |  |  |  | ü
454.2928 | [M+CH3CN+Na]+ | C24H38O4 | Diisooctyl phthalate | Diisooctyl phthalate, plasticiser | also observed w/ high conc. of formic acid | X | X | A | ü | ü | ü | ü
455.3343 | [AB5+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
459.28 | [A10B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
462.1466 | [M+NH4]+ | [C2H6SiO]6 | Polysiloxane | Polysiloxane (see m/z 445) |  | X | X | A, L | ü |  |  | ü
463.2668 | [A7B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
463.303 | [AB5+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
469.35 | [AB5+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
471.3316 | [AB6+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
472.2878 | [M+H]+ | SLPR | Peptide | porcine trypsin |  | X | X | G
475.3782 | 474.3704 | C6H4(COO(CH2)10CH3)2 | DUP | Diundecyl phthalate | [6 empty] | ü
475.3782 | 474.3704 | C6H4(COO(CH2)8CH(CH3)2)2 | DIUP | Diisoundecyl phthalate | [6 empty] | ü
481.2619 | [A10B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
483.3528 | [A8B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
485.3473 | [AB6+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
493.3136 | [AB6+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
494.5659 | [M-Cl]+ | C34H72NCl | DPDMA | Dipalmityldimethylammonium chloride, catalyst, personal care products additive |  | X | X | O, K |  |  |  | ü
497.2359 | [A10B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
499.3605 | [AB6+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
503.1075 | [M+H-CH4]+ | [C2H6SiO]7 | Polysiloxane | Polysiloxane, (neutral methane loss from m/z 519) | seen w/ silica capillary NanoSprayer | X | X | A, L | ü |  |  | ü
503.3062 | [A11B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
505.3347 | [A8B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
507.3292 | [AB6+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
513.3762 | [AB6+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
515.33 | [M+H]+ | IQVR | Peptide | trypsin-like artefact |  | X | X | E
515.3578 | [AB7+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
515.4129 | [M+H]+ | C30H58O4S | DDTDP | Didodecyl 3,3′-thiodipropionate, antioxidant |  | X | X | W |  |  |  | ü
519.1388 | [M+H]+ | [C2H6SiO]7 | Polysiloxane | Polysiloxane, followed by m/z 536 | Polydimethylcyclosiloxanes, widespread compound in industrial products, e.g. deodorants. They are ubiquitously present in ambient air. seen w/ silica capillary NanoSprayer | X | X | A, L | ü |  |  | ü
521.3086 | [A8B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
522 |  |  |  | Unknown† | Derived from rubber tip from disposable syringe plunger | X |  |  | ü | ü
522.5972 | [M-Cl]+ | C36H76NCl | SPDMA | Stearyl-palmityldimethylammonium chloride, catalyst, personal care product additive |  | X | X | O, K |  |  |  | ü
525.2881 | [A11B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
529.3735 | [AB7+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
531 | (M+H)+ |  |  | seen w/ Nunc polypropylene tubes |  | X |  |  | ü
531.4078 | [M+H]+ | C30H58O5S | DDTDP | Didodecyl 3,3′-thiodipropionate oxidized to sulfoxide, antioxidant |  | X | X | W |  |  |  | ü
531.4408 | 530.433 | C6H4(COO(CH2)12CH3)2 | DTDP | Ditridecyl phthalate | [6 empty] | ü
531.4408 | 530.433 | C6H4(COO(CH2)10CH(CH3)2)2 | DIUP | Diisotridecyl phthalate | [6 empty] | ü
531.4772 | [M+H]+ | C35H62O3 | Irganox | Irganox 1076, antioxidant in synthetic polymers, antioxidant |  | X | X | W |  |  |  | ü
536.1654 | [M+NH4]+ | [C2H6SiO]7 | Polysiloxane | Polysiloxane (see m/z 519) | If the silicon tube is instead connected to the API source housing background ions at m/z 536, 610 and 684 might be observed. Remove all silicon tubing from ESI housing and/or replace with tygon tubing, seen w/ silica capillary NanoSprayer | X | X | A, L | ü |  |  | ü
537.3398 | [AB7+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
537.879 | [M6-6H+3Fe+O]+ | C2H4O2 | Acetic acid-Fe-O- complex | during ESI with metal tips and acetic acid | (acetate)n+(Fe)n+ observed w/ high conc. of acetic acid | X |  | A, B, E, Y | ü |  |  | ü
541.2621 | [A11B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
541.3946 | [A9B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
543.3867 | [AB7+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
547.3324 | [A12B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
547.4027 | [M+H]+ | C30H58O6S | DDTDP | Didodecyl 3,3′-thiodipropionate oxidized to sulfone, antioxidant |  | X | X | W |  |  |  | ü
550 |  |  |  | Unknown | Derived from rubber tip from disposable syringe plunger | X |  |  |  | ü
550.6285 | [M-Cl]+ | C38H80NCl | DSDMA | Distearyldimethylammonium chloride, catalyst, personal care products additive | Derived from rubber tip from disposable syringe plunger | X | X | A, K, O | ü |  |  | ü
551.3554 | [AB7+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
553 | (M+Na)+ |  |  | seen w/ Nunc polypropylene tubes |  | X |  |  | ü
553.3897 | [M+Na]+ | C30H58O5S | DDTDP | Didodecyl 3,3′-thiodipropionate oxidized to sulfoxide, antioxidant |  | X | X | A, W |  |  |  | ü
553.4591 | [M+Na]+ | C35H62O3 | Irganox | Irganox 1076, antioxidant in synthetic polymers, antioxidant |  | X | X | A, W |  |  |  | ü
555.8896 | [M6-6H+H2O+3Fe+O]+ | C2H4O2 | Acetic acid-Fe-O- complex | during ESI with metal tips and acetic acid | observed w/ high conc. of acetic acid | X |  | A, B, E, Y | ü |  |  | ü
557.4024 | [AB7+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
559.3841 | [AB8+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
563.3766 | [A9B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
568.1351 | [M3+H]+ | C10H7NO3 | 4-HCCA | matrix cluster, trimer |  | X | X | H, I, J |  |  |  | ü
569 | (M+K)+ |  |  | seen w/ Nunc polypropylene tubes |  | X |  |  | ü
569.3144 | [A12B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
571.3562 | [M+H]+ | VSLPR | Peptide | porcine trypsin |  | X | X | G
573.3997 | [AB8+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
577.1263 | [M+H-CH4]+ | [C2H6SiO]8 | Polysiloxane | Polysiloxane, (neutral methane loss from m/z 593) |  | X | X | L |  |  |  | ü
579.3505 | [A9B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
581.366 | [AB8+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
585.2883 | [A12B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
585.3586 | [M+K]+ | C30H58O6S | Didodecyl 3,3'-thiodipropanoate sulfone product | Antioxidant |  | X |  | Xia, Y. Q., Patel, S., Bakhtiar, R., Franklin, R. B., & Doss, G. A. (2005). Identification of a new source of interference leached from polypropylene tubes in mass-selective analysis. Journal of the American Society for Mass Spectrometry, 16(3), 417-421, doi:10.1016/j.jasms.2004.11.020. |  |  |  | ü
587 | (M+H)+ |  |  | seen w/ Vanguard polypropylene tubes |  | X |  |  | ü
587.413 | [AB8+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
591.3586 | [A13B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
593.1576 | [M+H]+ | [C2H6SiO]8 | Polysiloxane | Polysiloxane, followed by m/z 610 | Polydimethylcyclosiloxanes, widespread compound in industrial products, e.g. deodorants. They are ubiquitously present in ambient air. | X | X | A, L | ü |  |  | ü
595.3817 | [AB8+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
597.9001 | [M7-6H+3Fe+O]+ | C2H4O2 | Acetic acid-Fe-O- complex | during ESI with metal tips and acetic acid |  | X |  | B, E, Y |  |  |  | ü
599 |  |  |  | Unknown |  | X |  |  |  | ü
599.4365 | [A10B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
601.4286 | [AB8+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
603.4103 | [AB9+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
605.2923 | [M+K+C2HF3O2]+ | C24H44N4O4 | Nylon 66 polymer | Filter |  | X |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. |  |  |  | ü
606.0915 | [M3+K]+ | C10H7NO3 | 4-HCCA | matrix cluster, trimer, potassiated |  | X | X | H, I, J |  |  |  | ü
609 | (M+Na)+ |  |  | seen w/ Vanguard polypropylene tubes |  | X |  |  | ü
610 |  |  |  | seen w/ silica capillary NanoSprayer |  | X |  |  | ü
610.1842 | [M+NH4]+ | [C2H6SiO]8 | Polysiloxane | Polysiloxane (see m/z 593) | If the silicon tube is instead connected to the API source housing background ions at m/z 536, 610 and 684 might be observed. Remove all silicon tubing from ESI housing and/or replace with tygon tubing | X | X | L |  |  |  | ü
613.3406 | [A13B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
615.4038 | [M+H]+ | C32H58N2O7S | CHAPS | 3-[(3-Cholamidopropyl)dimethylammonio]-1-propanesulfonate |  | X | X | E |  |  |  | ü
617.4259 | [AB9+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
621.4184 | [A10B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
621.9729 | [M6-6H+3Fe+O]+ | C3H6O2 | Propionic acid Fe-O complex | during ESI with metal tips and acetic acid |  | X |  | Y |  |  |  | ü
625 | (M+K)+ |  |  | seen w/ Vanguard polypropylene tubes |  | X |  |  | ü
625.3922 | [AB9+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
629.3145 | [A13B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
631.4392 | [AB9+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
633.3202 | [M+H]+ | QTIASN | Peptide | Bovine trypsin |  | X | X | K
635.3849 | [A14B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
637.3924 | [A10B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
638 |  |  |  | seen w/ silica capillary NanoSprayer |  | X |  |  | ü
639.4079 | [AB9+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
644.0468 | [M3-H+K2]+ | C10H7NO3 | 4-HCCA | Matrix compound |  |  | X | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. |  |  |  | ü
645.4548 | [AB9+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
646.3519 | [f+H]+ | GLTAER (from YLDGLTAER) | Peptide | keratin (human), ion source fragment |  | X | X | G, K
647.4365 | [AB10+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
651.1451 | [M+H-CH4]+ | [C2H6SiO]9 | Polysiloxane | Polysiloxane, (neutral methane loss from m/z 667) |  | X | X | L |  |  |  | ü
657.3668 | [A14B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
657.4784 | [A11B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
657.9213 | [M8-H6+Fe3+O]+ | C2H4O2 | Acetic Acid | Solvent |  | X |  | Ijames, C. F., Dutky, R. C., & Fales, H. M. (1995). Iron carboxylate oxygen-centered-triangle complexes detected during electrospray use of organic acid modifiers with a comment on the Finnigan TSQ-700 electrospray inlet system. Journal of the American Society for Mass Spectrometry, 6(12), 1226-1231, doi:10.1016/1044-0305(95)00579-x. |  |  |  | ü
659.3835 | [M+H]+ | SGIQVR | Peptide | Bovine trypsin |  | X | X | K
661.4521 | [AB10+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
666.0288 | [M3-H2+K2+Na]+ | C10H7NO3 | 4-HCCA | Matrix compound |  |  | X | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. |  |  |  | ü
667.1764 | [M+H]+ | [C2H6SiO]9 | Polysiloxane | Polysiloxane, followed by m/z 684 | Polydimethylcyclosiloxanes, widespread compound in industrial products, e.g. deodorants. They are ubiquitously present in ambient air. | X | X | L | ü |  |  | ü
669.4184 | [AB10+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
672.4039 | [M+H]+ | TVSLPR | Peptide | porcine trypsin |  | X | X | K
672.4039 | [M+H]+ | TVSLPR | Peptide | porcine trypsin |  | X | X | K
673.3407 | [A14B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
675 |  |  |  | seen w/ new PEEK tubing |  | X |  |  | ü
675.4654 | [AB10+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
679 | (M+H)+ |  |  | from purified water and solvent filters |  | X |  |  | ü
679.4111 | [A15B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
679.4603 | [A11B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
679.5117 | [M+H]+ | C36H66N6O6 | nylon | Cyclic oligomer of polyamide 66, (adipic acid/hexylmethylene diamine condensation) |  | X | X | M |  |  |  | ü
680.9949 | [M3-H3+K3]+ | C10H7NO3 | 4-HCCA | Matrix compound |  |  | X | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. |  |  |  | ü
683.4341 | [AB10+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
684 |  |  |  | seen w/ silica capillary NanoSprayer |  | X |  |  | ü
684.203 | [M+NH4]+ | [C2H6SiO]9 | Polysiloxane | Polysiloxane (see m/z 667) | If the silicon tube is instead connected to the API source housing background ions at m/z 536, 610 and 684 might be observed. Remove all silicon tubing from ESI housing and/or replace with tygon tubing | X | X | L |  |  |  | ü
689.481 | [AB10+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
691.4627 | [AB11+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
695.4342 | [A11B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
696 | (M+NH4)+ |  |  | from purified water and solvent filters |  | X |  |  | ü |  |  | ü
701 | (M+Na)+ |  |  | from purified water and solvent filters |  | X |  |  | ü
701.393 | [A15B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
701.4936 | [M+Na]+ | C36H66N6O6 | Nylon 66 polymer | Filter |  | X |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. |  |  |  | ü
704.3825 | [M+H]+ | LDSELK | Peptide | keratin |  | X | X | E
705.4784 | [AB11+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
713.4447 | [AB11+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
715.5202 | [A12B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
717.3669 | [A15B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
717.4675 | [M+K]+ | C36H66N6O6 | Nylon 66 polymer | Filter |  |  |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. |  |  |  | ü
719.4916 | [AB11+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
723.4373 | [A16B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
725.1639 | [M+H-CH4]+ | [C2H6SiO]10 | Polysiloxane | Polysiloxane, (neutral methane loss from m/z 741) |  | X | X | L |  |  |  | ü
727.4603 | [AB11+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
732.4654 | [M+H]+ | GLVLIAF | Peptide | bovine serum albumin (BSA) |  | X | X | E
733.5072 | [AB11+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
735.4889 | [AB12+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
737.5022 | [A12B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
741.1952 | [M+H]+ | [C2H6SiO]10 | Polysiloxane | Polysiloxane, followed by m/z 758 |  | X | X | L | ü |  |  | ü
742.4498 | [M+H]+ | GPFPILV | Peptide | bovine casein alpha-S1 |  | X | X | E
743.441 | [M+H]+ | ATVSLPR | Peptide | porcine trypsin |  | X | X | G
745.4192 | [A16B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
749.5046 | [AB12+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
753.4761 | [A12B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
757.4709 | [AB12+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
758.2217 | [M+NH4]+ | [C2H6SiO]10 | Polysiloxane | Polysiloxane (see m/z 741) |  | X | X | L |  |  |  | ü
758.4155 | [M+H]+ | PATLNSR | Peptide | porcine trypsin |  | X | X | G
761.3932 | [A16B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
763.5178 | [AB12+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
767.4635 | [A17B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
770.0465 | [M8-H6+Fe3+O]+ | C3H6O2 | Propionic acid |  |  | X |  | Ijames, C. F., Dutky, R. C., & Fales, H. M. (1995). Iron carboxylate oxygen-centered-triangle complexes detected during electrospray use of organic acid modifiers with a comment on the Finnigan TSQ-700 electrospray inlet system. Journal of the American Society for Mass Spectrometry, 6(12), 1226-1231, doi:10.1016/1044-0305(95)00579-x. |  |  |  | ü
771.4865 | [AB12+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
773.5621 | [A13B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
777.5335 | [AB12+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
779.5151 | [AB13+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
786 | (M+Ti)+ |  |  | titanium adduct |  | X |  |  | ü
789.4454 | [A17B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
793.5308 | [AB13+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
795.544 | [A13B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
798.5878 | [M2+NH4]+ | C24H38O4 | Diisooctyl phthalate | Diisooctyl phthalate, plasticiser | usually from contaminated solvents, can concentrate on the column and elute during a gradient | X | X | A, C, E | ü | ü | ü | ü
801.4971 | [AB13+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
802.4305 | [M+H]+ | LSSPATLN | Peptide | porcine trypsin, truncated |  | X | X | E
803.5432 | [M2+Na]+ | C24H38O4 | Diisooctyl phthalate | Diisooctyl phthalate, plasticiser | usually from contaminated solvents, can concentrate on the column and elute during a gradient | X | X | A, C, E | ü | ü | ü | ü
804.4098 | [M+H]+ | SEIDNVK | Peptide | keratin (human) |  | X | X | G
805.4163 | [M+H]+ | SAASLNSR | Peptide | bovine trypsin |  | X | X | G, K
805.4194 | [A17B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
807.3995 | [M+H]+ | LAADDFR | Peptide | keratin (human) |  | X | X | G, K
807.544 | [AB13+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
809.4403 | [M+H]+ | LASYLDK | Peptide | keratin |  | X | X | E
809.4869 | [AB10+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
811.4897 | [A18B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
811.518 | [A13B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
815.5127 | [AB13+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
819.5172 | [M2+K]+ | C24H38O4 | Diisooctyl phthalate | Diisooctyl phthalate, plasticiser |  | X | X | A, K |  |  |  | ü
821.5597 | [AB13+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
823.5414 | [AB14+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
824.4989 | [M+H]+ | PGVVSLPR | Peptide | trypsin-like artefact |  | X | X | E
827.4298 | [M+H]+ | FASFIDK | Peptide | keratin (human) |  | X | X | E, G, K
827.4621 | [M+H]+ | PEIQNVK | Peptide | keratin (human) |  | X | X | G, K
831.4604 | [M+K+C2HF3O2]+ | C36H66N6O6 | Nylon 66 polymer | Filter |  | X |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. |  |  |  | ü
831.604 | [A14B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
832.4887 | [M+H]+ | SISISVAR | Peptide | keratin (human) |  | X | X | G, K
833.0894 | [M4-H+K2]+ | C10H7NO3 | 4-HCCA | Matrix compound |  |  | X | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. |  |  |  | ü
833.4717 | [A18B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
837.557 | [AB14+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
839.0974 | [M4-2H+K+2Na]+ | C10H7NO3 | 4-HCCA | matrix cluster |  | X | X | H, I, J |  |  |  | ü
842.5094 | [M+H]+ | VATVSLPR | Peptide | porcine trypsin |  | X | X | E, G, K
845.1054 | [M4-3H+4Na]+ | C10H7NO3 | 4-HCCA | matrix cluster |  | X | X | H, I, J |  |  |  | ü
845.5233 | [AB14+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
848.4989 | [M+H]+ | AFIDKVR | Peptide | keratin (human) |  | X | X | G, K
849.4456 | [A18B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
851.5702 | [AB14+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
853.5131 | [AB11+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
853.5859 | [A14B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
855.0714 | [M4-2H+Na+2K]+ | C10H7NO3 | 4-HCCA | matrix cluster |  | X | X | H, I, J |  |  |  | ü
855.5159 | [A19B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
859.5389 | [AB14+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
861.0794 | [M4-3H+3Na+K]+ | C10H7NO3 | 4-HCCA | matrix cluster |  | X | X | H, I, J |  |  |  | ü
863.4734 | [f+H]+ | SFLINNR (from THNLEPYFESFINNLR) | Peptide | keratin (human), ion source fragment |  | X | X | G, K
865.4778 | [f+H]+ | IEIATYR (from LALDIEIATYR) | Peptide | keratin (human), ion source fragment |  | X | X | G, K
865.5495 | [AB10+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
865.5859 | [AB14+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
867.0874 | [M4-4H+5Na]+ | C10H7NO3 | 4-HCCA | matrix cluster |  | X | X | H, I, J |  |  |  | ü
867.5676 | [AB15+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
869.5598 | [A14B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
870.5407 | [M+H]+ | VATVSLPR N-term. methylated | Peptide | porcine trypsin - methylated |  | X | X | J
871.0453 | [M4-2H+3K]+ | C10H7NO3 | 4-HCCA | matrix cluster |  | X | X | H, I, J |  |  |  | ü
871.4996 | [M+H]+ | QATVSLPR | Peptide | trypsin-like artefact |  | X | X | E
874.4993 | [M+H]+ | SLVNLGGSK | Peptide | keratin (human) |  | X | X | E, G, K
877.4979 | [A19B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
879 |  |  |  | Unknown |  | X |  |  |  | ü
881.4727 | [M+H]+ | SLYGLGGSK | Peptide | keratin (human) |  | X | X | G, K
881.5832 | [AB15+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
883.5148 | [M+H]+ | RVYVHPI | Peptide | Angiotensin standards |  | X | X | E
889.5495 | [AB15+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
889.6458 | [A15B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
891.5652 | [AB10+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
893.4718 | [A19B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
893.5808 | [AB10+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
895.5965 | [AB15+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
897.5393 | [AB12+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
899.5309 | [M+H]+ | VQTVSLPR | Peptide | trypsin-like artefact |  | X | X | E
899.5421 | [A20B+H]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
903.5652 | [AB15+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
905.6798 | [M+H]+ | C48H88N8O8 | nylon | Cyclic oligomer of polyamide 66, (adipic acid/hexylmethylene diamine condensation) |  | X | X | M |  |  |  | ü
906.5043 | [M+H]+ | NKPGVYTK | Peptide | bovine trypsin |  | X | X | G, K
906.5043 | [M+H]+ | NKPGVYTK | Peptide | porcine trypsin |  | X | X | G, K
907.9934 | [M4-H4+K4]+ | C10H7NO3 | 4-HCCA | Matrix compound |  |  | X | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. |  |  |  | ü
909.5757 | [AB11+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
909.6121 | [AB15+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
911.5938 | [AB16+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
911.6278 | [A15B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
917.4992 | [M+H]+ | RVYVHPF | Peptide | Angiotensin standards (bovine) |  | X | X | E
921.5241 | [A20B+Na]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
925.6094 | [AB16+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
927.6017 | [A15B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
931.5148 | [M+H]+ | RVYIHPF | Peptide | Angiotensin standards (human) |  | X | X | E
933 |  |  |  | Unknown |  | X |  |  |  | ü
933.5757 | [AB16+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
935.5914 | [AB11+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
937.498 | [A20B+K]+ | [C2H4O]nH2O | PEG | Polyethylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
937.607 | [AB11+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
939.6227 | [AB16+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
941.5656 | [AB13+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
947.5914 | [AB16+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
947.6877 | [A16B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
950.473 | [M+H]+ | YVNWIQQ | Peptide | trypsin-like artefact, truncated |  | X | X | E
953.6019 | [AB12+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
953.6383 | [AB16+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
955.62 | [AB17+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
969.6356 | [AB17+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
969.6696 | [A16B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
973.5313 | [M+H]+ | IEISELNR | Peptide | keratin (human) |  | X | X | E, G, K
977.6019 | [AB17+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
979.5095 | [M+H]+ | GTSYPDVLK | Peptide | bovine trypsin |  | X | X | G, K
979.6176 | [AB12+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
981.6332 | [AB12+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
983.6489 | [AB17+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
985.5918 | [AB14+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
985.6436 | [A16B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
991.6176 | [AB17+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
994.1555 | [M5-H2O-2H+3Na]+ | C10H7NO3 | 4-HCCA | matrix cluster |  | X | X | H, I, J |  |  |  | ü
995.5197 | [M+H]+ | IKEWYEK | Peptide | keratin |  | X | X | E
997.6282 | [AB13+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
997.6645 | [AB17+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
999.6462 | [AB18+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
1000.5673 | [M+H]+ | LQAEIEGLK | Peptide | keratin |  | X | X | E
1002.583 | [M+H]+ | LVVSTQTALA, C-terminus | Peptide | bovine serum albumin (BSA) |  | X | X | E
1003.5531 | [M+H]+ | SEITELRR | Peptide | keratin |  | X | X | E
1005.7295 | [A17B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1013.6619 | [AB18+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
1020.503 | [M+H]+ | APILSDSSCK | Peptide | bovine trypsin |  | X | X | G, K
1020.536 | [M+H]+ | SIPYQVSLN | Peptide | trypsin-like artefact, truncated |  | X | X | E
1021.6282 | [AB18+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
1023.4742 | [M+H]+ | SSYPGQITGN | Peptide | trypsin-like artefact, truncated |  | X | X | E
1023.5105 | [M+H]+ | NEQFISASK | Peptide | bovine trypsin |  | X | X | G, K
1023.6438 | [AB13+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1025.6595 | [AB13+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1027.6751 | [AB18+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
1027.7115 | [A17B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1029.618 | [AB15+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1031.5983 | [M+H]+ | VLDELTLTK | Peptide | keratin (human) |  | X | X | E, G, K
1033.516 | [M+H]+ | TLLEGEESR | Peptide | keratin (human) |  | X | X | E, G, K
1035.6438 | [AB18+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
1036.0828 | [M5-2H2O-4H+4Na+K]+ | C10H7NO3 | 4-HCCA | matrix cluster |  | X | X | H, I, J |  |  |  | ü
1036.5323 | [M+H]+ | IRDWYQR | Peptide | keratin |  | X | X | E
1037.5262 | [M+H]+ | YLDGLTAER | Peptide | keratin (human) |  | X | X | G, K
1041.6544 | [AB14+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1041.6908 | [AB18+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
1043.6724 | [AB19+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
1043.6854 | [A17B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1044.114 | [M5-2H+Na+2K]+ | C10H7NO3 | 4-HCCA | matrix cluster |  | X | X | H, I, J |  |  |  | ü
1045.5637 | [M+H]+ | LSSPATLNSR | Peptide | porcine trypsin |  | X | X | G, K
1046.5953 | [M+H]+ | LKSAASLNSR | Peptide | bovine trypsin |  | X | X | G, K
1054.0933 | [M5-H2O-4H+4Na+K]+ | C10H7NO3 | 4-HCCA | matrix cluster |  | X | X | H, I, J |  |  |  | ü
1057.6881 | [AB19+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
1060.0879 | [M5-2H+3K]+ | C10H7NO3 | 4-HCCA | matrix cluster, pentamer, triply potassiated |  | X | X | H, I, J |  |  |  | ü
1060.5633 | [M+H]+ | TLLDIDNTR | Peptide | keratin (human) |  | X | X | E, G, K
1063.7714 | [A18B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1064.5735 | [M+H]+ | DIIRAVGAYS | Peptide | lysyl endopeptidase (Lys-C) |  | X | X | E
1064.6099 | [M+H]+ | LASYLDKVR | Peptide | keratin (human) |  | X | X | G, K
1065.4993 | [M+H]+ | STMQELNSR | Peptide | keratin |  | X | X | E
1065.6544 | [AB19+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
1066.0959 | [M5-3H+2K+2Na]+ | C10H7NO3 | 4-HCCA | matrix cluster, pentamer, potassiated, sodiated |  | X | X | H, I, J |  |  |  | ü
1066.6044 | [M+H]+ | FAAFIDKVR | Peptide | keratin (human) |  | X | X | G, K
1067.67 | [AB14+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1069.6857 | [AB14+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1071.6045 | [M+H]+ | LAGLEEALQK | Peptide | keratin (human) |  | X | X | G, K
1071.7013 | [AB19+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
1073.6442 | [AB16+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1079.67 | [AB19+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
1080.532 | [f+H]+ | GNEQFISASK (from LGEDNINVVEGNEQFISASK) | Peptide | bovine trypsin, in source fragment |  | X | X | G, K
1082.5993 | [M+H]+ | FASFIDKVR | Peptide | keratin (human) |  | X | X | G, K
1085.6806 | [AB15+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1085.717 | [AB19+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
1085.7534 | [A18B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1087.6986 | [AB20+H]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
1090.531 | [M+H]+ | VTMQNLNDR | Peptide | keratin (human) |  | X | X | G, K
1091.548 | [M+H]+ | GNEQFINAAK | Peptide | porcine trypsin |  | X | X | G, K |  |  |  | ü
1101.7143 | [AB20+H]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
1101.7273 | [A18B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1109.4898 | [M+H]+ | DAEAWFNEK | Peptide | keratin (human) |  | X | X | G, K
1109.6806 | [AB20+Na]+ | [C14H22O][C2H4O]n | Triton | X-100, X-114, X-405, or X-45 Detergents |  | X | X | D, K |  |  |  | ü
1111.5605 | [M+H]+ | VCNYVSWIK | Peptide | bovine trypsin |  | X | X | G, K
1111.6962 | [AB15+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1113.7119 | [AB15+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1115.7275 | [AB20+Na]+ | [C14H28O][C2H4O]n | Triton, reduced | X-100R, X-114R, X-405R, or X-45R Detergents |  | X | X | D, K |  |  |  | ü
1117.6704 | [AB17+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1118.5086 | [M+H]+ | HGNSHQGEPR | Peptide | keratin (human) |  | X | X | G, K
1121.5837 | [M+H]+ | QEYEQLIAK | Peptide | keratin (human) |  | X | X | G, K
1121.8133 | [A19B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1123.6962 | [AB20+Na]+ | [C15H24O][C2H4O]n | Triton | 101 Detergents |  | X | X | D, K |  |  |  | ü
1126.564 | [M+H]+ | IITHPNFNGN | Peptide | porcine trypsin, methylated, truncated |  | X | X | E, J
1129.6212 | [M+H]+ | LSELEAALQR | Peptide | Keratin |  | X | X | E
1129.7068 | [AB16+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1129.7432 | [AB20+Na]+ | [C15H30O][C2H4O]n | Triton, reduced | 101R Detergents |  | X | X | D, K |  |  |  | ü
1141.5194 | [M+H]+ | DYQELMNTK | Peptide | Keratin |  | X | X | E
1143.5503 | [M+H]+ | VCNYVSWIK (W - oxid.) | Peptide | bovine trypsin |  | X | X | G, K
1143.7952 | [A19B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1153.5735 | [M+H]+ | SSGTSYPDVLK | Peptide | bovine trypsin |  | X | X | G, K
1155.7224 | [AB16+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1157.5909 | [M+H]+ | QGVDADINGLR | Peptide | keratin (human) |  | X | X | G, K
1157.7381 | [AB16+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1159.7692 | [A19B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1161.6966 | [AB18+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1165.5848 | [M+H]+ | LENEIQTYR | Peptide | keratin (human) |  | X | X | E, G, K
1168.5819 | [M+H]+ | VCcamNYVSWIK | Peptide | bovine trypsin |  | X | X | G, K
1173.733 | [AB17+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1175.5224 | [M+H]+ | MFCAGYLEGGK | Peptide | bovine trypsin |  | X | X | J
1175.634 | [M+H]+ | TLDNDIMLIK | Peptide | porcine trypsin |  | X | X | E, G, K
1179.6004 | [M+H]+ | YEELQITAGR | Peptide | keratin (human) |  | X | X | E, G, K
1179.8551 | [A20B+H]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1183.6157 | [M+H]+ | NVEIDPEIQK | Peptide | keratin (human) |  | X | X | G, K
1190.6086 | [M+H]+ | QVLDNLTMEK | Peptide | kereatin |  | X | X | E
1191.5173 | [M+H]+ | MoxFCAGYLEGGK | Peptide | bovine trypsin |  | X | X | J
1191.629 | [M+H]+ | TLDNDIMoxLIK | Peptide | porcine trypsin |  | X | X | E, G, K
1193.6161 | [M+H]+ | YEELQVTVGR | Peptide | keratin (human) |  | X | X | G, K
1199.7487 | [AB17+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1200.5718 | [M+H]+ | VCcamNYVSWIK (W - oxid.) | Peptide | bovine trypsin |  | X | X | G, K
1201.7643 | [AB17+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1201.8371 | [A20B+Na]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1203.6402 | [M+H]+ | TLDNDIMLIR | Peptide | trypsin-like artefact |  | X | X | E
1205.7228 | [AB19+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1208.6157 | [M+H]+ | TAAENDFVTLK | Peptide | keratin (human) |  | X | X | G, K
1210.595 | [M+H]+ | SSGGTSYPDVLK | Peptide | keratin (human) |  | X | X | G, K
1210.595 | [M+H]+ | SSGGTSYPDVLK | Peptide | bovine trypsin |  | X | X | G, K
1217.7592 | [AB18+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1217.811 | [A20B+K]+ | [C3H6O]nH2O | PPG | Polypropylene glycol, ubiquitous polyether |  | X | X | D, K |  |  |  | ü
1219.6351 | [M+H]+ | TLDNDIMoxLIR | Peptide | trypsin-like artefact |  | X | X | E
1220.5906 | [M+H]+ | EGNEQFINAAK | Peptide | porcine trypsin |  | X | X | G, K
1225.1254 | [M6-2H2O-4H+4Na+K]+ | C10H7NO3 | 4-HCCA | matrix cluster, hexamer, potassiated, sodiated |  | X | X | H, I, J |  |  |  | ü
1229.8002 | [M2+H]+ | C32H58N2O7S | CHAPS | 3-[(3-Cholamidopropyl)dimethylammonio]-1-propanesulfonate |  | X | X | E |  |  |  | ü
1232.5438 | [M+H]+ | MFCcamAGYLEGGK | Peptide | bovine trypsin |  | X | X | J
1232.5978 | [M+H]+ | SGGGGGGGLGSGGSIR | Peptide | keratin (human) |  | X | X | G, K
1235.5287 | [M+H]+ | FSSSSGYGGGSSR | Peptide | keratin (human) |  | X | X | G, K
1243.7749 | [AB18+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1245.7905 | [AB18+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1248.5388 | [M+H]+ | MoxFCcamAGYLEGGK | Peptide | bovine trypsin |  | X | X | J
1249.7491 | [AB20+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1251.7096 | [M+H]+ | YIPIQYVLSR | Peptide | bovine casein alpha-S1 |  | X | X | E
1254.6073 | [M+H]+ | GFSSGSAVVSGGSR | Peptide | keratin (human) |  | X | X | G, K
1261.7854 | [AB19+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1262.5971 | [M+H]+ | SLLEGEGSSGGGGR | Peptide | keratin |  | X | X | E
1262.7579 | [M+H]+ | LLHGVATVSLPR | Peptide | trypsin-like artefact |  | X | X | E
1265.6372 | [M+H]+ | TNAENEFVTIK | Peptide | keratin (human) |  | X | X | G, K
1267.7045 | [M+H]+ | YLGYLEQLLR | Peptide | bovine casein alpha-S1 |  | X | X | E
1277.6332 | [M+H]+ | SDQSRLDSELK | Peptide | keratin (human) |  | X | X | G, K
1277.71 | [M+H]+ | LALDIEIATYR | Peptide | keratin (human) |  | X | X | E, G, K
1278.5855 | [M+H]+ | GSCcamGIGGGIGGGSSR | Peptide | keratin |  | X | X | E
1278.6437 | [M+H]+ | YEQLAEQNRK | Peptide | keratin (human) |  | X | X | G, K
1282.6691 | [M+H]+ | DRVYVHPFHL | Peptide | Angiotensin standards (bovine) |  | X | X | E
1287.8011 | [AB19+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1289.8167 | [AB19+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1293.7753 | [AB21+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1296.6848 | [M+H]+ | DRVYIHPFHL | Peptide | Angiotensin standards (human) |  | X | X | E
1301.6583 | [M+H]+ | ALEEANADLEVK | Peptide | keratin |  | X | X | E
1302.7012 | [M+H]+ | NSKIEISELNR | Peptide | keratin (human) |  | X | X | G, K
1302.7151 | [M+H]+ | SLDLDSIIAEVK | Peptide | keratin (human) |  | X | X | E, G, K
1305.7161 | [M+H]+ | HLVDEPQNLIK | Peptide | bovine serum albumin (BSA) |  | X | X | E
1305.8117 | [AB20+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1307.6776 | [M+H]+ | IKFEMEQNLR | Peptide | keratin (human) |  | X | X | E, G, K
1308.643 | [M+H]+ | VEGNEQFISASK | Peptide | bovine trypsin |  | X | X | G, K
1308.6543 | [M+H]+ | NKYEDEINKR | Peptide | keratin (human) |  | X | X | G, K
1315.6852 | [M+H]+ | DQIVDLTVGNNK | Peptide | keratin |  | X | X | E
1316.7243 | [M+H]+ | TLDNDIMLIRL | Peptide | trypsin-like artefact |  | X | X | E
1320.5828 | [M+H]+ | HGGGGGGFGGGGFGSR | Peptide | keratin (human) |  | X | X | G, K
1323.6726 | [M+H]+ | IKFEMoxEQNLR | Peptide | keratin (human) |  | X | X | G, K
1329.6393 | [M+H]+ | NVQDAIADAEQR | Peptide | keratin (human) |  | X | X | G, K
1331.8273 | [AB20+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1332.7192 | [M+H]+ | TLDNDIMoxLIRL | Peptide | trypsin-like artefact |  | X | X | E
1333.843 | [AB20+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1336.7107 | [M+H]+ | TAAENDFVTLKK | Peptide | keratin (human) |  | X | X | G, K
1337.8015 | [AB22+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1340.6692 | [M+H]+ | SKAEAESLYQSK | Peptide | keratin |  | X | X | E
1344.6754 | [M+H]+ | ASLEAAIADAEQR | Peptide | keratin |  | X | X | E
1349.8379 | [AB21+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1350.6873 | [M+H]+ | IGLGGRGGSGGSYGR | Peptide | keratin |  | X | X | E
1357.6958 | [M+H]+ | LNDLEDALQQAK | Peptide | keratin (human) |  | X | X | E, G, K
1357.7182 | [M+H]+ | QSVEADINGLRR | Peptide | keratin (human) |  | X | X | G, K
1365.6393 | [M+H]+ | SQYEQLAEQNR | Peptide | keratin (human) |  | X | X | G, K
1366.6961 | [M+H]+ | RISNSTSPTSFVA | Peptide | lysyl endopeptidase (Lys-C) |  | X | X | E
1367.6702 | [M+H]+ | VFTSWTGGGAAASR | Peptide | lysyl endopeptidase (Lys-C) |  | X | X | E
1371.7114 | [M+H]+ | LNDLEEALQQAK | Peptide | keratin (human) |  | X | X | E, G, K
1374.6933 | [M+H]+ | DLNMDNIVAEIK | Peptide | keratin (human) |  | X | X | G
1375.8535 | [AB21+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1377.8692 | [AB21+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1381.6482 | [M+H]+ | ALEESNYELEGK | Peptide | keratin (human) |  | X | X | E, G
1381.8277 | [AB23+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1383.6903 | [M+H]+ | SLNNQFASFIDK | Peptide | keratin (human) |  | X | X | E, G
1384.73 | [M+H]+ | FFVAPFPEVFGK | Peptide | bovine casein alpha-S1 |  | X | X | E
1390.6809 | [M+H]+ | QSLEASLAETEGR | Peptide | keratin |  | X | X | E
1393.7322 | [M+H]+ | TNAENEFVTIKK | Peptide | keratin (human) |  | X | X | E, G
1393.8641 | [AB22+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1399.6926 | [M+H]+ | TVMENFVAFVDK | Peptide | bovine serum albumin (BSA) |  | X | X | E
1399.7005 | [f+H]+ | PYFESFINNLR (from THNLEPYFESFINNLR) | Peptide | keratin (human), ion-source fragment |  | X | X | G, K
1418.7274 | [M+H]+ | VDALNDEINFLR | Peptide | keratin |  | X | X | E
1419.7478 | [M+H]+ | LEGLTDEINFLR | Peptide | keratin |  | X | X | E
1419.7954 | [M+H]+ | LGLDIEIATYRR | Peptide | keratin (human) |  | X | X | G
1419.8797 | [AB22+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1421.8954 | [AB22+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1424.7645 | [M+H]+ | AAKIITHPNFNGN | Peptide | porcine trypsin, methylated |  | X | X | J
1425.8539 | [AB24+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1433.7206 | [M+H]+ | LQGIVSWGSGCAQK | Peptide | bovine trypsin |  | X | X | G
1434.77 | [M+H]+ | IRLENEIQTYR | Peptide | keratin (human) |  | X | X | G
1437.8903 | [AB23+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1438.7437 | [M+H]+ | NNQFASFIDKVR | Peptide | keratin (human) |  | X | X | G
1439.8118 | [M+H]+ | RHPEYAVSVLLR | Peptide | bovine serum albumin (BSA) |  | X | X | E
1442.8002 | [M+H]+ | LPDIFEAQIAGLR | Peptide | keratin |  | X | X | E
1443.642 | [M+H]+ | YICcamDNQDTISSK | Peptide | bovine serum albumin (BSA) |  | X | X | E
1447.7751 | [M+H]+ | AIGGGLSSVGGGSSTIK | Peptide | keratin (human) |  | X | X | G
1449.7155 | [M+H]+ | LQGIVSWGSGCAQK (W - oxid. I) | Peptide | bovine trypsin |  | X | X | G
1453.8373 | [M+H]+ | EVTINQSLLAPLR | Peptide | keratin |  | X | X | E
1463.906 | [AB23+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1465.7104 | [M+H]+ | LQGIVSWGSGCAQK (W - oxid. II) | Peptide | bovine trypsin |  | X | X | G
1465.9216 | [AB23+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1469.8801 | [AB25+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1475.7489 | [M+H]+ | WELLQQVDTSTR | Peptide | keratin (human) |  | X | X | E, G
1475.7853 | [M+H]+ | FLEQQNQVLQTK | Peptide | keratin (human) |  | X | X | E, G
1476.8057 | [M+H]+ | FLEQQNKVLETK | Peptide | keratin |  | X | X | E
1479.7954 | [M+H]+ | LGEYGFQNALIVR | Peptide | bovine serum albumin (BSA) |  | X | X | E
1481.9165 | [AB24+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1490.742 | [M+H]+ | LQGIVSWGSGCcamAQK | Peptide | bovine trypsin |  | X | X | G
1491.7438 | [M+H]+ | WELLQQVDTSTR (W - oxid. I) | Peptide | keratin (human) |  | X | X | G
1491.7518 | [M+H]+ | FYAPELLYYANK | Peptide | bovine serum albumin (BSA) |  | X | X | E
1493.5995 | [M+H]+ | DSCQGDSGGPVVCSGK ox. S-S bond | Peptide | porcine trypsin |  | X | X | J
1493.7343 | [M+H]+ | SQYEQLAEQNRK | Peptide | keratin (human) |  | X | X | E, G
1495.6152 | [M+H]+ | DSCQGDSGGPVVCSGK | Peptide | bovine trypsin |  | X | X | G
1507.7387 | [M+H]+ | WELLQQVDTSTR (W - oxid. II) | Peptide | keratin (human) |  | X | X | G
1507.9322 | [AB24+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1509.9478 | [AB24+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1513.9063 | [AB26+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1521.7544 | [M+H]+ | NVVEGNEQFISASK | Peptide | bovine trypsin |  | X | X | G
1522.7319 | [M+H]+ | LQGIVSWoxGSGCcamAQK | Peptide | bovine trypsin |  | X | X | G
1525.9427 | [AB25+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1549.6772 | [M+H]+ | SGGGGGGGGCcamGGGGGVSSLR | Peptide | keratin |  | X | X | E
1551.9584 | [AB25+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1553.974 | [AB25+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1557.9326 | [AB27+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1566.7394 | [M+H]+ | LGEHNIDVLEGNEQ | Peptide | porcine trypsin, truncated |  | X | X | E, G
1567.7427 | [M+H]+ | DAFLGSFLYEYSR | Peptide | bovine serum albumin (BSA) |  | X | X | E
1569.9689 | [AB26+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1585.7584 | [M+H]+ | VQALEEANNDLENK | Peptide | keratin (human) |  | X | X | E, G
1595.9846 | [AB26+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1598.0003 | [AB26+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1599.8337 | [M+H]+ | NKLNDLEDALQQAK | Peptide | keratin |  | X | X | E
1601.9588 | [AB28+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1613.9952 | [AB27+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1623.8853 | [M+H]+ | EQFINAAKIITHPN | Peptide | porcine trypsin, methylated |  | X | X | J
1638.7657 | [M+H]+ | SLNNQFASFIDKVR | Peptide | keratin (human) |  | X | X | G
1639.9378 | [M+H]+ | KVPQVSTPTLVEVSR | Peptide | bovine serum albumin (BSA) |  | X | X | E
1640.0108 | [AB27+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1642.0265 | [AB27+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1645.985 | [AB29+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1657.7929 | [M+H]+ | SGGGFSSGSAGIINYQR | Peptide | keratin (human) |  | X | X | E, G
1658.0214 | [AB28+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1665.9044 | [M+H]+ | TKPSQARGFHPRAGR | Peptide | keratin (human) |  | X | X | G
1676.7762 | [M+H]+ | LGEDNINVVEGNEQF | Peptide | bovine trypsin |  | X | X | J
1684.037 | [AB28+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1686.0527 | [AB28+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1687.9701 | [M+H]+ | SLVNLGGSKSISISVAR | Peptide | keratin (human) |  | X | X | G
1690.0112 | [AB30+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1702.0476 | [AB29+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1707.7721 | [M+H]+ | GSLGGGFSSGGFSGGSFSR | Peptide | keratin (human) |  | X | X | E, G
1708.7132 | [M+H]+ | FSSCGGGGGSFGAGGGFGSR | Peptide | keratin (human) |  | X | X | G
1713.8078 | [M+H]+ | LGEHNIDVLEGNEQF | Peptide | porcine trypsin, truncated |  | X | X | E, G
1716.8511 | [M+H]+ | QISNLQQSISDAEQR | Peptide | keratin (human) |  | X | X | E, G
1725.8629 | [M+H]+ | VCNYVSWIKQTIASN | Peptide | bovine trypsin |  | X | X | G
1728.0632 | [AB29+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1730.0789 | [AB29+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1734.0374 | [AB31+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1744.9228 | [M+H]+ | NNLEPILEGYISNLR | Peptide | keratin (human) |  | X | X | G
1746.0738 | [AB30+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1757.8527 | [M+H]+ | VCNYVSWIKQTIASN (W - oxid.) | Peptide | bovine trypsin |  | X | X | G
1765.7347 | [M+H]+ | FSSCcamGGGGGSFGAGGGFGSR | Peptide | keratin (human) |  | X | X | E, G
1772.0895 | [AB30+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1774.1051 | [AB30+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1774.8503 | [M+H]+ | IVGGYTCGANTVPY & CLK via S-S bond | Peptide | bovine trypsin |  | X | X | J
1774.897 | [M+H]+ | NIDVLEGNEQFINAAK | Peptide | porcine trypsin |  | X | X | G
1778.0636 | [AB32+Na]+ | [C18H34O6][C2H4O]n | Tween | Tween 20 |  | X | X | D, K |  |  |  | ü
1790.1 | [AB31+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1791.7277 | [M+H]+ | GGSGGSYGGGGSGGGYGGGSGSR | Peptide | keratin (human) |  | X | X | E, G
1792.8786 | [M+H]+ | LEAELGNMQGLVEDFK | Peptide | keratin |  | X | X | E
1797.0116 | [M+H]+ | NVQALEIELQSQLALK | Peptide | keratin |  | X | X | E
1816.1157 | [AB31+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1818.1313 | [AB31+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1830.8406 | [M+H]+ | SDQYGRVFTSWTGGGAAA | Peptide | lysyl endopeptidase (Lys-C) |  | X | X | E
1834.1262 | [AB32+Na]+ | [C22H42O6][C2H4O]n | Tween | Tween 40 |  | X | X | D, K |  |  |  | ü
1837.9654 | [M+H]+ | HGVQELEIELQSQLSK | Peptide | keratin |  | X | X | E
1838.9144 | [M+H]+ | SISISVAGGGGGFGAAGGFGGR | Peptide | keratin (human) |  | X | X | G
1847.8051 | [M+H]+ | SNMDNMFESYINNLR | Peptide | keratin |  | X | X | E
1851.9269 | [M+H]+ | TLNDMRQEYEQLIAK | Peptide | keratin |  | X | X | E
1860.1419 | [AB32+Na]+ | [C24H44O6][C2H4O]n | Tween | Tween 80 |  | X | X | D, K |  |  |  | ü
1862.1575 | [AB32+Na]+ | [C24H46O6][C2H4O]n | Tween | Tween 60 |  | X | X | D, K |  |  |  | ü
1940.9349 | [M+H]+ | LGEHNIDVLEGNEQFIN | Peptide | porcine trypsin, truncated, methylated |  | X | X | E, J
1987.078 | [M+H]+ | TLDNDIMLIKLSSPATLN, K methylated | Peptide | porcine trypsin, methylated |  | X | X | J
1993.9767 | [M+H]+ | THNLEPYFESFINNLR | Peptide | keratin (human) |  | X | X | E, K
1993.9825 | [M+H]+ | SSGGSSSVKFVSTTYSGVTR | Peptide | keratin (human) |  | X | X | G
2003.0729 | [M+H]+ | TLDNDIMoxLIKLSSPATLN, K methylated | Peptide | porcine trypsin, methylated |  | X | X | J
2021.9887 | [M+H]+ | SSGGSSSVRFVSTTYSGVTR | Peptide | keratin (human) |  | X | X | G
2055.0149 | [M+H]+ | RVLGQLHGGPSSCcamSATGTNR | Peptide | lysyl endopeptidase (Lys-C) |  | X | X | E
2082.9648 | [M+H]+ | AETECcamQNTEYQQLLDIK | Peptide | keratin |  | X | X | E
2083.0091 | [M+H]+ | LGEHNIDVLEGNEQFINAA | Peptide | porcine trypsin, truncated |  | X | X | E
2109.0128 | [M+H]+ | ELQSQISDTSVVLSMDNSR | Peptide | keratin |  | X | X | E
2150.0778 | [M+H]+ | THNLEPYFESFINNLRR | Peptide | keratin (human) |  | X | X | K
2155.0852 | [M+H]+ | IITHPNFNGNTLDNDIMLI | Peptide | trypsin-like artefact, truncated |  | X | X | E
2163.0564 | [M+H]+ | LGEDNINVVEGNEQFISASK | Peptide | bovine trypsin |  | X | X | G
2171.0247 | [M+H]+ | SDLEMQYETLQEELMALK | Peptide | keratin |  | X | X | E
2184.1255 | [M+H]+ | NKLNDLEDALQQAKEDLAR | Peptide | keratin |  | X | X | E
2185.0384 | [M+Na]+ | LGEDNINVVEGNEQFISASK sodiated | Peptide | bovine trypsin |  | X | X | G, K
2193.0025 | [M+H]+ | MFCAGYLEGGK & APILSDSSCK via S-S bond | Peptide | bovine trypsin |  | X | X | J
2193.9944 | [M+K]+ | SAYPGQITSNMFCAGYLEGGK | Peptide | bovine trypsin |  | X | X | G, K
2201.0123 | [M+H]+ | LGEDNINVVEGNEQFISASK potassiated | Peptide | bovine trypsin |  | X | X | G, K
2208.9974 | [M+H]+ | MoxFCAGYLEGGK & APILSDSSCK via S-S bond | Peptide | bovine trypsin |  | X | X | J
2209.9893 | [M+H]+ | SAYPGQITSNMoxFCAGYLEGGK | Peptide | bovine trypsin |  | X | X | G, K
2211.104 | [M+H]+ | LGEHNIDVLEGNEQFINAAK | Peptide | porcine trypsin |  | X | X | G, K
2225.1197 | [M+H]+ | LGEHNIDVLEGNEQFINAAK, K part. methylated | Peptide | porcine trypsin |  | X | X | G, K
2225.1561 | [f+H]+ | SIVHPSYNSNTLNNDIMoxLIK -[CH3SOH] | Peptide | bovine trypsin, ion source fragment |  | X | X | G, K
2239.1353 | [M+H]+ | LGEHNIDVLEGNEQFINAAK, K methylated | Peptide | porcine trypsin |  | X | X | G, K
2240.1697 | [M+H]+ | ADLEFQIESLTEELAYLKK | Peptide | keratin |  | X | X | E
2251.0158 | [M+H]+ | SAYPGQITSNMFCcamAGYLEGGK | Peptide | bovine trypsin |  | X | X | G
2267.0107 | [M+H]+ | SAYPGQITSNMoxFCcamAGYLEGGK | Peptide | bovine trypsin |  | X | X | G
2273.1594 | [M+H]+ | SIVHPSYNSNTLNNDIMLIK | Peptide | bovine trypsin |  | X | X | G, K
2283.1802 | [M+H]+ | IITHPNFNGNTLDNDIMLIK | Peptide | bovine trypsin |  | X | X | G, K
2283.1802 | [M+H]+ | IITHPNFNGNTLDNDIMLIK | Peptide | porcine trypsin |  | X | X | G, K
2286.1248 | [M+H]+ | AEAESLYQSKYEELQITAGR | Peptide | keratin (human) |  | X | X | G
2289.1544 | [M+H]+ | SIVHPSYNSNTLNNDIMoxLIK (M - oxid.) | Peptide | bovine trypsin |  | X | X | G, K
2295.1414 | [M+Na]+ | SIVHPSYNSNTLNNDIMLIK sodiated | Peptide | bovine trypsin |  | X | X | G, K
2297.1958 | [M+H]+ | IITHPNFNGNTLDNDIMLIK, K part. methylated | Peptide | porcine trypsin, methylated |  | X | X | J
2305.1493 | [M+H]+ | SIVHPSYNSNTLNNDIMoxoxLIK (M - 2 oxid.) | Peptide | bovine trypsin |  | X | X | G, K
2311.1363 | [M+Na]+ | SIVHPSYNSNTLNNDIMoxLIK (M - oxid.) sodiated | Peptide | bovine trypsin |  | X | X | G, K
2321.2071 | [M+H]+ | LQGIVSWGSGCAQKNKPGVYTK | Peptide | bovine trypsin |  | X | X | G, K
2329.1742 | [M+H]+ | QISNLQQSISDAEQRGENALK | Peptide | keratin |  | X | X | E
2343.189 | [M+Na]+ | LQGIVSWGSGCAQKNKPGVYTK sodiated | Peptide | bovine trypsin |  | X | X | G, K
2353.1969 | [M+H]+ | LQGIVSWGSGCAQKNKPGVYTK (W - oxid.) | Peptide | bovine trypsin |  | X | X | G, K
2367.2627 | [M+H]+ | NQILNLTTDNANILLQIDNAR | Peptide | keratin (human) |  | X | X | G
2375.1788 | [M+Na]+ | LQGIVSWGSGCAQKNKPGVYTK (W - oxid.) sodiated | Peptide | bovine trypsin |  | X | X | G, K
2383.9519 | [M+H]+ | GGGGGGYGSGGSSYGSGGGSYGSGGGGGGGR | Peptide | keratin (human) |  | X | X | G
2400.2744 | [M+H]+ | IIKHPNYSSWTLNNDIMLIK | Peptide | bovine trypsin |  | X | X | G, K
2501.0672 | [M+H]+ | AAFGGSGGRGSSSGGGYSSGSSSYGSGGR | Peptide | keratin (human) |  | X | X | G
2501.2518 | [M+H]+ | SKAEAESLYQSKYEELQITAGR | Peptide | keratin (human) |  | X | X | G
2510.1318 | [M+H]+ | EIETYHNLLEGGQEDFESSGAGK | Peptide | keratin (human) |  | X | X | E, G
2514.3385 | [M+H]+ | SIVHPSYNSNTLNNDIMLIKLK | Peptide | bovine trypsin |  | X | X | G, K
2530.3334 | [M+H]+ | SIVHPSYNSNTLNNDIMoxLIKLK | Peptide | bovine trypsin |  | X | X | G, K
2550.2327 | [M+H]+ | VASISLPTSCASAGTQCLISGWGNTK, -S-S-bond | Peptide | bovine trypsin |  | X | X | J
2552.2483 | [M+H]+ | VASISLPTSCASAGTQCLISGWGNTK | Peptide | bovine trypsin |  | X | X | J
2564.1809 | [M+H]+ | SGSHFCGGSLINSQWVVSAAHCYK, -S-S- bond | Peptide | porcine trypsin, methylated |  | X | X | J
2564.3679 | [M+H]+ | EVTQLRHGVQELEIELQSQLSK | Peptide | keratin |  | X | X | E
2566.2784 | [M+H]+ | SKEEAEALYHSKYEELQVTVGR | Peptide | keratin (human) |  | X | X | G
2581.1947 | [M+H]+ | FSGECcamAPNVSVSVSTSHTTISGGGSR | Peptide | keratin |  | X | X | E
2612.1809 | [M+H]+ | SGYHFCGGSLINSQWVVSAAHCYK, -S-S- bond | Peptide | bovine trypsin |  | X | X | J
2613.3494 | [M+H]+ | NKPGVYTKVCNYVSWIKQTIASN | Peptide | bovine trypsin |  | X | X | J
2614.1966 | [M+H]+ | SGYHFCGGSLINSQWVVSAAHCYK | Peptide | bovine trypsin |  | X | X | J
2666.2913 | [M+H]+ | VASISLPTSCcamASAGTQCcamLISGWGNTK | Peptide | bovine trypsin |  | X | X | J
2670.3708 | [M+H]+ | NKPGVYTKVCcamNYVSWIKQTIASN | Peptide | bovine trypsin |  | X | X | J
2705.1611 | [M+H]+ | GGGGSFGYSYGGGSGGGFSASSLGGGFGGGSR | Peptide | keratin |  | X | X | E
2707.4162 | [M+H]+ | IQVRLGEHNIDVLEGNEQFINAAK | Peptide | porcine trypsin, methylated |  | X | X | J
2720.3634 | [M+H]+ | IVGGYTCAAN & TKSSGSSYPSLLQCLK, via S-S bond | Peptide | porcine trypsin, methylated |  | X | X | J
2728.2395 | [M+H]+ | SGYHFCcamGGSLINSQWVVSAAHCcamYK | Peptide | bovine trypsin |  | X | X | J
2746.4192 | [M+H]+ | YCcamVQLSQIQAQISALEEQLQQIR | Peptide | keratin |  | X | X | E
2748.3947 | [M+H]+ | IVGGYTCAAN & TKSSGSSYPSLLQCLK, via S-S bond, N-term. Methyl. | Peptide | porcine trypsin, methylated |  | X | X | J
2807.3141 | [M+H]+ | SGSHFCGGSLINSQWVVSAAHCYKSR, -S-S- bond | Peptide | porcine trypsin, methylated |  | X | X | J
2872.3928 | [M+H]+ | NVSTGDVNVEMNAAPGVDLTQLLNNMR | Peptide | keratin |  | X | X | E
2902.4105 | [M+H]+ | NYSPYYNTIDDLKDQIVDLTVGNNK | Peptide | keratin |  | X | X | E
2904.3826 | [M+H]+ | NVSTGDVNVEMoxNAAPGVDLTQLLNNMoxR | Peptide | keratin |  | X | X | E
2914.5058 | [M+H]+ | LGEHNIDVLEGNEQFINAAKIITHPN, K methylated | Peptide | porcine trypsin, methylated |  | X | X | J
2932.5163 | [M+H]+ | FLEQQNQVLQTKWELLQQVDTSTR | Peptide | keratin (human) |  | X | X | G
3211.4745 | [M+H]+ | SAYPGQITSNMFCAGYLEGGK & APILSDSSCK via S-S bond | Peptide | bovine trypsin |  | X | X | J
3223.2816 | [M+H]+ | GGSGGSHGGGSGFGGESGGSYGGGEEASGSGGGYGGGSGK | Peptide | keratin |  | X | X | E
3227.4694 | [M+H]+ | SAYPGQITSNMoxFCAGYLEGGK & APILSDSSCK via S-S bond | Peptide | bovine trypsin |  | X | X | J
3312.3081 | [M+H]+ | GSYGSGGSSYGSGGGSYGSGGGGGGHGSYGSGSSSGGYR | Peptide | keratin (human) |  | X | X | G
3337.7321 | [M+H]+ | IITHPNFNGNTLDNDIMLIRLSSPATLNSR | Peptide | trypsin-like artefact, truncated |  | X | X | E
3346.6815 | [M+H]+ | LGEHNIDVLEGNEQFINAAKIITHPNFNGN, K methylated | Peptide | porcine trypsin, methylated |  | X | X | J
3353.7271 | [M+H]+ | IITHPNFNGNTLDNDIMoxLIRLSSPATLNSR | Peptide | trypsin-like artefact, truncated |  | X | X | E
Note on K93: unidentified contaminant
Note on K99: Unknown† Derived from rubber tip from disposable syringe plunger

## Negative
adapted from supplementary Data:
Keller, B.O.; Sui, J.; Young, A.B.; Whittal, R.M. Interferences and contaminants encountered in modern mass spectrometry.
Analytica Chimica Acta (Review/tutorial, Special Issue on Mass Spectrometry), 2008.
List of potential interference- or contaminant ions in modern mass spectrometry, negative mode
Monoisotopic ion mass (singly charged) | Ion type | Formula for M or subunit or sequence |  | Possible origin and other comments |  | References | Waters | MaConDa | Agilent
[7 empty] | http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf | http://www.maconda.bham.ac.uk/index.php | Common LC/MS Contaminants - CIGS
26.0036 | f- | CN | CN- | Fragment from acetonitrile |  | A,F | ü | ü
44.9982 | [M-H]- | HCOOH | (HCOO)- | Formic acid, FA (anion), Formate |  | A,F | ü | ü
59.0139 | [M-H]- | CH3COOH | (CH3CO2)- | Acetic Acid, Ac (anion), Acetate |  | A,F | ü | ü
69 |  |  | (CF3CO2Na)- | sodium trifluoracetate, TFA |  | A | ü
78.9591 | [M-H3O]- | H3PO4 | PO3- | Phosphoric acid (phosphate, anion) also from oligonucleotides or phosphopeptides |  | A,F | ü | ü
79.9574 | [M-H2O]- | H2SO4 | SO3- | Sulfuric acid (sulfate, anion) also from other sulfated materials |  | A,F | ü | ü
94.9808 | [M-H]- | CH3SO3H | CH3SO3- | Methanesulfuric acid (methanesulfonate, anion) |  | A,F | ü | ü
96.9601 | [M-H]- | H2SO4 | HSO4- | Sulfuric acid (sulfate, anion) |  | A,F | ü
96.9696 | [M-H]- | H3PO4 | H2PO4- | Phosphoric acid (phosphate, anion) also from oligonucleotides or phosphopeptides |  | A,F | ü | ü
112.9856 | [M-H]- | CF3COOH |  | Trifluoroacetic acid, TFA (anion) | Common LC buffer used in ESI Pos mode. Will see acetate anion as a background contaminate in ESI Neg mode. Rinse LC system. If need to continue to use TFA as buffer, don’t scan down to 113. | A,F | ü | ü
112.9856 | [M2+Na-2H]- | HCOOH |  | Formic acid, dimer, FA (anion) |  | F |  | ü
126.905 | A- | I |  | Iodine anion (Iodide) |  | V |  | ü
127 |  |  | (CF3CO2Na)- | sodium trifluoracetate, TFA |  |  | ü
141.0197 |  | C4H5O2F3 |  | Acetate adducts of a fluorinated polymer (PTFE). ; From lithium grease used on 1100 LC pump modules. |  |  |  |  | ü
162.9824 | [M-H]- | CF3CF2COOH |  | Pentafluoropropionic acid, Perfluoropropionic acid |  | E |  | ü
212.0751 | [M-H]- | C10H15NO2S |  | n-butyl benzenesulfonamide |  | A, K |  | ü
223.0278 |  | C6H6O2F6 |  | Acetate adducts of a fluorinated polymer (PTFE). ; From lithium grease used on 1100 LC pump modules. |  |  |  |  | ü
226.9785 | [M2-H]- | CF3COOH | (CF3CO2Na)- | Trifluoroacetic acid, TFA (anion) |  | A | ü | ü
248.9604 | [M2+Na-2H]- | CF3COOH | (CF3CO2Na)- | Trifluoroacetic acid, TFA, sodiated (anion) |  | A | ü | ü
255.233 | [M-H]- | C16H32O2 |  | Palmitic acid (carboxylate anion) C16:0 |  | K |  | ü
281.2486 | [M-H]- | C18H34O2 |  | Oleic acid (carboxylate anion) C18:1 |  | K |  | ü
283.2643 | [M-H]- | C18H36O2 |  | Stearic acid (carboxylate anion) 18:0 |  | K |  | ü
305.0345 |  | C8H7O2F9 |  | Acetate adducts of a fluorinated polymer (PTFE). ; From lithium grease used on 1100 LC pump modules. |  |  |  |  | ü
316.7391 | [63CuI+I]- | CuI |  | Copper catalyst iodide adduct |  | K |  | ü
318.7373 | [65CuI+I]- | CuI |  | Copper catalyst iodide adduct |  | K |  | ü
387.0421 |  | C10H8O2F12 |  | Acetate adducts of a fluorinated polymer (PTFE). ; From lithium grease used on 1100 LC pump modules. |  |  |  |  | ü
469.043 |  | C12H9O2F15 |  | Acetate adducts of a fluorinated polymer (PTFE). ; From lithium grease used on 1100 LC pump modules. |  |  |  |  | ü
551.034 |  | C14H10O2F18 |  | Acetate adducts of a fluorinated polymer (PTFE). ; From lithium grease used on 1100 LC pump modules. |  |  |  |  | ü

## Repeating Units
REPEATING UNITS, positive mode
Observed Masses | Mass diff. | Formula for M or subunit or sequence | Compound ID or species | Origin | mode +/- |  | Reference(s) | Waters | Leiden Univ. | Agilent | NewObjective
[8 empty] | http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf | http://www.lc-ms.nl/ | Common LC/MS Contaminants - CIGS | http://www.newobjective.com/downloads/technotes/PV-3.pdf
| 14.0157 | [CH2] |  | Alkane chains, waxes, fatty acids, methylation | + | +14n | F |  | ü
| 15.9949 | O |  | Oxidation | + | +15n | F |  | ü
| 18.0106 | H2O |  | Water clusters | +/- | +18n | F | ü | ü
| 28.0313 | [C2H4] |  | Natural alkane chains such as fatty acids | + | +28n | F |  | ü
| 32.0262 | [CH3OH] |  | Methanol clusters | + | +32n | F | ü | ü
| 41.0265 | [CH3CN] |  | Acetonitrile clusters | + | +41n | F | ü | ü
| 42.047 | [C3H6] |  | Propyl repeating units, propylation | + | +42n | F |  | ü
| 44.0262 | [C2H4O] |  | Polyethylene glycol, PEG, and related components such as Tritons and Tween buffers; Possible polymer contamination. If you have shot detergent containing samples on the system that could explain the background. Also PEGs and other ethoxylated polymers give +44 ion series. The PEG's could also be from the water or extracted polymer from plastic ware/silicon coatings. | + | +44n | D, F | ü | ü | ü
| 49.9968 | [CF2] |  | From perfluoro compounds | + | +50n | F | ü | ü
| 53.0032 | NH4Cl |  | Ammonium chloride adducts (NH4Cl) salt adducts/clusters | + | +53n | F | ü | ü
| 56.0626 | [C4H8] |  | Butyl repeating units, butylation | + | +56n | F |  | ü
| 57.9586 | NaCl |  | NaCl, sodium chloride clusters (distinguishable from PPG by Cl Isotope pattern) | + | +58n | F | ü | ü
| 58.0419 | [C3H6O] |  | Polypropylene glycol and related compounds, PPG, and related compounds | + | +58n | D, F | ü | ü
| 63 | NH4Cl |  | Salts |  |  |  | ü
| 63.032 | [CHOONH4] |  | Ammonium formate adducts/clusters | + | +63n | F |  | ü
| 67.9874 | [NaHCO2] |  | Sodium formate clusters |  |  | F | ü
| 67.9874 | [CHOONa] |  | Sodium formate adducts/clusters | +/- | +68n | F |  | ü
| 72.0395 | OH |  | Replacement of OH by OSi(CH3)3; -OH replacement with -OSi(CH3)3,(=[C3H8Si]), trimethylsiloxane, endcapping reagent | + | +72n | F | ü | ü
| 73.9326 | KCl |  | KCl adducts (distinguishable by Cl isotope pattern) | + | +74n | F | ü | ü
| 74.0188 | [O-Si(CH3)2] |  | Polysiloxane, silicone rubber polymer (typical series at m/z's 355, 429, 503, 593, 667, 741, 815); Si(CH3)2O (from rubber) | +/- | +74n | F, E, L | ü | ü
| 77.0477 | NH4Ac |  | Ammonium acetate salts (CH3COONH4) | + | +77n |  | ü | ü
| 78.0139 | [C2H6OS] |  | DMSO adducts/clusters, dimethylsulfoxide solvent | + | +78n | F | ü | ü
| 82.0031 | [NaCH3CO2] |  | Sodium acetate clusters | +/- | +82n | F | ü | ü
| 84.0516 | [C2D6OS] |  | Deuterated DMSO adducts/clusters, NMR solvent |  |  | F
| 106.9051 | 107Ag |  | 107Ag, silver clusters in MALDI analysis of non-polar polymers, together with 109Ag (~1:1) |  |  | U
| 108.9048 | 109Ag |  | 109Ag, silver clusters in MALDI analysis of non-polar polymers, togethre with 107Ag ( ~1:1) |  |  | U
| 113.9929 | CF3COOH |  | TFA (trifluoroacetic acid) adducts (CF3COOH) | - | +114n |  |  | ü | ü
| 121.9383 | [NaClO4] |  | Sodium perchlorate adducts (NaClO4) | - | +122n |  |  | ü
| 135.9748 | NaCF3CO2 |  | Sodium TFA (trifluoroacetic acid) adducts (CF3COONa) | +/- | +136n | F | ü | ü
| 162.0528 | [C6H10O5] |  | Polysaccharides (C6H10O5) | + | +162n | F | ü | ü
| 226.1681 | [C12H22N2O2] |  | Cyclic oligomers from polyamide 66 (series observed with m/z 453, 679, 905); Nylon HPLC solvent filters can produce nylon (6,6monomer) peaks at masses of 226 Da a dimer 452, trimer 678 and tetramer 905 Da. The contaminant is very hard to get rid off since it binds very well to C18. | + | +226n | M |  | ü
| 259.8099 | CsI |  | Cesium iodide clusters, used as calibration |  |  | V
| 288.1371 | [CH3(CH2)11OSO3Na] |  | SDS (sodium dodecylsulfate) adducts (CH3(CH2)11OSO3Na) | - | +288n |  |  | ü
ESI Pos
Observed Masses | Mass diff. |  | PTFE
104.9919 | 82.0022 | C2HF3Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
186.9941 | 82.0027 | C4H2F6Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
268.9968 | 82.0028 | C6H3F9Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
350.9996 | 82.0016 | C8H4F12Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
433.0012 | 82.0001 | C10H5F15Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
515.0013 | 81.9856 | C12H6F18Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
596.9869 | 81.993 | C14H7F21Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
678.9799 | 81.9935 | C16H8F24Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
760.9734 | 82.0039 | C18H9F27Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
842.9773 | 81.9971 | C20H10F30Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
924.9744 |  | C22H11F33Na |  | Sodium adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules. | [5 empty] | ü
(M+H) | Exact mass |  | Phthalate
149.0233 |  | C8H4O3 | Pthalic Anhydride | fragment ion originating from phthalate esters | [5 empty] | ü
209 |  |  | Pthalic Anhydride | Phthalates Plastics.(Nalgene bottles, tubing) Rinse plasticware with isopropal alcohol. Minimize use of low-density polyethylene plasticware. Use Nalgene (high-density polyethylene or polypropylene) or teflon plasticware. | [5 empty] | ü
237 |  |  | Pthalic Anhydride | Plasticizers – Tygon tubing, solvents like methanol, and nitrogen generators. | [5 empty] | ü
255 |  |  | Pthalic Anhydride | [6 empty] | ü
195.0652 | 194.0574 | C6H4(COOCH3)2 | DMP | Dimethyl phthalate |  |  |  |  | ü | ü
223.0965 | 222.0887 | C6H4(COOC2H5)2 | DEP | Diethyl phthalate |  |  |  |  | ü
247.0965 | 246.0887 | C6H4(COOCH2CH=CH2)2 | DAP | Diallyl phthalate |  |  |  |  | ü
251.1278 | 250.12 | C6H4(COO(CH2)2CH3)2 | DPP | Di-n-propyl phthalate |  |  |  |  | ü
279.1591 | 278.1513 | C6H4(COO(CH2)3CH3)2 | DBP | Di-n-butyl phthalate |  |  |  |  | ü
279.1591 | 278.1513 | C6H4(COOCH2CH(CH3)2)2 | DIBP | Diisobutyl phthalate |  |  |  |  | ü
305.1747 | 304.1669 | CH3(CH2)3OOCC6H4COOC6H11 | BCP | Butyl cyclohexyl phthalate |  |  |  |  | ü
307.1904 | 306.1826 | C6H4(COO(CH2)4CH3)2 | DNPP | Di-n-pentyl phthalate |  |  |  |  | ü
331.1904 | 330.1826 | C6H4(COOC6H11)2 | DCP | Dicyclohexyl phthalate |  |  |  |  | ü
313.1434 | 312.1356 | CH3(CH2)3OOCC6H4COOCH2C6H5 | BBP | Butyl benzyl phthalate |  |  |  |  | ü
335.2217 | 334.2139 | C6H4(COO(CH2)5CH3)2 | DNHP | Di-n-hexyl phthalate |  |  |  |  | ü
335.2217 | 334.2139 | C6H4(COO(CH2)3CH(CH3)2)2 | DIHxP | Diisohexyl phthalate |  |  |  |  | ü
363.253 | 362.2452 | C6H4(COO(CH2)4CH(CH3)2)2 | DIHpP | Diisoheptyl phthalate |  |  |  |  | ü
363.253 | 362.2452 | CH3(CH2)3OOCC6H4COO(CH2)9CH3 | BDP | Butyl decyl phthalate |  |  |  |  | ü
391.2843 | 390.2765 | C6H4(COOCH2CH(C2H5)(CH2)3CH3)2 | DEHP, DOP | Di(2-ethylhexyl) phthalate | ubiquitous plasticiser bis(2-ethylhexyl) phthalate |  |  |  | ü | ü
391.2843 | 390.2765 | C6H4(COO(CH2)7CH3)2 | DNOP | Di(n-octyl) phthalate |  |  |  |  | ü
391.2843 | 390.2765 | C6H4(COO(CH2)5CH(CH3)2)2 | DIOP | Diisooctyl phthalate |  |  |  |  | ü
419.3156 | 418.3078 | CH3(CH2)7OOCC6H4COO(CH2)9CH3 | ODP | n-Octyl n-decyl phthalate |  |  |  |  | ü
419.3156 | 418.3078 | C6H4(COO(CH2)6CH(CH3)2)2 | DINP | Diisononyl phthalate |  |  |  |  | ü
447.3469 | 446.3391 | C6H4(COO(CH2)7CH(CH3)2)2 | DIDP | Diisodecyl phthalate |  |  |  |  | ü
475.3782 | 474.3704 | C6H4(COO(CH2)10CH3)2 | DUP | Diundecyl phthalate |  |  |  |  | ü
475.3782 | 474.3704 | C6H4(COO(CH2)8CH(CH3)2)2 | DIUP | Diisoundecyl phthalate |  |  |  |  | ü
531.4408 | 530.433 | C6H4(COO(CH2)12CH3)2 | DTDP | Ditridecyl phthalate |  |  |  |  | ü
531.4408 | 530.433 | C6H4(COO(CH2)10CH(CH3)2)2 | DIUP | Diisotridecyl phthalate |  |  |  |  | ü
Observed in ESI positive and APCI positive ion mode | Polysiloxane
sisweb.com | mass diff. | (Si(CH3)2O)n | [O-Si(CH3)2-]n | Polydimethylcyclosiloxanes, widespread compound in industrial products, e.g. deodorants. They are ubiquitously present in ambient air. | +17 Da ammonium adducts | -16 Da loss of methane
75.0261 |  | [(Si [CH3]2 O)1 + H]+ | [6 empty] | ü
149.0449 | 74.0188 | [(Si [CH3]2 O)2 + H]+
223.0637 | 74.0188 | [(Si [CH3]2 O)3 + H]+
297.0824 | 74.0188 | [(Si [CH3]2 O)4 + H]+
371.1012 | 74.0188 | [(Si [CH3]2 O)5 + H]+ |  | Polysiloxane, followed by m/z 388 | 388.1278 | 355.0699 |  | ü |  | ü | ü
445.12 | 74.0188 | [(Si [CH3]2 O)6 + H]+ |  | Polysiloxane, followed by m/z 462 | 462.1466 | 429.0887 |  | ü |  | ü | ü
519.1388 | 74.0188 | [(Si [CH3]2 O)7 + H]+ |  | Polysiloxane, followed by m/z 536 | 536.1654 | 503.1075 |  |  |  | ü
593.1576 | 74.0188 | [(Si [CH3]2 O)8 + H]+ |  | Polysiloxane, followed by m/z 610 | 610.1842 | 577.1263 |  |  |  | ü
667.1764 | 74.0188 | [(Si [CH3]2 O)9 + H]+ |  | Polysiloxane, followed by m/z 684 | 684.203 | 651.1451 |  |  |  | ü
741.1952 | 74.0188 | [(Si [CH3]2 O)10 + H]+ |  | Polysiloxane, followed by m/z 758 | 758.2218 | 725.1639 |  |  |  | ü
388.1278 |  | [(Si [CH3]2 O)5 + H + NH3]+ | +17 Da ammonium adducts
462.1466 | 74.0188 | [(Si [CH3]2 O)6 + H + NH3]+ | +17 Da ammonium adducts
536.1654 | 74.0188 | [(Si [CH3]2 O)7 + H + NH3]+ | +17 Da ammonium adducts
610.1842 | 74.0188 | [(Si [CH3]2 O)8 + H + NH3]+ | +17 Da ammonium adducts
684.203 | 74.0188 | [(Si [CH3]2 O)9 + H + NH3]+ | +17 Da ammonium adducts
758.2218 | 74.0188 | [(Si [CH3]2 O)10 + H + NH3]+ | +17 Da ammonium adducts
355.0699 |  | [(Si [CH3]2 O)5 + H]+ | -16 Da loss of methane
429.0887 | 74.0188 | [(Si [CH3]2 O)6 + H]+ | -16 Da loss of methane
503.1075 | 74.0188 | [(Si [CH3]2 O)7 + H]+ | -16 Da loss of methane
577.1263 | 74.0188 | [(Si [CH3]2 O)8 + H]+ | -16 Da loss of methane
651.1451 | 74.0188 | [(Si [CH3]2 O)9 + H]+ | -16 Da loss of methane
725.1639 | 74.0188 | [(Si [CH3]2 O)10 + H]+ | -16 Da loss of methane
+59 series
| +59 |  |  | Polymers of +59 might be iron in some form, presumably leaching from the steel in acid. We saw a lot of that before we eliminated as much metal as possible. We also saw +59 adducts onto larger, acidic peptides.
453.3 |  | Contamination |  | Nylon HPLC solvent filters can produce nylon (6,6monomer) peaks at masses of 226 Da a dimer 452, trimer 678 and tetramer 905 Da. The contaminant is very hard to get rid off since it binds very well to C18.
679.5 | 226.2 | Contamination |  | Nylon HPLC solvent filters can produce nylon (6,6monomer) peaks at masses of 226 Da a dimer 452, trimer 678 and tetramer 905 Da. The contaminant is very hard to get rid off since it binds very well to C18.
905.7 | 226.2 | Contamination |  | Nylon HPLC solvent filters can produce nylon (6,6monomer) peaks at masses of 226 Da a dimer 452, trimer 678 and tetramer 905 Da. The contaminant is very hard to get rid off since it binds very well to C18.
Peak Clusters at +21, -17, -35, -52
| +21 |  |  | +21 is sodium
| -17 |  |  | -17, -35, -52 are various losses of ammonia and water (if you look at the -17 ion you will probably see a small amount of the -18, as well).
| -35 |  |  | -17, -35, -52 are various losses of ammonia and water (if you look at the -17 ion you will probably see a small amount of the -18, as well).
| -52 |  |  | -17, -35, -52 are various losses of ammonia and water (if you look at the -17 ion you will probably see a small amount of the -18, as well).
Iron-Acetate clusters 538, 555 and 534
538 |  | Fe3O(CH3COO)6(L)x |  | Mobile phases with acetic acid can react with iron in the LC/MS system to produce Iron-acetate clusters (as established by Ijames, Dutky, and Fales: J Am Soc Mass Spec 6 1226 1995.) The general formula is Fe3O(CH3COO)6(L)x where L is a ligand from the mobile phase (water or organic component, acetonitrile solvent or water. 538 could be 6 acetates with no additional ligand. The 555 presumably has additional water.
555
534 |  |  |  | 533.88 could be a chromium cluster similar to the iron 537.88. Stainless steel in the system contains both iron and chromium. Potential sources in the ESI probe could be the ESI spray needle or the with the metal needle kit. This is not common try replacing the needle.
REPEATING UNITS, negative mode
Observed Masses | Mass difference | Formula for M or subunit or sequence | Compound ID or species | Origin |  |  | Reference(s)
| 14.0157 | -[CH2]-, |  | -[CH2]-, alkane chains, waxes, fatty acids, methylation |  |  | F
| 15.9949 | O, |  | O, oxidation |  |  | F
| 18.0106 | H2O, |  | H2O, water clusters |  |  | F
| 28.0313 | -[C2H4]-, |  | -[C2H4]-, natural alkane chains such as fatty acids |  |  | F
| 49.9968 | -[CF2]-, |  | -[CF2]-, from perfluoro compounds |  |  | F
| 67.9874 | NaHCO2, |  | NaHCO2, sodium formate clusters |  |  | F
| 74.0188 | -[O-Si(CH3)2]-, |  | -[O-Si(CH3)2]-, polysiloxane, silicone rubber polymer |  |  | F
| 82.0031 | NaCH3CO2, |  | NaCH3CO2, sodium acetate clusters |  |  | F
| 113.9929 | CF3COOH, |  | CF3COOH, trifluoroacetic acid, TFA adducts/clusters |  |  | F
| 121.9383 | NaClO4, |  | NaClO4, sodium perchlorate adducts/clusters |  |  | F, E
| 135.9748 | NaCF3CO2, |  | NaCF3CO2, sodium trifluoroacetate clusters |  |  | F
| 288.1371 | C12H25SO4Na, |  | C12H25SO4Na, sodium dodecylsulfate (SDS), adducts/clusters |  |  | F
ESI Neg |  | Acetate adducts of a fluorinated polymer (PTFE). From lithium grease used on 1100 LC pump modules.
141.0197 | 82.0081 | C4H5O2F3
223.0278 | 82.0067 | C6H6O2F6
305.0345 | 82.0076 | C8H7O2F9
387.0421 | 82.0009 | C10H8O2F12
469.043 | 81.991 | C12H9O2F15
551.034 |  | C14H10O2F18

## Adducts
Adducts, losses and replacements
Mass difference ± |  | Origin | Reference(s) | Waters | NewObjective | Leiden Univ.
|  |  |  | http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf | http://www.newobjective.com/downloads/technotes/PV-3.pdf | http://www.lc-ms.nl/
0.98402 |  | OH ↔ NH2, e.g. de-amidiation, CHNO compounds | F |  |  | ü
1 | (M+1) | (M+H)+ |  | ü
1.9957 |  | F ↔ OH, halogen exchange with hydroxy group (typically -F + OH) | F |  |  | ü
2.0156 |  | ± 2H, opening or forming of double bond | F |  |  | ü
4.9554 |  | Na+↔ NH4+, salt adduct | F |  |  | ü
7.0047 |  | F ↔ CN, halogen exchange with cyano group | F |  |  | ü
8.9658 |  | Cl ↔ CN, halogen exchange with cyano group | F |  |  | ü
13.9793 |  | O ↔ 2H, e.g. Oxidation follwed by H2O elimination | F |  |  | ü
14.0157 |  | ± CH2, alkane chains, waxes, fatty acids, methylation | F |  |  | ü
14.0157 |  | -[C3H6ON] ↔ -[C2H4ON], acrylamide versus iodoacetamide in cysteine alkylation (gels) | T |  |  | ü
15.9739 |  | Na+↔ K+, salt adduct | F |  |  | ü
15.9772 |  | S ↔ O, sulfur compounds | F |  |  | ü
15.9949 |  | ± O, e.g. oxidation/reduction | F |  |  | ü
17.0265 |  | ± NH3, ammonium adduct/neutral ammonium loss | F |  |  | ü
17.0265 |  | NH4+↔ H+, salt adduct | F |  |  | ü
17.9661 |  | Cl ↔ OH, halogen exchange with hydroxy group (typically -Cl + OH) | F |  |  | ü
17.9906 |  | F ↔ H, halogen exchange | F |  |  | ü
18 | (M+18) | (M+NH4)+ |  | ü | ü
18.0106 |  | ± H2O, water addition/loss | F |  |  | ü
18.9946 |  | CN ↔ COOH, nitrile compounds | F |  |  | ü
20.9293 |  | K+↔ NH4+, salt adduct | F |  |  | ü
21.9819 |  | Na+↔ H+, salt adduct | F |  |  | ü
23 | (M+23) | (M+Na)+ |  | ü | ü
24.9952 |  | CN ↔ H, nitrile compounds | F |  |  | ü
27.0109 |  | ± HCN, nitrile compounds | F |  |  | ü
27.9949 |  | ± CO | Z |  |  | ü
28.0061 |  | - 2N, nitrogen loss, e.g. azido compounds (N2) | Z |  |  | ü
28.0313 |  | ± C2H4, natural alkane chains such as fatty acids | F |  |  | ü
29.9742 |  | NO2 ↔ NH2, nitro compounds | F |  |  | ü
29.998 |  | -NO, nitroso compounds | F |  |  | ü
31.9721 |  | ± S, sulfur compounds | F |  |  | ü
31.9898 |  | ± 2O, oxygen loss | F |  |  | ü
33 | (M+33) | (M+CH3OH+H)+ |  | ü
33.0215 |  | - NH2OH, loss from hydroxamic acids | F |  |  | ü
33.961 |  | Cl ↔ H, halogen exchange | F |  |  | ü
33.9877 |  | ± H2S, sulfur compounds | F |  |  | ü
37.9559 |  | K+↔ H+, salt adduct | F |  |  | ü
37.9892 |  | 2CN ↔ 2COOH, nitrile compounds | F |  |  | ü
39 | (M+39) | (M+K)+ |  | ü | ü
40.0313 |  | +(C3H6O - H2O), acetone condensation after dehydration | F |  |  | ü
42 | (M+42) | (M+ACN+H)+ |  | ü
42.0106 |  | ± COCH2 | F |  |  | ü
42.047 |  | ± C3H6, propylation | F |  |  | ü
43.0058 |  | ± CONH (wrong calc. in ref.) | F |  |  | ü
43.9495 |  | 79Br ↔ Cl, halogen exchange (typically -Br +Cl) | F |  |  | ü
43.9898 |  | ± CO2 | F |  |  | ü
45.9474 |  | 81Br ↔ Cl, halogen exchange (typically -Br +Cl) | F |  |  | ü
47.967 |  | ± SO, sulfur compounds | F |  |  | ü
52.9153 |  | 79Br ↔ CN, halogen exchange with cyano group | F |  |  | ü
54.9132 |  | 81Br ↔ CN, halogen exchange with cyano group | F |  |  | ü
56.0626 |  | ± C4H8, butylation | F |  |  | ü
58.0055 |  | ± CO2CH2 | F |  |  | ü
58.0419 |  | +C3H6O, acetone condensation | F |  |  | ü
61.9156 |  | 79Br ↔ OH, halogen exchange with hydroxy group (typically -Br + OH) | F |  |  | ü
63.9136 |  | 81Br ↔ OH, halogen exchange with hydroxy group (typically -Br + OH) | F |  |  | ü
63.9619 |  | ± SO2, sulfur compounds | F |  |  | ü
63.9983 |  | -CH3SOH, specific loss from singly oxidized methionine residues | T |  |  | ü
64 | (M+64) | (M+ACN+Na)+ |  | ü
77.9105 |  | 79Br ↔ H, halogen exchange | F |  |  | ü
79 | (M+79) | (M+DMSO+H)+ |  | ü
79.9085 |  | 81Br ↔ H, halogen exchange | F |  |  | ü
79.9568 |  | ± SO3, sulfur compounds | F |  |  | ü
83 | (M+83) | (M+2ACN+H)+ |  | ü
91.9356 |  | I ↔ Cl, halogen exchange (typically -I +Cl) | F |  |  | ü
97.9674 |  | ± H2SO4, sulfur compounds | F |  |  | ü
97.9769 |  | ± H3PO4, phosphorous compounds | F |  |  | ü
100.9014 |  | I ↔ CN, halogen exchange with cyano group | F |  |  | ü
109.9017 |  | I ↔ OH, halogen exchange with hydroxy group (typically -I + OH) | F |  |  | ü
125.8966 |  | I ↔ H, halogen exchange | F |  |  | ü
146.0579 |  | ± [Deoxy-Hexose-H2O, C6O4H10], e.g. Fucose | S |  |  | ü
162.0528 |  | ± [Hexose-H2O, C6O5H10], e.g. Glucose, Galactose, Mannose, Fructose | F |  |  | ü
164.0685 |  | ± [Deoxy-Hexose-H2O, C6O5H12], e.g. Fucose | S |  |  | ü
176.0321 |  | ± [Glucuronic acid-H2O, C6O6H8] | F |  |  | ü
180.0634 |  | ± [Hexose, C6O6H12], e.g. Glucose, Galactose, Mannose, Fructose | F |  |  | ü
194.0427 |  | ± [Glucuronic acid, C6O7H10] | F |  |  | ü
203.0794 |  | ± [HexNAc-H2O, C8O5NH13], N-acetylhexoseamine | S |  |  | ü
206.0579 |  | + sinapic acid - H2O, (C11H12O5-H2O), MALDI-matrix for proteins, possible photo- or condensation adduct | R |  |  | ü
221.0899 |  | ± [HexNAc, C8O6NH15], N-acetylhexoseamine | S |  |  | ü
289.0732 |  | ± [Glutathione-H2O, C10O5N3SH15] | F |  |  | ü
291.0954 |  | ± [Neu5Ac-H2O, C11O8NH17], N-acetylneuraminic acid, sialic acid, NANA | S |  |  | ü
305.0682 |  | ± [Glutathione+O-H2O, C10O6N3SH15] | F |  |  | ü
307.0838 |  | ± [Glutathione, C10O6N3SH17] | F |  |  | ü
309.106 |  | ± [Neu5Ac, C11O9NH19], N-acetylneuraminic acid, sialic acid, NANA | S |  |  | ü
324.1056 |  | ± [Sucrose-H2O, C12O10H20] | F |  |  | ü
342.1162 |  | ± [Sucrose, C12O11H22] | F |  |  | ü
484.171 |  | +Tetraphenyl-tetramethyl-trisiloxane, (C28H32O2Si3), ( from silicon-based diffusion pump oil (DC7040, Dow corning) | P |  |  | ü
546.1867 |  | +Pentaphenyl-trimethyl-trisiloxane, (C33H34O2Si3), ( from silicon-based diffusion pump oil (DC7050, Dow corning) | P |  |  | ü

## Solvents
Solvents | Abbreviation | Formula | M.W. | M.W. deuterated
Acetic acid | - | C2H4O2 | 60 | 64
Acetone | - | C3H6O | 58 | 64
Acetonitrile | ACN | C2H3N | 41 | 44
Benzene | - | C6H6 | 78 | 84
Butanone (2-) | - | C4H8O | 72.1 | 80.1
Carbondisulfide | - | CS2 | 75.9 | 75.9
Chloroform | - | CHCl3 | 117.9 | 118.9
Dibutylphthalate (Phthalic acid dibutyl ester) | - | C16H22O4 | 278.2 | -
Dichloromethane | DCM | CH2Cl2 | 83.9 | 86
Diethylenglycoldimethylether (Diglyme) | - | C6H14O3 | 134.1 | 148.2
Diethylether | - | C4H10O | 74.1 | 84.1
Diisopropylether | - | C6H14O | 102.1 | 116.2
Dimethoxyethane (1,2 -) = Glykoldimethylether | DME | C4H10O2 | 90.1 | 100.1
Dimethylformamide | DMF | C3H7NO | 73.1 | 80.1
Dimethylsulfoxide | DMSO | C2H6OS | 78 | 84.1
Dioxan (1,4-) | - | C4H8O2 | 88.1 | 96.1
Ethanol | EtOH | C2H6O | 46 | 52.1
Ethylacetate (Acetic acid ethyl ester) | EtAc | C4H8O2 | 88.1 | 96.1
Formic acid | - | CH2O2 | 46 | 48
Glycerol | - | C3H8O3 | 92 | -
Hexamethyl-phosphoric acid triamide | HMPA | C6H18N3OP | 179.1 | 197.2
Hexane | - | C6H14 | 86.1 | 100.2
Isopropanol (2-propanol) | - | C3H8O | 60.1 | 68.1
Jonol (2,6-Di-(tert-butyl)-4-methylphenol) stab. | - | C15H24O | 220.2 | -
Methanol | MeOH | CH4O | 32 | 36
Propanol | - | C3H8O | 60.1 | 68.1
Pyridine | - | C5H5N | 79 | 84.1
Sulfur | - | S8 | 255.8 | 255.8
Tetrachlorocarbon | - | CCl4 | 151.9 | 151.9
Tetrachloroethylene | - | C2Cl4 | 163.9 | 163.9
Tetrahydrofuran | THF | C4H8O | 72.1 | 80.1
Tetramethylsilane | TMS | C4H12Si | 88.1 | -
Toluene | - | C7H8 | 92.1 | 100.1
Trichloroethylene | - | C2HCl3 | 129.9 | 130.9
Trifluoroacetic acid | TFA | C2F3O2H | 114 | 115
Triethylamine | TEA | C6H15N | 101.1 | -
Tripropylamine | TPA | C9H21N | 143.2 | -
Xylene | - | C8H10 | 106.1 | 116.1

## Carbs
All info on this sheet are from ionsource:
ionsource.com | http://ionsource.com/
Monosaccharide | Abbreviation | Generic | Composition | Monosaccharide |  | Monosaccharide Residue Mass (-H20)
Mass
- | - | - | - | Monoisotopic | Average | Monoisotopic | Average
- | - | - | - | - | - | - | -
Galactose | Gal | Hex | C6O6H12 | 180.0634 | 180.2 | 162.0528 | 162.1 | ü
Mannose | Man | Hex | C6O6H12 | 180.0634 | 180.2 | 162.0528 | 162.1 | ü
Fucose | Fuc | dHex | C6O5H12 | 164.0685 | 164.2 | 146.0579 | 146.1 | ü
Sialic Acid | NANA | - | C11O9NH19 | 309.106 | 309.3 | 291.0954 | 291.3 | ü
N-Acetlygalactoseamine | GalNAc | HexNAc | C8O6NH15 | 221.0899 | 221.2 | 203.0794 | 203.2 | ü | Fuc |  | Gal |  | Man |  |  | NANA | GalNAc | GlcNAc
N-Acetlyglucoseamine | GlcNAc | HexNAc | C8O6NH15 | 221.0899 | 221.2 | 203.0794 | 203.2 | ü
Sugar | Composition | 1+ | 2+ | 3+ | 4+ | 5+
dHexose | C6O4H10 | 146.1 | 73 | 48.7 | 36.5 | 29.2
Hexose | C6O5H10 | 162.1 | 81 | 54 | 40.4 | 32.4
HexNAc | C8O5NH13 | 203.2 | 101.6 | 67.7 | 50.8 | 40.6
NANA | C11O8NH17 | 291.3 | 145.6 | 97.1 | 72.8 | 58.2
HexNac+Hex | C14O10NH23 | 365.3 | 182.6 | 121.8 | 91.3 | 73.1
Monosaccharide Mass and Low Mass MS/MS (CID) Carbohydrate Marker Ions
Monosaccharide / | Composition | Monoisotopic mass |  | Low Mass CID Marker Ions
Disaccharide |  | Intact | Residue
Hexose, Hex | C6O6H12 | 180.0634 | 162.0528 | 163
(Galactose, Mannose)
Deoxy-Hexose, dHex | C6O5H12 | 164.0685 | 146.0579 | 147
(Fucose)
N-acetylhexoseamine, HexNAc | C8O6NH15 | 221.0899 | 203.0794 | 204
(N-acetylglucoseamine,
N-acetylgalactoseamine)
N-acetylneuraminic | C11O9NH19 | 309.106 | 291.0954 | 292/274/256
acid, Neu5Ac, NANA
(sialic acid)
Hex-HexNAc | C14O11NH25 | 383.1428 | 365.1322 | 366
Hex3HexNac2 | C46O35N2H76 | 910.3278 | 892.3172 | -
(Common N-linked Core)
Reference: (1) Conboy, J. J. and Henion, J. D., The determination of glycopeptides by liquid chromatography/mass spectrometry with collision-induced dissociation, JASMS, 3, 804,1992. (2) Huddleston, M. J., Bean, M. F., Carr, S. A., Collisional fragmentation of glycopeptides by electrospray ionization LC/MS and LC/MS/MS. Methods for selective detection of glycopeptides in protein digests, Anal. Chem., 65, 877, 1993.
di-
| Enter the | [10 empty] | Elemental Composition
| # of residues | mono | average |  | C | O | N | H |  | mono | average | C | O | N | H
Sialic Acid | 2 | 582.1908 | 582.52 |  | 22 | 16 | 2 | 34 | C11O8NH17 | 291.0954 | 291.26 | 11 | 8 | 1 | 17
Galactose | 2 | 324.1056 | 324.28 |  | 12 | 10 | 0 | 20 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
N-Acetylglucosamine | 4 | 812.3176 | 812.76 |  | 32 | 20 | 4 | 52 | C8O5NH13 | 203.0794 | 203.19 | 8 | 5 | 1 | 13
Mannose | 3 | 486.1584 | 486.42 |  | 18 | 15 | 0 | 30 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
Fucose | 1 | 146.0579 | 146.14 |  | 6 | 4 | 0 | 10 | C6O4H10 | 146.0579 | 146.14 | 6 | 4 | 0 | 10
mass |  | 2350.8303 | 2352.12 | Glycan -> | 90 | 65 | 6 | 146
Elemental
Composition
Glyco Peptide Mass Calculation
| Peptide Mass |  | Resulting GlycoPeptide Mass
| mono | average |  | mono | average
Enter Peptide Mass | 0 | 0 |  | 2350.8303 | 2352.12
tri-
| Enter the | [10 empty] | Elemental Composition
| # of residues | mono | average |  | C | O | N | H |  | mono | average | C | O | N | H
Sialic Acid | 3 | 873.2862 | 873.78 |  | 33 | 24 | 3 | 51 | C11O8NH17 | 291.0954 | 291.26 | 11 | 8 | 1 | 17
Galactose | 3 | 486.1584 | 486.42 |  | 18 | 15 | 0 | 30 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
N-Acetylglucosamine | 5 | 1015.397 | 1015.95 |  | 40 | 25 | 5 | 65 | C8O5NH13 | 203.0794 | 203.19 | 8 | 5 | 1 | 13
Mannose | 3 | 486.1584 | 486.42 |  | 18 | 15 | 0 | 30 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
Fucose | 1 | 146.0579 | 146.14 |  | 6 | 4 | 0 | 10 | C6O4H10 | 146.0579 | 146.14 | 6 | 4 | 0 | 10
mass |  | 3007.0579 | 3008.71 | Glycan -> | 115 | 83 | 8 | 186
Elemental
Composition
Glyco Peptide Mass Calculation
| Peptide Mass |  | Resulting GlycoPeptide Mass
| mono | average |  | mono | average
Enter Peptide Mass | 0 | 0 |  | 3007.0579 | 3008.71
tetra-
| Enter the | [10 empty] | Elemental Composition
| # of residues | mono | average |  | C | O | N | H |  | mono | average | C | O | N | H
Sialic Acid | 4 | 1164.3816 | 1165.04 |  | 44 | 32 | 4 | 68 | C11O8NH17 | 291.0954 | 291.26 | 11 | 8 | 1 | 17
Galactose | 4 | 648.2112 | 648.56 |  | 24 | 20 | 0 | 40 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
N-Acetylglucosamine | 6 | 1218.4764 | 1219.14 |  | 48 | 30 | 6 | 78 | C8O5NH13 | 203.0794 | 203.19 | 8 | 5 | 1 | 13
Mannose | 3 | 486.1584 | 486.42 |  | 18 | 15 | 0 | 30 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
Fucose | 1 | 146.0579 | 146.14 |  | 6 | 4 | 0 | 10 | C6O4H10 | 146.0579 | 146.14 | 6 | 4 | 0 | 10
mass |  | 3663.2855 | 3665.3 | Glycan -> | 140 | 101 | 10 | 226
Elemental
Composition
Glyco Peptide Mass Calculation
| Peptide Mass |  | Resulting GlycoPeptide Mass
| mono | average |  | mono | average
Enter Peptide Mass | 0 | 0 |  | 3663.2855 | 3665.3
man7
| Enter the | [10 empty] | Elemental Composition
| # of residues | mono | average |  | C | O | N | H |  | mono | average | C | O | N | H
Sialic Acid | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | C11O8NH17 | 291.0954 | 291.26 | 11 | 8 | 1 | 17
Galactose | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
N-Acetylglucosamine | 2 | 406.1588 | 406.38 |  | 16 | 10 | 2 | 26 | C8O5NH13 | 203.0794 | 203.19 | 8 | 5 | 1 | 13
Mannose | 7 | 1134.3696 | 1134.98 |  | 42 | 35 | 0 | 70 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
Fucose | 1 | 146.0579 | 146.14 |  | 6 | 4 | 0 | 10 | C6O4H10 | 146.0579 | 146.14 | 6 | 4 | 0 | 10
mass |  | 1686.5863 | 1687.5 | Glycan -> | 64 | 49 | 2 | 106
Elemental
Composition
Glyco Peptide Mass Calculation
| Peptide Mass |  | Resulting GlycoPeptide Mass
| mono | average |  | mono | average
Enter Peptide Mass | 0 | 0 |  | 1686.5863 | 1687.5
hybrid-
| Enter the | [10 empty] | Elemental Composition
| # of residues | mono | average |  | C | O | N | H |  | mono | average | C | O | N | H
Sialic Acid | 1 | 291.0954 | 291.26 |  | 11 | 8 | 1 | 17 | C11O8NH17 | 291.0954 | 291.26 | 11 | 8 | 1 | 17
Galactose | 1 | 162.0528 | 162.14 |  | 6 | 5 | 0 | 10 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
N-Acetylglucosamine | 3 | 609.2382 | 609.57 |  | 24 | 15 | 3 | 39 | C8O5NH13 | 203.0794 | 203.19 | 8 | 5 | 1 | 13
Mannose | 6 | 972.3168 | 972.84 |  | 36 | 30 | 0 | 60 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
Fucose | 1 | 146.0579 | 146.14 |  | 6 | 4 | 0 | 10 | C6O4H10 | 146.0579 | 146.14 | 6 | 4 | 0 | 10
mass |  | 2180.7611 | 2181.95 | Glycan -> | 83 | 62 | 4 | 136
Elemental
Composition
Glyco Peptide Mass Calculation
| Peptide Mass |  | Resulting GlycoPeptide Mass
| mono | average |  | mono | average
Enter Peptide Mass | 0 | 0 |  | 2180.7611 | 2181.95
Core
| Enter the | [10 empty] | Elemental Composition
| # of residues | mono | average |  | C | O | N | H |  | mono | average | C | O | N | H
Sialic Acid | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | C11O8NH17 | 291.0954 | 291.26 | 11 | 8 | 1 | 17
Galactose | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
N-Acetylglucosamine | 2 | 406.1588 | 406.38 |  | 16 | 10 | 2 | 26 | C8O5NH13 | 203.0794 | 203.19 | 8 | 5 | 1 | 13
Mannose | 3 | 486.1584 | 486.42 |  | 18 | 15 | 0 | 30 | C6O5H10 | 162.0528 | 162.14 | 6 | 5 | 0 | 10
Fucose | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | C6O4H10 | 146.0579 | 146.14 | 6 | 4 | 0 | 10
mass |  | 892.3172 | 892.8 | Glycan -> | 34 | 25 | 2 | 56
Elemental
Composition
Glyco Peptide Mass Calculation
| Peptide Mass |  | Resulting GlycoPeptide Mass
| mono | average |  | mono | average
Enter Peptide Mass | 0 | 0 |  | 892.3172 | 892.8
N-linked:
All N-linked carbohydrates are linked through
N-Acetylglucosamine and the amino acid asparagine
as shown in Figure 1. The N-linked amino acid
consensus sequence is Asn-any AA- Ser or Thr.
The middle amino acid can not be proline (Pro).
"O-Linked "
-
Most O-linked carbohydrate covalent attachments
to proteins involve a linkage between the
monosaccharide N- Acetylgalactosamine and the
amino acids serine or threonine. Currently there is
not an O-linked amino acid consensus sequence.

## Masses
Element | Abbr. | atomic masses *) | # of atoms | Abbr. |  | element | mono mass |  | 12 | 1.0078 | 14.0031 | 15.9949 | 31.9721 | 30.9738
[6 empty] | element | avg mass |  | 12.0107 | 1.0079 | 14.0067 | 15.9994 | 32.065 | 30.9738
Carbon-12 | C-12 | 12 | 1 | C-12
Carbon-13 | C-13 | 13.0034 | 0 | C-13
Hydrogen | H | 1.0078 | 4 | H |  |  | amino acid | code | c | h | n | o | s | p | calculated avg mass | calculated mono mass
Deuterium | D | 2.0141 | 0 | D |  |  | alanine | A | 3 | 5 | 1 | 1 | 0 |  | 71.0779 | 71.0371
Oxygen | O | 15.9949 | 0 | O |  |  | arginine | R | 6 | 12 | 4 | 1 | 0 |  | 156.1857 | 156.1011
Oxygen-18 | O-18 | 17.9992 | 0 | O-18 |  |  | asparagine | N | 4 | 6 | 2 | 2 | 0 |  | 114.1026 | 114.0429
Nitrogen-14 | N-14 | 14.0031 | 0 | N-14 |  |  | aspartic acid | D | 4 | 5 | 1 | 3 | 0 |  | 115.0874 | 115.0269
Nitrogen-15 | N-15 | 15.0001 | 0 | N-15 |  |  | cysteine | C | 3 | 5 | 1 | 1 | 1 |  | 103.1429 | 103.0092
Sulfur-32 | S-32 | 31.9721 | 0 | S-32 |  |  | glutamic acid | E | 5 | 7 | 1 | 3 | 0 |  | 129.114 | 129.0426
Sulfur-34 | S-34 | 33.9679 | 0 | S-34 |  |  | glutamine | Q | 5 | 8 | 2 | 2 | 0 |  | 128.1292 | 128.0586
Phosphorus | P | 30.9738 | 0 | P |  |  | glycine | G | 2 | 3 | 1 | 1 | 0 |  | 57.0513 | 57.0215
Silicium | Si | 27.9769 | 0 | Si |  |  | histidine | H | 6 | 7 | 3 | 1 | 0 |  | 137.1393 | 137.0589
Sodium | Na | 22.9898 | 0 | Na |  |  | isoleucine | I | 6 | 11 | 1 | 1 | 0 |  | 113.1576 | 113.0841
Potassium | K | 38.9637 | 0 | K |  |  | leucine | L | 6 | 11 | 1 | 1 | 0 |  | 113.1576 | 113.0841
Chlorine-35 | Cl-35 | 34.9689 | 0 | Cl-35 |  |  | lysing | K | 6 | 12 | 2 | 1 | 0 |  | 128.1723 | 128.095
Chlorine-37 | Cl-37 | 36.9659 | 0 | Cl-37 |  |  | methionine | M | 5 | 9 | 1 | 1 | 1 |  | 131.1961 | 131.0405
Bromine-79 | Br-79 | 78.9183 | 0 | Br-79 |  |  | phenylalanine | F | 9 | 9 | 1 | 1 | 0 |  | 147.1739 | 147.0684
Bromine 81 | Br-81 | 80.9163 | 0 | Br-81 |  |  | proline | P | 5 | 7 | 1 | 1 | 0 |  | 97.1152 | 97.0528
Iodine | Iodine | 126.9045 | 0 | Iodine |  |  | serine | S | 3 | 5 | 1 | 2 | 0 |  | 87.0773 | 87.032
Fluorine | F | 18.9984 | 0 | F |  |  | threonine | T | 4 | 7 | 1 | 2 | 0 |  | 101.1039 | 101.0477
Iron-54 | Fe-54 | 53.9396 | 0 | Fe-54 |  |  | tryptophan | W | 11 | 10 | 2 | 1 | 0 |  | 186.2099 | 186.0793
Iron-56 | Fe-56 | 55.9349 | 0 | Fe-56 |  |  | tyrosine | Y | 9 | 9 | 1 | 2 | 0 |  | 163.1733 | 163.0633
Lithium-6 | Li-6 | 6.0151 | 0 | Li-6 |  |  | valine | V | 5 | 9 | 1 | 1 | 0 |  | 99.1311 | 99.0684
Lithium-7 | Li-7 | 7.016 | 0 | Li-7 |  |  | ornithine | O | 5 | 12 | 2 | 2 | 0 |  | 132.161 | 132.0899
Boron-10 | B-10 | 10.0129 | 0 | B-10
Boron-11 | B-11 | 11.0093 | 0 | B-11
Copper-63 | Cu-63 | 62.9296 | 0 | Cu-63 |  | Reference: http://www.unimod.org/masses.html
Copper-65 | Cu-65 | 64.9278 | 0 | Cu-65 |  | 2009-04-16 Jimmy Eng, UW Proteomics Resource
Silver-107 | Ag-107 | 106.9051 | 0 | Ag-107
Silver-109 | Ag-109 | 108.9048 | 0 | Ag-109
Tin-120 | Sn-120 | 119.9022 | 0 | Sn-120
Cesium | Cs-133 | 132.9055 | 0 | Cs-133
Proton | H+ | 1.0073 | 0 | H+
Electron | e- | 0.000549 | 0 | e-
monoisotopic mass | (exact) | 16.0313 | (theoretical)
Measured value |  | 16
Mass error in mDa |  | 31.3
Mass error in ppm |  | 1952.4306
*) atomic masses adopted from the National Institute of Standarts and Technology (NIST)
(http://physics.nist.gov/PhysRefData/Compositions/index.html)
http://www.nist.gov/pml/data/comp.cfm
Isotope |  |  |  | Relative | Isotopic | Standard | Notes
|  |  |  | Atomic Mass | Composition | Atomic Weight
1 | H | 1 |  | 1.007 825 032 07(10) | 0.999 885(70) | 1.007 94(7) | g,m,r,b,w
| D | 2 |  | 2.014 101 777 8(4) | 0.000 115(70)
| T | 3 |  | 3.016 049 2777(25)
2 | He | 3 |  | 3.016 029 3191(26) | 0.000 001 34(3) | 4.002 602(2) | g,r,a
|  | 4 |  | 4.002 603 254 15(6) | 0.999 998 66(3)
3 | Li | 6 |  | 6.015 122 795(16) | 0.0759(4) | 6.941(2) | g,m,r,c,i
|  | 7 |  | 7.016 004 55(8) | 0.9241(4)
4 | Be | 9 |  | 9.012 182 2(4) | 1.0000 | 9.012 182(3)
5 | B | 10 |  | 10.012 937 0(4) | 0.199(7) | 10.811(7) | g,m,r
|  | 11 |  | 11.009 305 4(4) | 0.801(7)
6 | C | 12 |  | 12.000 000 0(0) | 0.9893(8) | 12.0107(8) | g,r
|  | 13 |  | 13.003 354 8378(10) | 0.0107(8)
|  | 14 |  | 14.003 241 989(4)
7 | N | 14 |  | 14.003 074 004 8(6) | 0.996 36(20) | 14.0067(2) | g,r,a,d
|  | 15 |  | 15.000 108 898 2(7) | 0.003 64(20)
8 | O | 16 |  | 15.994 914 619 56(16) | 0.997 57(16) | 15.9994(3) | g,r,e,w
|  | 17 |  | 16.999 131 70(12) | 0.000 38(1)
|  | 18 |  | 17.999 161 0(7) | 0.002 05(14)
9 | F | 19 |  | 18.998 403 22(7) | 1.0000 | 18.998 403 2(5)
10 | Ne | 20 |  | 19.992 440 1754(19) | 0.9048(3) | 20.1797(6) | g,m,a
|  | 21 |  | 20.993 846 68(4) | 0.0027(1)
|  | 22 |  | 21.991 385 114(19) | 0.0925(3)
11 | Na | 23 |  | 22.989 769 2809(29) | 1.0000 | 22.989 769 28(2)
12 | Mg | 24 |  | 23.985 041 700(14) | 0.7899(4) | 24.3050(6)
|  | 25 |  | 24.985 836 92(3) | 0.1000(1)
|  | 26 |  | 25.982 592 929(30) | 0.1101(3)
13 | Al | 27 |  | 26.981 538 63(12) | 1.0000 | 26.981 538 6(8)
14 | Si | 28 |  | 27.976 926 5325(19) | 0.922 23(19) | 28.0855(3) | r
|  | 29 |  | 28.976 494 700(22) | 0.046 85(8)
|  | 30 |  | 29.973 770 17(3) | 0.030 92(11)
15 | P | 31 |  | 30.973 761 63(20) | 1.0000 | 30.973 762(2)
16 | S | 32 |  | 31.972 071 00(15) | 0.9499(26) | 32.065(5) | g,r
|  | 33 |  | 32.971 458 76(15) | 0.0075(2)
|  | 34 |  | 33.967 866 90(12) | 0.0425(24)
|  | 36 |  | 35.967 080 76(20) | 0.0001(1)
17 | Cl | 35 |  | 34.968 852 68(4) | 0.7576(10) | 35.453(2) | m
|  | 37 |  | 36.965 902 59(5) | 0.2424(10)
18 | Ar | 36 |  | 35.967 545 106(29) | 0.003 365(30) | 39.948(1) | g,r,a
|  | 38 |  | 37.962 732 4(4) | 0.000 632(5)
|  | 40 |  | 39.962 383 1225(29) | 0.996 003(30)
19 | K | 39 |  | 38.963 706 68(20) | 0.932 581(44) | 39.0983(1) | g
|  | 40 |  | 39.963 998 48(21) | 0.000 117(1)
|  | 41 |  | 40.961 825 76(21) | 0.067 302(44)
20 | Ca | 40 |  | 39.962 590 98(22) | 0.969 41(156) | 40.078(4) | g,f
|  | 42 |  | 41.958 618 01(27) | 0.006 47(23)
|  | 43 |  | 42.958 766 6(3) | 0.001 35(10)
|  | 44 |  | 43.955 481 8(4) | 0.020 86(110)
|  | 46 |  | 45.953 6926(24) | 0.000 04(3)
|  | 48 |  | 47.952 534(4) | 0.001 87(21)
21 | Sc | 45 |  | 44.955 911 9(9) | 1.0000 | 44.955 912(6)
22 | Ti | 46 |  | 45.952 631 6(9) | 0.0825(3) | 47.867(1)
|  | 47 |  | 46.951 763 1(9) | 0.0744(2)
|  | 48 |  | 47.947 946 3(9) | 0.7372(3)
|  | 49 |  | 48.947 870 0(9) | 0.0541(2)
|  | 50 |  | 49.944 791 2(9) | 0.0518(2)
23 | V | 50 |  | 49.947 1585(11) | 0.002 50(4) | 50.9415(1)
|  | 51 |  | 50.943 9595(11) | 0.997 50(4)
24 | Cr | 50 |  | 49.946 0442(11) | 0.043 45(13) | 51.9961(6)
|  | 52 |  | 51.940 507 5(8) | 0.837 89(18)
|  | 53 |  | 52.940 649 4(8) | 0.095 01(17)
|  | 54 |  | 53.938 880 4(8) | 0.023 65(7)
25 | Mn | 55 |  | 54.938 045 1(7) | 1.0000 | 54.938 045(5)
26 | Fe | 54 |  | 53.939 610 5(7) | 0.058 45(35) | 55.845(2)
|  | 56 |  | 55.934 937 5(7) | 0.917 54(36)
|  | 57 |  | 56.935 394 0(7) | 0.021 19(10)
|  | 58 |  | 57.933 275 6(8) | 0.002 82(4)
27 | Co | 59 |  | 58.933 195 0(7) | 1.0000 | 58.933 195(5)
28 | Ni | 58 |  | 57.935 342 9(7) | 0.680 769(89) | 58.6934(4)
|  | 60 |  | 59.930 786 4(7) | 0.262 231(77)
|  | 61 |  | 60.931 056 0(7) | 0.011 399(6)
|  | 62 |  | 61.928 345 1(6) | 0.036 345(17)
|  | 64 |  | 63.927 966 0(7) | 0.009 256(9)
29 | Cu | 63 |  | 62.929 597 5(6) | 0.6915(15) | 63.546(3) | r
|  | 65 |  | 64.927 789 5(7) | 0.3085(15)
30 | Zn | 64 |  | 63.929 142 2(7) | 0.482 68(321) | 65.38(2)
|  | 66 |  | 65.926 0334(10) | 0.279 75(77)
|  | 67 |  | 66.927 1273(10) | 0.041 02(21)
|  | 68 |  | 67.924 8442(10) | 0.190 24(123)
|  | 70 |  | 69.925 3193(21) | 0.006 31(9)
31 | Ga | 69 |  | 68.925 5736(13) | 0.601 08(9) | 69.723(1)
|  | 71 |  | 70.924 7013(11) | 0.398 92(9)
32 | Ge | 70 |  | 69.924 2474(11) | 0.2038(18) | 72.64(1)
|  | 72 |  | 71.922 0758(18) | 0.2731(26)
|  | 73 |  | 72.923 4589(18) | 0.0776(8)
|  | 74 |  | 73.921 1778(18) | 0.3672(15)
|  | 76 |  | 75.921 4026(18) | 0.0783(7)
33 | As | 75 |  | 74.921 5965(20) | 1.0000 | 74.921 60(2)
34 | Se | 74 |  | 73.922 4764(18) | 0.0089(4) | 78.96(3)
|  | 76 |  | 75.919 2136(18) | 0.0937(29)
|  | 77 |  | 76.919 9140(18) | 0.0763(16)
|  | 78 |  | 77.917 3091(18) | 0.2377(28)
|  | 80 |  | 79.916 5213(21) | 0.4961(41)
|  | 82 |  | 81.916 6994(22) | 0.0873(22)
35 | Br | 79 |  | 78.918 3371(22) | 0.5069(7) | 79.904(1)
|  | 81 |  | 80.916 2906(21) | 0.4931(7)
36 | Kr | 78 |  | 77.920 3648(12) | 0.003 55(3) | 83.798(2) | g,m,a
|  | 80 |  | 79.916 3790(16) | 0.022 86(10)
|  | 82 |  | 81.913 4836(19) | 0.115 93(31)
|  | 83 |  | 82.914 136(3) | 0.115 00(19)
|  | 84 |  | 83.911 507(3) | 0.569 87(15)
|  | 86 |  | 85.910 610 73(11) | 0.172 79(41)
37 | Rb | 85 |  | 84.911 789 738(12) | 0.7217(2) | 85.4678(3) | g
|  | 87 |  | 86.909 180 527(13) | 0.2783(2)
38 | Sr | 84 |  | 83.913 425(3) | 0.0056(1) | 87.62(1) | g,r,f
|  | 86 |  | 85.909 2602(12) | 0.0986(1)
|  | 87 |  | 86.908 8771(12) | 0.0700(1)
|  | 88 |  | 87.905 6121(12) | 0.8258(1)
39 | Y | 89 |  | 88.905 8483(27) | 1.0000 | 88.905 85(2)
40 | Zr | 90 |  | 89.904 7044(25) | 0.5145(40) | 91.224(2) | g
|  | 91 |  | 90.905 6458(25) | 0.1122(5)
|  | 92 |  | 91.905 0408(25) | 0.1715(8)
|  | 94 |  | 93.906 3152(26) | 0.1738(28)
|  | 96 |  | 95.908 2734(30) | 0.0280(9)
41 | Nb | 93 |  | 92.906 3781(26) | 1.0000 | 92.906 38(2)
42 | Mo | 92 |  | 91.906 811(4) | 0.1477(31) | 95.96(2) | g
|  | 94 |  | 93.905 0883(21) | 0.0923(10)
|  | 95 |  | 94.905 8421(21) | 0.1590(9)
|  | 96 |  | 95.904 6795(21) | 0.1668(1)
|  | 97 |  | 96.906 0215(21) | 0.0956(5)
|  | 98 |  | 97.905 4082(21) | 0.2419(26)
|  | 100 |  | 99.907 477(6) | 0.0967(20)
43 | Tc | 97 |  | 96.906 365(5) |  | [98]
|  | 98 |  | 97.907 216(4)
|  | 99 |  | 98.906 2547(21)
44 | Ru | 96 |  | 95.907 598(8) | 0.0554(14) | 101.07(2) | g
|  | 98 |  | 97.905 287(7) | 0.0187(3)
|  | 99 |  | 98.905 9393(22) | 0.1276(14)
|  | 100 |  | 99.904 2195(22) | 0.1260(7)
|  | 101 |  | 100.905 5821(22) | 0.1706(2)
|  | 102 |  | 101.904 3493(22) | 0.3155(14)
|  | 104 |  | 103.905 433(3) | 0.1862(27)
45 | Rh | 103 |  | 102.905 504(3) | 1.0000 | 102.905 50(2)
46 | Pd | 102 |  | 101.905 609(3) | 0.0102(1) | 106.42(1) | g
|  | 104 |  | 103.904 036(4) | 0.1114(8)
|  | 105 |  | 104.905 085(4) | 0.2233(8)
|  | 106 |  | 105.903 486(4) | 0.2733(3)
|  | 108 |  | 107.903 892(4) | 0.2646(9)
|  | 110 |  | 109.905 153(12) | 0.1172(9)
47 | Ag | 107 |  | 106.905 097(5) | 0.518 39(8) | 107.8682(2) | g
|  | 109 |  | 108.904 752(3) | 0.481 61(8)
48 | Cd | 106 |  | 105.906 459(6) | 0.0125(6) | 112.411(8) | g
|  | 108 |  | 107.904 184(6) | 0.0089(3)
|  | 110 |  | 109.903 0021(29) | 0.1249(18)
|  | 111 |  | 110.904 1781(29) | 0.1280(12)
|  | 112 |  | 111.902 7578(29) | 0.2413(21)
|  | 113 |  | 112.904 4017(29) | 0.1222(12)
|  | 114 |  | 113.903 3585(29) | 0.2873(42)
|  | 116 |  | 115.904 756(3) | 0.0749(18)
49 | In | 113 |  | 112.904 058(3) | 0.0429(5) | 114.818(3)
|  | 115 |  | 114.903 878(5) | 0.9571(5)
50 | Sn | 112 |  | 111.904 818(5) | 0.0097(1) | 118.710(7) | g,e
|  | 114 |  | 113.902 779(3) | 0.0066(1)
|  | 115 |  | 114.903 342(3) | 0.0034(1)
|  | 116 |  | 115.901 741(3) | 0.1454(9)
|  | 117 |  | 116.902 952(3) | 0.0768(7)
|  | 118 |  | 117.901 603(3) | 0.2422(9)
|  | 119 |  | 118.903 308(3) | 0.0859(4)
|  | 120 |  | 119.902 1947(27) | 0.3258(9)
|  | 122 |  | 121.903 4390(29) | 0.0463(3)
|  | 124 |  | 123.905 2739(15) | 0.0579(5)
51 | Sb | 121 |  | 120.903 8157(24) | 0.5721(5) | 121.760(1) | g
|  | 123 |  | 122.904 2140(22) | 0.4279(5)
52 | Te | 120 |  | 119.904 020(10) | 0.0009(1) | 127.60(3) | g,h
|  | 122 |  | 121.903 0439(16) | 0.0255(12)
|  | 123 |  | 122.904 2700(16) | 0.0089(3)
|  | 124 |  | 123.902 8179(16) | 0.0474(14)
|  | 125 |  | 124.904 4307(16) | 0.0707(15)
|  | 126 |  | 125.903 3117(16) | 0.1884(25)
|  | 128 |  | 127.904 4631(19) | 0.3174(8)
|  | 130 |  | 129.906 2244(21) | 0.3408(62)
53 | I | 127 |  | 126.904 473(4) | 1.0000 | 126.904 47(3)
54 | Xe | 124 |  | 123.905 8930(20) | 0.000 952(3) | 131.293(6) | g,m,a
|  | 126 |  | 125.904 274(7) | 0.000 890(2)
|  | 128 |  | 127.903 5313(15) | 0.019 102(8)
|  | 129 |  | 128.904 779 4(8) | 0.264 006(82)
|  | 130 |  | 129.903 508 0(8) | 0.040 710(13)
|  | 131 |  | 130.905 0824(10) | 0.212 324(30)
|  | 132 |  | 131.904 1535(10) | 0.269 086(33)
|  | 134 |  | 133.905 394 5(9) | 0.104 357(21)
|  | 136 |  | 135.907 219(8) | 0.088 573(44)
55 | Cs | 133 |  | 132.905 451 933(24) | 1.0000 | 132.905 451 9(2)
56 | Ba | 130 |  | 129.906 3208(30) | 0.001 06(1) | 137.327(7)
|  | 132 |  | 131.905 0613(11) | 0.001 01(1)
|  | 134 |  | 133.904 508 4(4) | 0.024 17(18)
|  | 135 |  | 134.905 688 6(4) | 0.065 92(12)
|  | 136 |  | 135.904 575 9(4) | 0.078 54(24)
|  | 137 |  | 136.905 827 4(5) | 0.112 32(24)
|  | 138 |  | 137.905 247 2(5) | 0.716 98(42)
57 | La | 138 |  | 137.907 112(4) | 0.000 90(1) | 138.905 47(7) | g
|  | 139 |  | 138.906 3533(26) | 0.999 10(1)
58 | Ce | 136 |  | 135.907 172(14) | 0.001 85(2) | 140.116(1) | g,f
|  | 138 |  | 137.905 991(11) | 0.002 51(2)
|  | 140 |  | 139.905 4387(26) | 0.884 50(51)
|  | 142 |  | 141.909 244(3) | 0.111 14(51)
59 | Pr | 141 |  | 140.907 6528(26) | 1.0000 | 140.907 65(2)
60 | Nd | 142 |  | 141.907 7233(25) | 0.272(5) | 144.242(3) | g,f
|  | 143 |  | 142.909 8143(25) | 0.122(2)
|  | 144 |  | 143.910 0873(25) | 0.238(3)
|  | 145 |  | 144.912 5736(25) | 0.083(1)
|  | 146 |  | 145.913 1169(25) | 0.172(3)
|  | 148 |  | 147.916 893(3) | 0.057(1)
|  | 150 |  | 149.920 891(3) | 0.056(2)
61 | Pm | 145 |  | 144.912 749(3) |  | [145]
|  | 147 |  | 146.915 1385(26)
62 | Sm | 144 |  | 143.911 999(3) | 0.0307(7) | 150.36(2) | g
|  | 147 |  | 146.914 8979(26) | 0.1499(18)
|  | 148 |  | 147.914 8227(26) | 0.1124(10)
|  | 149 |  | 148.917 1847(26) | 0.1382(7)
|  | 150 |  | 149.917 2755(26) | 0.0738(1)
|  | 152 |  | 151.919 7324(27) | 0.2675(16)
|  | 154 |  | 153.922 2093(27) | 0.2275(29)
63 | Eu | 151 |  | 150.919 8502(26) | 0.4781(6) | 151.964(1) | g
|  | 153 |  | 152.921 2303(26) | 0.5219(6)
64 | Gd | 152 |  | 151.919 7910(27) | 0.0020(1) | 157.25(3) | g
|  | 154 |  | 153.920 8656(27) | 0.0218(3)
|  | 155 |  | 154.922 6220(27) | 0.1480(12)
|  | 156 |  | 155.922 1227(27) | 0.2047(9)
|  | 157 |  | 156.923 9601(27) | 0.1565(2)
|  | 158 |  | 157.924 1039(27) | 0.2484(7)
|  | 160 |  | 159.927 0541(27) | 0.2186(19)
65 | Tb | 159 |  | 158.925 3468(27) | 1.0000 | 158.925 35(2)
66 | Dy | 156 |  | 155.924 283(7) | 0.000 56(3) | 162.500(1) | g
|  | 158 |  | 157.924 409(4) | 0.000 95(3)
|  | 160 |  | 159.925 1975(27) | 0.023 29(18)
|  | 161 |  | 160.926 9334(27) | 0.188 89(42)
|  | 162 |  | 161.926 7984(27) | 0.254 75(36)
|  | 163 |  | 162.928 7312(27) | 0.248 96(42)
|  | 164 |  | 163.929 1748(27) | 0.282 60(54)
67 | Ho | 165 |  | 164.930 3221(27) | 1.0000 | 164.930 32(2)
68 | Er | 162 |  | 161.928 778(4) | 0.001 39(5) | 167.259(3) | g
|  | 164 |  | 163.929 200(3) | 0.016 01(3)
|  | 166 |  | 165.930 2931(27) | 0.335 03(36)
|  | 167 |  | 166.932 0482(27) | 0.228 69(9)
|  | 168 |  | 167.932 3702(27) | 0.269 78(18)
|  | 170 |  | 169.935 4643(30) | 0.149 10(36)
69 | Tm | 169 |  | 168.934 2133(27) | 1.0000 | 168.934 21(2)
70 | Yb | 168 |  | 167.933 897(5) | 0.0013(1) | 173.054(5) | g
|  | 170 |  | 169.934 7618(26) | 0.0304(15)
|  | 171 |  | 170.936 3258(26) | 0.1428(57)
|  | 172 |  | 171.936 3815(26) | 0.2183(67)
|  | 173 |  | 172.938 2108(26) | 0.1613(27)
|  | 174 |  | 173.938 8621(26) | 0.3183(92)
|  | 176 |  | 175.942 5717(28) | 0.1276(41)
71 | Lu | 175 |  | 174.940 7718(23) | 0.9741(2) | 174.9668(1) | g
|  | 176 |  | 175.942 6863(23) | 0.0259(2)
72 | Hf | 174 |  | 173.940 046(3) | 0.0016(1) | 178.49(2) | f
|  | 176 |  | 175.941 4086(24) | 0.0526(7)
|  | 177 |  | 176.943 2207(23) | 0.1860(9)
|  | 178 |  | 177.943 6988(23) | 0.2728(7)
|  | 179 |  | 178.945 8161(23) | 0.1362(2)
|  | 180 |  | 179.946 5500(23) | 0.3508(16)
73 | Ta | 180 |  | 179.947 4648(24) | 0.000 12(2) | 180.947 88(2)
|  | 181 |  | 180.947 9958(19) | 0.999 88(2)
74 | W | 180 |  | 179.946 704(4) | 0.0012(1) | 183.84(1)
|  | 182 |  | 181.948 204 2(9) | 0.2650(16)
|  | 183 |  | 182.950 223 0(9) | 0.1431(4)
|  | 184 |  | 183.950 931 2(9) | 0.3064(2)
|  | 186 |  | 185.954 3641(19) | 0.2843(19)
75 | Re | 185 |  | 184.952 9550(13) | 0.3740(2) | 186.207(1)
|  | 187 |  | 186.955 7531(15) | 0.6260(2)
76 | Os | 184 |  | 183.952 4891(14) | 0.0002(1) | 190.23(3) | g,f
|  | 186 |  | 185.953 8382(15) | 0.0159(3)
|  | 187 |  | 186.955 7505(15) | 0.0196(2)
|  | 188 |  | 187.955 8382(15) | 0.1324(8)
|  | 189 |  | 188.958 1475(16) | 0.1615(5)
|  | 190 |  | 189.958 4470(16) | 0.2626(2)
|  | 192 |  | 191.961 4807(27) | 0.4078(19)
77 | Ir | 191 |  | 190.960 5940(18) | 0.373(2) | 192.217(3)
|  | 193 |  | 192.962 9264(18) | 0.627(2)
78 | Pt | 190 |  | 189.959 932(6) | 0.000 14(1) | 195.084(9)
|  | 192 |  | 191.961 0380(27) | 0.007 82(7)
|  | 194 |  | 193.962 680 3(9) | 0.329 67(99)
|  | 195 |  | 194.964 791 1(9) | 0.338 32(10)
|  | 196 |  | 195.964 951 5(9) | 0.252 42(41)
|  | 198 |  | 197.967 893(3) | 0.071 63(55)
79 | Au | 197 |  | 196.966 568 7(6) | 1.0000 | 196.966 569(4)
80 | Hg | 196 |  | 195.965 833(3) | 0.0015(1) | 200.59(2)
|  | 198 |  | 197.966 769 0(4) | 0.0997(20)
|  | 199 |  | 198.968 279 9(4) | 0.1687(22)
|  | 200 |  | 199.968 326 0(4) | 0.2310(19)
|  | 201 |  | 200.970 302 3(6) | 0.1318(9)
|  | 202 |  | 201.970 643 0(6) | 0.2986(26)
|  | 204 |  | 203.973 493 9(4) | 0.0687(15)
81 | Tl | 203 |  | 202.972 3442(14) | 0.2952(1) | 204.3833(2)
|  | 205 |  | 204.974 4275(14) | 0.7048(1)
82 | Pb | 204 |  | 203.973 0436(13) | 0.014(1) | 207.2(1) | g,r,f
|  | 206 |  | 205.974 4653(13) | 0.241(1)
|  | 207 |  | 206.975 8969(13) | 0.221(1)
|  | 208 |  | 207.976 6521(13) | 0.524(1)
83 | Bi | 209 |  | 208.980 3987(16) | 1.0000 | 208.980 40(1)
84 | Po | 209 |  | 208.982 4304(20) |  | [209]
|  | 210 |  | 209.982 8737(13)
85 | At | 210 |  | 209.987 148(8) |  | [210]
|  | 211 |  | 210.987 4963(30)
86 | Rn | 211 |  | 210.990 601(7) |  | [222]
|  | 220 |  | 220.011 3940(24)
|  | 222 |  | 222.017 5777(25)
87 | Fr | 223 |  | 223.019 7359(26) |  | [223]
88 | Ra | 223 |  | 223.018 5022(27) |  | [226]
|  | 224 |  | 224.020 2118(24)
|  | 226 |  | 226.025 4098(25)
|  | 228 |  | 228.031 0703(26)
89 | Ac | 227 |  | 227.027 7521(26) |  | [227]
90 | Th | 230 |  | 230.033 1338(19)
|  | 232 |  | 232.038 0553(21) | 1.0000 | 232.038 06(2) | g
91 | Pa | 231 |  | 231.035 8840(24) | 1.0000 | 231.035 88(2)
92 | U | 233 |  | 233.039 6352(29)
|  | 234 |  | 234.040 9521(20) | 0.000 054(5) | 238.028 91(3) | g,m,c
|  | 235 |  | 235.043 9299(20) | 0.007 204(6)
|  | 236 |  | 236.045 5680(20)
|  | 238 |  | 238.050 7882(20) | 0.992 742(10)
93 | Np | 236 |  | 236.046 570(50) |  | [237]
|  | 237 |  | 237.048 1734(20)
94 | Pu | 238 |  | 238.049 5599(20) |  | [244]
|  | 239 |  | 239.052 1634(20)
|  | 240 |  | 240.053 8135(20)
|  | 241 |  | 241.056 8515(20)
|  | 242 |  | 242.058 7426(20)
|  | 244 |  | 244.064 204(5)
95 | Am | 241 |  | 241.056 8291(20) |  | [243]
|  | 243 |  | 243.061 3811(25)
96 | Cm | 243 |  | 243.061 3891(22) |  | [247]
|  | 244 |  | 244.062 7526(20)
|  | 245 |  | 245.065 4912(22)
|  | 246 |  | 246.067 2237(22)
|  | 247 |  | 247.070 354(5)
|  | 248 |  | 248.072 349(5)
97 | Bk | 247 |  | 247.070 307(6) |  | [247]
|  | 249 |  | 249.074 9867(28)
98 | Cf | 249 |  | 249.074 8535(24) |  | [251]
|  | 250 |  | 250.076 4061(22)
|  | 251 |  | 251.079 587(5)
|  | 252 |  | 252.081 626(5)
99 | Es | 252 |  | 252.082 980(50) |  | [252]
100 | Fm | 257 |  | 257.095 105(7) |  | [257]
101 | Md | 258 |  | 258.098 431(5) |  | [258]
|  | 260 |  | 260.103 65(34)#
102 | No | 259 |  | 259.101 03(11)# |  | [259]
103 | Lr | 262 |  | 262.109 63(22)# |  | [262]
104 | Rf | 265 |  | 265.116 70(46)# |  | [265]
105 | Db | 268 |  | 268.125 45(57)# |  | [268]
106 | Sg | 271 |  | 271.133 47(70)# |  | [271]
107 | Bh | 272 |  | 272.138 03(65)# |  | [272]
108 | Hs | 270 |  | 270.134 65(31)# |  | [270]
109 | Mt | 276 |  | 276.151 16(73)# |  | [276]
110 | Ds | 281 |  | 281.162 06(78)# |  | [281]
111 | Rg | 280 |  | 280.164 47(80)# |  | [280]
112 | Cn | 285 |  | 285.174 11(78)# |  | [285]
113 | Uut | 284 |  | 284.178 08(86)# |  | [284]
114 | Uuq | 289 |  | 289.187 28(79)# |  | [289]
115 | Uup | 288 |  | 288.192 49(92)# |  | [288]
116 | Uuh | 293 |  |  |  | [293]
117 | Uus | 292 |  | 292.207 55(101)# |  | [292]
118 | Uuo | 294 |  |  |  | [294]

## Unimod
UNIMOD |  | http://www.unimod.org |  | interactive list of protein modifications for mass spectrometry
last updated: | 2011-07-14
| Accession # | PSI-MS Name | Interim name | Description | Monoisotopic mass | Average mass | Composition
View | 765 |  | Met-loss | Removal of initiator methionine from protein N-terminus | -131.0405 | -131.1961 | H(-9) C(-5) N(-1) O(-1) S(-1)
View | 676 |  | Trp->Gly | Trp->Gly substitution | -129.0578 | -129.1586 | H(-7) C(-9) N(-1)
View | 313 | Lys-loss | -K | Loss of C-terminal K from Heavy Chain of MAb | -128.095 | -128.1723 | H(-12) C(-6) N(-2) O(-1)
View | 1224 |  | Trp->Ala | Trp->Ala substitution | -115.0422 | -115.132 | H(-5) C(-8) N(-1)
View | 1239 |  | Tyr->Gly | Tyr->Gly substitution | -106.0419 | -106.1219 | H(-6) C(-7) O(-1)
View | 646 |  | Arg->Gly | Arg->Gly substitution | -99.0796 | -99.1344 | H(-9) C(-4) N(-3)
View | 673 |  | Trp->Ser | Trp->Ser substitution | -99.0473 | -99.1326 | H(-5) C(-8) N(-1) O
View | 400 | Tyr->Dha | DehydroalaY | Dehydroalanine (from Tyrosine) | -94.0419 | -94.1112 | H(-6) C(-6) O(-1)
View | 1237 |  | Tyr->Ala | Tyr->Ala substitution | -92.0262 | -92.0954 | H(-4) C(-6) O(-1)
View | 1093 |  | Phe->Gly | Phe->Gly substitution | -90.0469 | -90.1225 | H(-6) C(-7)
View | 766 |  | Met-loss+Acetyl | Removal of initiator methionine from protein N-terminus, then acetylation of the new N-terminus | -89.0299 | -89.1594 | H(-7) C(-3) N(-1) S(-1)
View | 1232 |  | Trp->Pro | Trp->Pro substitution | -89.0265 | -89.0947 | H(-3) C(-6) N(-1)
View | 1235 |  | Trp->Val | Trp->Val substitution | -87.0109 | -87.0788 | H(-1) C(-6) N(-1)
View | 1189 |  | Arg->Ala | Arg->Ala substitution | -85.064 | -85.1078 | H(-7) C(-3) N(-3)
View | 1234 |  | Trp->Thr | Trp->Thr substitution | -85.0316 | -85.106 | H(-3) C(-7) N(-1) O
View | 674 |  | Trp->Cys | Trp->Cys substitution | -83.0701 | -83.067 | H(-5) C(-8) N(-1) S
View | 1117 |  | His->Gly | His->Gly substitution | -80.0374 | -80.088 | H(-4) C(-4) N(-2)
View | 679 |  | Tyr->Ser | Tyr->Ser substitution | -76.0313 | -76.096 | H(-4) C(-6)
View | 1090 |  | Phe->Ala | Phe->Ala substitution | -76.0313 | -76.096 | H(-4) C(-6)
View | 1147 |  | Met->Gly | Met->Gly substitution | -74.019 | -74.1447 | H(-6) C(-3) S(-1)
View | 677 |  | Trp->Xle | Trp->Leu/Ile substitution | -72.9952 | -73.0523 | H C(-5) N(-1)
View | 1231 |  | Trp->Asn | Trp->Asn substitution | -72.0364 | -72.1073 | H(-4) C(-7) O
View | 564 |  | Glu->Gly | Glu->Gly substitution | -72.0211 | -72.0627 | H(-4) C(-3) O(-2)
View | 1135 |  | Lys->Gly | Lys->Gly substitution | -71.0735 | -71.121 | H(-9) C(-4) N(-1)
View | 1225 |  | Trp->Asp | Trp->Asp substitution | -71.0524 | -71.1225 | H(-5) C(-7) N(-1) O(2)
View | 1181 |  | Gln->Gly | Gln->Gly substitution | -71.0371 | -71.0779 | H(-5) C(-3) N(-1) O(-1)
View | 636 |  | Arg->Ser | Arg->Ser substitution | -69.0691 | -69.1084 | H(-7) C(-3) N(-3) O
View | 1113 |  | His->Ala | His->Ala substitution | -66.0218 | -66.0614 | H(-2) C(-3) N(-2)
View | 1242 |  | Tyr->Pro | Tyr->Pro substitution | -66.0106 | -66.0581 | H(-2) C(-4) O(-1)
View | 1246 |  | Tyr->Val | Tyr->Val substitution | -63.9949 | -64.0422 | C(-4) O(-1)
View | 1245 |  | Tyr->Thr | Tyr->Thr substitution | -62.0157 | -62.0694 | H(-2) C(-5)
View | 683 |  | Tyr->Cys | Tyr->Cys substitution | -60.0541 | -60.0304 | H(-4) C(-6) O(-1) S
View | 566 |  | Phe->Ser | Phe->Ser substitution | -60.0364 | -60.0966 | H(-4) C(-6) O
View | 1142 |  | Met->Ala | Met->Ala substitution | -60.0034 | -60.1182 | H(-4) C(-2) S(-1)
View | 639 |  | Arg->Pro | Arg->Pro substitution | -59.0483 | -59.0705 | H(-5) C(-1) N(-3)
View | 1233 |  | Trp->Gln | Trp->Gln substitution | -58.0207 | -58.0807 | H(-2) C(-6) O
View | 556 |  | Asp->Gly | Asp->Gly substitution | -58.0055 | -58.0361 | H(-2) C(-2) O(-2)
View | 560 |  | Glu->Ala | Glu->Ala substitution | -58.0055 | -58.0361 | H(-2) C(-2) O(-2)
View | 822 |  | Gly-loss+Amide | Enzymatic glycine removal leaving an amidated C-terminus | -58.0055 | -58.0361 | H(-2) C(-2) O(-2)
View | 1229 |  | Trp->Lys | Trp->Lys substitution | -57.9843 | -58.0376 | H(2) C(-5)
View | 1131 |  | Lys->Ala | Lys->Ala substitution | -57.0578 | -57.0944 | H(-7) C(-3) N(-1)
View | 1226 |  | Trp->Glu | Trp->Glu substitution | -57.0367 | -57.0959 | H(-3) C(-6) N(-1) O(2)
View | 1193 |  | Arg->Val | Arg->Val substitution | -57.0327 | -57.0546 | H(-3) C(-1) N(-3)
View | 1159 |  | Asn->Gly | Asn->Gly substitution | -57.0215 | -57.0513 | H(-3) C(-2) N(-1) O(-1)
View | 1177 |  | Gln->Ala | Gln->Ala substitution | -57.0215 | -57.0513 | H(-3) C(-2) N(-1) O(-1)
View | 1129 |  | Xle->Gly | Leu/Ile->Gly substitution | -56.0626 | -56.1063 | H(-8) C(-4)
View | 638 |  | Arg->Thr | Arg->Thr substitution | -55.0534 | -55.0818 | H(-5) C(-2) N(-3) O
View | 1230 |  | Trp->Met | Trp->Met substitution | -55.0388 | -55.0138 | H(-1) C(-6) N(-1) S
View | 644 |  | Arg->Cys | Arg->Cys substitution | -53.0919 | -53.0428 | H(-7) C(-3) N(-3) S
View | 1121 |  | His->Ser | His->Ser substitution | -50.0269 | -50.062 | H(-2) C(-3) N(-2) O
View | 1098 |  | Phe->Pro | Phe->Pro substitution | -50.0157 | -50.0587 | H(-2) C(-4)
View | 1248 |  | Tyr->Xle | Tyr->Leu/Ile substitution | -49.9793 | -50.0156 | H(2) C(-3) O(-1)
View | 680 |  | Tyr->Asn | Tyr->Asn substitution | -49.0204 | -49.0706 | H(-3) C(-5) N
View | 1228 |  | Trp->His | Trp->His substitution | -49.0204 | -49.0706 | H(-3) C(-5) N
View | 682 |  | Tyr->Asp | Tyr->Asp substitution | -48.0364 | -48.0859 | H(-4) C(-5) O
View | 11 | Met->Hsl | Hse_lact | Homoserine lactone | -48.0034 | -48.1075 | H(-4) C(-1) S(-1)
View | 526 | Dethiomethyl | Dethiomethyl | Prompt loss of side chain from oxidised Met | -48.0034 | -48.1075 | H(-4) C(-1) S(-1)
View | 570 |  | Phe->Val | Phe->Val substitution | -48 | -48.0428 | C(-4)
View | 1101 |  | Phe->Thr | Phe->Thr substitution | -46.0207 | -46.07 | H(-2) C(-5) O
View | 140 | a-type-ion | a-type_Ion | ISD a-series (C-Term) | -46.0055 | -46.0254 | H(-2) C(-1) O(-2)
View | 552 |  | Cys->Gly | Cys->Gly substitution | -45.9877 | -46.0916 | H(-2) C(-1) S(-1)
View | 567 |  | Phe->Cys | Phe->Cys substitution | -44.0592 | -44.031 | H(-4) C(-6) S
View | 1207 |  | Thr->Gly | Thr->Gly substitution | -44.0262 | -44.0526 | H(-4) C(-2) O(-1)
View | 1152 |  | Met->Ser | Met->Ser substitution | -44.0085 | -44.1188 | H(-4) C(-2) O S(-1)
View | 553 |  | Asp->Ala | Asp->Ala substitution | -43.9898 | -44.0095 | C(-1) O(-2)
View | 344 | Arg->GluSA | Argglutamicsealde | Arginine oxidation to glutamic semialdehyde | -43.0534 | -43.0711 | H(-5) C(-1) N(-3) O
View | 645 |  | Arg->Xle | Arg->Leu/Ile substitution | -43.017 | -43.028 | H(-1) N(-3)
View | 1155 |  | Asn->Ala | Asn->Ala substitution | -43.0058 | -43.0247 | H(-1) C(-1) N(-1) O(-1)
View | 1192 |  | Arg->Asn | Arg->Asn substitution | -42.0582 | -42.083 | H(-6) C(-2) N(-2) O
View | 1125 |  | Xle->Ala | Leu/Ile->Ala substitution | -42.047 | -42.0797 | H(-6) C(-3)
View | 672 |  | Val->Gly | Val->Gly substitution | -42.047 | -42.0797 | H(-6) C(-3)
View | 372 | Arg->Orn | Arg2Orn | Ornithine from Arginine | -42.0218 | -42.04 | H(-2) C(-1) N(-2)
View | 1086 |  | Glu->Ser | Glu->Ser substitution | -42.0106 | -42.0367 | H(-2) C(-2) O(-1)
View | 1190 |  | Arg->Asp | Arg->Asp substitution | -41.0742 | -41.0983 | H(-7) C(-2) N(-3) O(2)
View | 1138 |  | Lys->Ser | Lys->Ser substitution | -41.0629 | -41.095 | H(-7) C(-3) N(-1) O
View | 1184 |  | Gln->Ser | Gln->Ser substitution | -41.0265 | -41.0519 | H(-3) C(-2) N(-1)
View | 1170 |  | Pro->Gly | Pro->Gly substitution | -40.0313 | -40.0639 | H(-4) C(-3)
View | 580 |  | His->Pro | His->Pro substitution | -40.0061 | -40.0241 | C(-1) N(-2)
View | 1227 |  | Trp->Phe | Trp->Phe substitution | -39.0109 | -39.036 | H(-1) C(-2) N(-1)
View | 1123 |  | His->Val | His->Val substitution | -37.9905 | -38.0082 | H(2) C(-1) N(-2)
View | 1122 |  | His->Thr | His->Thr substitution | -36.0112 | -36.0354 | C(-2) N(-2) O
View | 1243 |  | Tyr->Gln | Tyr->Gln substitution | -35.0048 | -35.044 | H(-1) C(-4) N
View | 1240 |  | Tyr->Lys | Tyr->Lys substitution | -34.9684 | -35.001 | H(3) C(-3) N O(-1)
View | 1114 |  | His->Cys | His->Cys substitution | -34.0497 | -33.9964 | H(-2) C(-3) N(-2) S
View | 1238 |  | Tyr->Glu | Tyr->Glu substitution | -34.0207 | -34.0593 | H(-2) C(-4) O
View | 368 | Cys->Dha | DehydroalaC | Dehydroalanine (from Cysteine) | -33.9877 | -34.0809 | H(-2) S(-1)
View | 1150 |  | Met->Pro | Met->Pro substitution | -33.9877 | -34.0809 | H(-2) S(-1)
View | 568 |  | Phe->Xle | Phe->Leu/Ile substitution | -33.9843 | -34.0162 | H(2) C(-3)
View | 1097 |  | Phe->Asn | Phe->Asn substitution | -33.0255 | -33.0712 | H(-3) C(-5) N O
View | 382 | Cys->PyruvicAcid | pyruvicC | pyruvic acid from N-term cys | -33.0037 | -33.0961 | H(-3) N(-1) O S(-1)
View | 1091 |  | Phe->Asp | Phe->Asp substitution | -32.0415 | -32.0865 | H(-4) C(-5) O(2)
View | 1241 |  | Tyr->Met | Tyr->Met substitution | -32.0228 | -31.9772 | C(-4) O(-1) S
View | 1084 |  | Glu->Pro | Glu->Pro substitution | -31.9898 | -31.9988 | O(-2)
View | 615 |  | Met->Val | Met->Val substitution | -31.9721 | -32.065 | S(-1)
View | 1055 |  | Cys->Ala | Cys->Ala substitution | -31.9721 | -32.065 | S(-1)
View | 1137 |  | Lys->Pro | Lys->Pro substitution | -31.0422 | -31.0571 | H(-5) C(-1) N(-1)
View | 630 |  | Gln->Pro | Gln->Pro substitution | -31.0058 | -31.014 | H(-1) N(-1) O(-1)
View | 360 | Pro->Pyrrolidinone | Pyrrolidinone | Proline oxidation to pyrrolidinone | -30.0106 | -30.026 | H(-2) C(-1) O(-1)
View | 657 |  | Ser->Gly | Ser->Gly substitution | -30.0106 | -30.026 | H(-2) C(-1) O(-1)
View | 659 |  | Thr->Ala | Thr->Ala substitution | -30.0106 | -30.026 | H(-2) C(-1) O(-1)
View | 10 | Met->Hse | HSe | Homoserine | -29.9928 | -30.0922 | H(-2) C(-1) O S(-1)
View | 610 |  | Met->Thr | Met->Thr substitution | -29.9928 | -30.0922 | H(-2) C(-1) O S(-1)
View | 675 |  | Trp->Arg | Trp->Arg substitution | -29.9782 | -30.0242 | H(2) C(-5) N(2)
View | 565 |  | Glu->Val | Glu->Val substitution | -29.9742 | -29.9829 | H(2) O(-2)
View | 1139 |  | Lys->Val | Lys->Val substitution | -29.0265 | -29.0412 | H(-3) C(-1) N(-1)
View | 1186 |  | Gln->Val | Gln->Val substitution | -28.9902 | -28.9982 | H N(-1) O(-1)
View | 642 |  | Arg->Gln | Arg->Gln substitution | -28.0425 | -28.0565 | H(-4) C(-1) N(-2) O
View | 667 |  | Val->Ala | Val->Ala substitution | -28.0313 | -28.0532 | H(-4) C(-2)
View | 1143 |  | Met->Cys | Met->Cys substitution | -28.0313 | -28.0532 | H(-4) C(-2)
View | 640 |  | Arg->Lys | Arg->Lys substitution | -28.0061 | -28.0134 | N(-2)
View | 369 | Pro->Pyrrolidone | OxPro | Pyrrolidone from Proline | -27.9949 | -28.0101 | C(-1) O(-1)
View | 1075 |  | Asp->Ser | Asp->Ser substitution | -27.9949 | -28.0101 | C(-1) O(-1)
View | 1087 |  | Glu->Thr | Glu->Thr substitution | -27.9949 | -28.0101 | C(-1) O(-1)
View | 1191 |  | Arg->Glu | Arg->Glu substitution | -27.0585 | -27.0717 | H(-5) C(-1) N(-3) O(2)
View | 594 |  | Lys->Thr | Lys->Thr substitution | -27.0473 | -27.0684 | H(-5) C(-2) N(-1) O
View | 616 |  | Asn->Ser | Asn->Ser substitution | -27.0109 | -27.0253 | H(-1) C(-1) N(-1)
View | 1185 |  | Gln->Thr | Gln->Thr substitution | -27.0109 | -27.0253 | H(-1) C(-1) N(-1)
View | 601 |  | Xle->Ser | Leu/Ile->Ser substitution | -26.052 | -26.0803 | H(-6) C(-3) O
View | 1078 |  | Glu->Cys | Glu->Cys substitution | -26.0334 | -25.9711 | H(-2) C(-2) O(-2) S
View | 624 |  | Pro->Ala | Pro->Ala substitution | -26.0157 | -26.0373 | H(-2) C(-2)
View | 681 |  | Tyr->His | Tyr->His substitution | -26.0044 | -26.034 | H(-2) C(-3) N(2) O(-1)
View | 1132 |  | Lys->Cys | Lys->Cys substitution | -25.0858 | -25.0294 | H(-7) C(-3) N(-1) S
View | 643 |  | Arg->Met | Arg->Met substitution | -25.0606 | -24.9896 | H(-3) C(-1) N(-3) S
View | 1178 |  | Gln->Cys | Gln->Cys substitution | -25.0494 | -24.9863 | H(-3) C(-2) N(-1) O(-1) S
View | 585 |  | His->Xle | His->Leu/Ile substitution | -23.9748 | -23.9816 | H(4) N(-2)
View | 348 | His->Asn | His2Asn | His->Asn substitution | -23.016 | -23.0366 | H(-1) C(-2) N(-1) O
View | 1236 |  | Trp->Tyr | Trp->Tyr substitution | -23.016 | -23.0366 | H(-1) C(-2) N(-1) O
View | 349 | His->Asp | His2Asp | His->Asp substitution | -22.032 | -22.0519 | H(-2) C(-2) N(-2) O(2)
View | 899 |  | Met->Hpg | methionine replacement by homopropargylglycine | -21.9877 | -22.0702 | H(-2) C S(-1)
View | 641 |  | Arg->His | Arg->His substitution | -19.0422 | -19.0464 | H(-5) N(-1)
View | 1099 |  | Phe->Gln | Phe->Gln substitution | -19.0098 | -19.0446 | H(-1) C(-4) N O
View | 1095 |  | Phe->Lys | Phe->Lys substitution | -18.9735 | -19.0016 | H(3) C(-3) N
View | 1092 |  | Phe->Glu | Phe->Glu substitution | -18.0258 | -18.0599 | H(-2) C(-4) O(2)
View | 23 | Dehydrated | Phospho+PL | Dehydration | -18.0106 | -18.0153 | H(-2) O(-1)
View | 27 | Glu->pyro-Glu | Pyro_glu | Pyro-glu from E | -18.0106 | -18.0153 | H(-2) O(-1)
View | 402 | Cys->Oxoalanine | oxoalanine | oxoalanine | -17.9928 | -18.0815 | H(-2) O S(-1)
View | 1072 |  | Asp->Pro | Asp->Pro substitution | -17.9742 | -17.9722 | H(2) C O(-2)
View | 614 |  | Met->Xle | Met->Leu/Ile substitution | -17.9564 | -18.0384 | H(2) C S(-1)
View | 28 | Gln->pyro-Glu | Pyro-glu | Pyro-glu from Q | -17.0265 | -17.0305 | H(-3) N(-1)
View | 385 | Ammonia-loss | N-oxobutanoic | Loss of ammonia | -17.0265 | -17.0305 | H(-3) N(-1)
View | 1149 |  | Met->Asn | Met->Asn substitution | -16.9976 | -17.0934 | H(-3) C(-1) N O S(-1)
View | 1161 |  | Asn->Pro | Asn->Pro substitution | -16.9902 | -16.9875 | H C N(-1) O(-1)
View | 604 |  | Xle->Pro | Leu/Ile->Pro substitution | -16.0313 | -16.0425 | H(-4) C(-1)
View | 1096 |  | Phe->Met | Phe->Met substitution | -16.0279 | -15.9778 | C(-4) S
View | 1144 |  | Met->Asp | Met->Asp substitution | -16.0135 | -16.1087 | H(-4) C(-1) O(2) S(-1)
View | 447 | Deoxy | semialdehyde | reduction | -15.9949 | -15.9994 | O(-1)
View | 648 |  | Ser->Ala | Ser->Ala substitution | -15.9949 | -15.9994 | O(-1)
View | 678 |  | Tyr->Phe | Tyr->Phe substitution | -15.9949 | -15.9994 | O(-1)
View | 548 |  | Cys->Ser | Cys->Ser substitution | -15.9772 | -16.0656 | O S(-1)
View | 559 |  | Asp->Val | Asp->Val substitution | -15.9585 | -15.9563 | H(4) C O(-2)
View | 1081 |  | Glu->Xle | Glu->Leu/Ile substitution | -15.9585 | -15.9563 | H(4) C O(-2)
View | 403 | Ser->LacticAcid | lactic | lactic acid from N-term Ser | -15.0109 | -15.0146 | H(-1) N(-1)
View | 600 |  | Lys->Xle | Lys->Leu/Ile substitution | -15.0109 | -15.0146 | H(-1) N(-1)
View | 991 |  | ISD_z+2_ion | ISD (z+2)-series | -15.0109 | -15.0146 | H(-1) N(-1)
View | 1083 |  | Glu->Asn | Glu->Asn substitution | -14.9997 | -15.0113 | H(-1) C(-1) N O(-1)
View | 635 |  | Gln->Xle | Gln->Leu/Ile substitution | -14.9745 | -14.9716 | H(3) C N(-1) O(-1)
View | 1164 |  | Asn->Val | Asn->Val substitution | -14.9745 | -14.9716 | H(3) C N(-1) O(-1)
View | 595 |  | Lys->Asn | Lys->Asn substitution | -14.052 | -14.0696 | H(-6) C(-2) O
View | 544 |  | Ala->Gly | Ala->Gly substitution | -14.0157 | -14.0266 | H(-2) C(-1)
View | 562 |  | Glu->Asp | Glu->Asp substitution | -14.0157 | -14.0266 | H(-2) C(-1)
View | 605 |  | Xle->Val | Leu/Ile->Val substitution | -14.0157 | -14.0266 | H(-2) C(-1)
View | 658 |  | Thr->Ser | Thr->Ser substitution | -14.0157 | -14.0266 | H(-2) C(-1)
View | 1183 |  | Gln->Asn | Gln->Asn substitution | -14.0157 | -14.0266 | H(-2) C(-1)
View | 1076 |  | Asp->Thr | Asp->Thr substitution | -13.9793 | -13.9835 | H(2) O(-1)
View | 1133 |  | Lys->Asp | Lys->Asp substitution | -13.068 | -13.0849 | H(-7) C(-2) N(-1) O(2)
View | 1179 |  | Gln->Asp | Gln->Asp substitution | -13.0316 | -13.0418 | H(-3) C(-1) N(-1) O
View | 617 |  | Asn->Thr | Asn->Thr substitution | -12.9952 | -12.9988 | H N(-1)
View | 588 |  | Xle->Thr | Leu/Ile->Thr substitution | -12.0364 | -12.0538 | H(-4) C(-2) O
View | 1220 |  | Val->Ser | Val->Ser substitution | -12.0364 | -12.0538 | H(-4) C(-2) O
View | 1067 |  | Asp->Cys | Asp->Cys substitution | -12.0178 | -11.9445 | C(-1) O(-2) S
View | 1156 |  | Asn->Cys | Asn->Cys substitution | -11.0337 | -10.9597 | H(-1) C(-1) N(-1) O(-1) S
View | 1126 |  | Xle->Cys | Leu/Ile->Cys substitution | -10.0749 | -10.0147 | H(-6) C(-3) S
View | 623 |  | Pro->Ser | Pro->Ser substitution | -10.0207 | -10.0379 | H(-2) C(-2) O
View | 1094 |  | Phe->His | Phe->His substitution | -10.0095 | -10.0346 | H(-2) C(-3) N(2)
View | 1195 |  | Arg->Phe | Arg->Phe substitution | -9.0327 | -9.0118 | H(-3) C(3) N(-3)
View | 582 |  | His->Gln | His->Gln substitution | -9.0003 | -9.0101 | H C(-1) N(-1) O
View | 1119 |  | His->Lys | His->Lys substitution | -8.9639 | -8.967 | H(5) N(-1)
View | 1115 |  | His->Glu | His->Glu substitution | -8.0163 | -8.0253 | C(-1) N(-2) O(2)
View | 1244 |  | Tyr->Arg | Tyr->Arg substitution | -6.9622 | -6.9876 | H(3) C(-3) N(3) O(-1)
View | 1120 |  | His->Met | His->Met substitution | -6.0184 | -5.9432 | H(2) C(-1) N(-2) S
View | 1063 |  | Cys->Pro | Cys->Pro substitution | -5.9564 | -6.0277 | H(2) C(2) S(-1)
View | 896 |  | Met->Aha | Methionine replacement by azido homoalanine | -4.9863 | -5.0794 | H(-3) C(-1) N(3) S(-1)
View | 662 |  | Thr->Pro | Thr->Pro substitution | -3.9949 | -3.9887 | C O(-1)
View | 1066 |  | Cys->Val | Cys->Val substitution | -3.9408 | -4.0118 | H(4) C(2) S(-1)
View | 1151 |  | Met->Gln | Met->Gln substitution | -2.9819 | -3.0668 | H(-1) N O S(-1)
View | 613 |  | Met->Lys | Met->Lys substitution | -2.9455 | -3.0238 | H(3) C N S(-1)
View | 939 |  | Cys->methylaminoAla | carbamidomethylated Cys that undergoes beta-elimination and Michael addition of methylamine | -2.9455 | -3.0238 | H(3) C N S(-1)
View | 401 | Didehydro | didehydro | 2-amino-3-oxo-butanoic_acid | -2.0156 | -2.0159 | H(-2)
View | 1217 |  | Val->Pro | Val->Pro substitution | -2.0156 | -2.0159 | H(-2)
View | 1145 |  | Met->Glu | Met->Glu substitution | -1.9979 | -2.0821 | H(-2) O(2) S(-1)
View | 1210 |  | Thr->Val | Thr->Val substitution | -1.9793 | -1.9728 | H(2) C O(-1)
View | 1065 |  | Cys->Thr | Cys->Thr substitution | -1.9615 | -2.039 | H(2) C O S(-1)
View | 1069 |  | Asp->Xle | Asp->Leu/Ile substitution | -1.9429 | -1.9298 | H(6) C(2) O(-2)
View | 352 | Lys->Allysine | Lysaminoadipicsealde | Lysine oxidation to aminoadipic semialdehyde | -1.0316 | -1.0311 | H(-3) N(-1) O
View | 374 | Dehydro | Cystine | Half of a disulfide bridge | -1.0078 | -1.0079 | H(-1)
View | 2 | Amidated | Amide | Amidation | -0.98402 | -0.9848 | H N O(-1)
View | 555 |  | Asp->Asn | Asp->Asn substitution | -0.98402 | -0.9848 | H N O(-1)
View | 561 |  | Glu->Gln | Glu->Gln substitution | -0.98402 | -0.9848 | H N O(-1)
View | 622 |  | Asn->Xle | Asn->Leu/Ile substitution | -0.95886 | -0.945 | H(5) C(2) N(-1) O(-1)
View | 563 |  | Glu->Lys | Glu->Lys substitution | -0.94763 | -0.9417 | H(5) C N O(-2)
View | 597 |  | Lys->Gln | Lys->Gln substitution | -0.036386 | -0.0431 | H(-4) C(-1) O
View | 631 |  | Gln->Lys | Gln->Lys substitution | 0.036386 | 0.0431 | H(4) C O(-1)
View | 596 |  | Lys->Glu | Lys->Glu substitution | 0.94763 | 0.9417 | H(-5) C(-1) N(-1) O(2)
View | 589 |  | Xle->Asn | Leu/Ile->Asn substitution | 0.95886 | 0.945 | H(-5) C(-2) N O
View | 7 | Deamidated | Deamidation | Deamidation | 0.98402 | 0.9848 | H(-1) N(-1) O
View | 621 |  | Asn->Asp | Asn->Asp substitution | 0.98402 | 0.9848 | H(-1) N(-1) O
View | 632 |  | Gln->Glu | Gln->Glu substitution | 0.98402 | 0.9848 | H(-1) N(-1) O
View | 994 |  | Label:15N(1) | SILAS 15N(1) | 0.99704 | 0.9934 | N(-1) 15N
View | 1127 |  | Xle->Asp | Leu/Ile->Asp substitution | 1.9429 | 1.9298 | H(-6) C(-2) O(2)
View | 1203 |  | Thr->Cys | Thr->Cys substitution | 1.9615 | 2.039 | H(-2) C(-1) O(-1) S
View | 1221 |  | Val->Thr | Val->Thr substitution | 1.9793 | 1.9728 | H(-2) C(-1) O
View | 995 |  | Label:15N(2) | SILAS 15N(2) | 1.9941 | 1.9868 | N(-2) 15N(2)
View | 1082 |  | Glu->Met | Glu->Met substitution | 1.9979 | 2.0821 | H(2) O(-2) S
View | 258 | Label:18O(1) | O18_STY | O18 Labeling | 2.0042 | 1.9998 | O(-1) 18O
View | 1174 |  | Pro->Val | Pro->Val substitution | 2.0156 | 2.0159 | H(2)
View | 598 |  | Lys->Met | Lys->Met substitution | 2.9455 | 3.0238 | H(-3) C(-1) N(-1) S
View | 1182 |  | Gln->Met | Gln->Met substitution | 2.9819 | 3.0668 | H N(-1) O(-1) S
View | 170 | Delta:H(1)O(-1)18O(1) | GN18Olabeling | glycosylated asparagine 18O labeling | 2.9883 | 2.9845 | H(-1) N(-1) 18O
View | 366 | Deamidated:18O(1) | Deamidation_O18 | Deamidation in presence of O18 | 2.9883 | 2.9845 | H(-1) N(-1) 18O
View | 996 |  | Label:15N(3) | SILAS 15N(3) | 2.9911 | 2.9802 | N(-3) 15N(3)
View | 262 | Label:2H(3) | D3_Label | Trideuteration | 3.0188 | 3.0185 | H(-3) 2H(3)
View | 1213 |  | Val->Cys | Val->Cys substitution | 3.9408 | 4.0118 | H(-4) C(-2) S
View | 897 |  | Label:15N(4) | SILAC 15N(4) | 3.9881 | 3.9736 | N(-4) 15N(4)
View | 351 | Trp->Kynurenin | kynurenin | tryptophan oxidation to kynurenin | 3.9949 | 3.9887 | C(-1) O
View | 627 |  | Pro->Thr | Pro->Thr substitution | 3.9949 | 3.9887 | C(-1) O
View | 193 | Label:18O(2) | double_O18 | O18 label at both C-terminal oxygens | 4.0085 | 3.9995 | O(-2) 18O(2)
View | 862 |  | Label:13C(1)2H(3) | SILAC | 4.0222 | 4.0111 | H(-3) 2H(3) C(-1) 13C
View | 481 | Label:2H(4) | Lys4 | 4,4,5,5-D4 Lysine | 4.0251 | 4.0246 | H(-4) 2H(4)
View | 772 | Label:13C(5) | 13C5 | 13C(5) Silac label | 5.0168 | 4.9633 | C(-5) 13C(5)
View | 1166 |  | Pro->Cys | Pro->Cys substitution | 5.9564 | 6.0277 | H(-2) C(-2) S
View | 950 |  | Cation:Li | Replacement of proton by lithium | 6.0082 | 5.9331 | H(-1) Li
View | 268 | Label:13C(5)15N(1) | 13C5-15N1 | 13C(5) 15N(1) Silac label | 6.0138 | 5.9567 | C(-5) 13C(5) N(-1) 15N
View | 1148 |  | Met->His | Met->His substitution | 6.0184 | 5.9432 | H(-2) C N(2) S(-1)
View | 188 | Label:13C(6) | 13C6 | 13C(6) Silac label | 6.0201 | 5.9559 | C(-6) 13C(6)
View | 1194 |  | Arg->Tyr | Arg->Tyr substitution | 6.9622 | 6.9876 | H(-3) C(3) N(-3) O
View | 695 | Label:13C(6)15N(1) | Label:13C(6)15N(1) | 13C(6) 15N(1) Silac label | 7.0172 | 6.9493 | C(-6) 13C(6) N(-1) 15N
View | 259 | Label:13C(6)15N(2) | 13C6-15N2 | 13C(6) 15N(2) Silac label | 8.0142 | 7.9427 | C(-6) 13C(6) N(-2) 15N(2)
View | 1080 |  | Glu->His | Glu->His substitution | 8.0163 | 8.0253 | C N(2) O(-2)
View | 1136 |  | Lys->His | Lys->His substitution | 8.9639 | 8.967 | H(-5) N
View | 633 |  | Gln->His | Gln->His substitution | 9.0003 | 9.0101 | H(-1) C N O(-1)
View | 184 | Label:13C(9) | 13C9 | 13C(9) Silac label | 9.0302 | 8.9339 | C(-9) 13C(9)
View | 1100 |  | Phe->Arg | Phe->Arg substitution | 9.0327 | 9.0118 | H(3) C(-3) N(3)
View | 267 | Label:13C(6)15N(4) | 13C6-15N4 | 13C(6) 15N(4) Silac label | 10.0083 | 9.9296 | C(-6) 13C(6) N(-4) 15N(4)
View | 1116 |  | His->Phe | His->Phe substitution | 10.0095 | 10.0346 | H(2) C(3) N(-2)
View | 652 |  | Ser->Pro | Ser->Pro substitution | 10.0207 | 10.0379 | H(2) C(2) O(-1)
View | 854 |  | Label:13C(8)15N(2) | 13C(8) 15N(2) Silac label | 10.0209 | 9.9281 | C(-8) 13C(8) N(-2) 15N(2)
View | 269 | Label:13C(9)15N(1) | 13C9-15N1 | 13C(9) 15N(1) Silac label | 10.0272 | 9.9273 | C(-9) 13C(9) N(-1) 15N
View | 1059 |  | Cys->Xle | Cys->Leu/Ile substitution | 10.0749 | 10.0147 | H(6) C(3) S(-1)
View | 1062 |  | Cys->Asn | Cys->Asn substitution | 11.0337 | 10.9597 | H C N O S(-1)
View | 944 |  | Label:15N(2)2H(9) | SILAC label | 11.0506 | 11.0423 | H(-9) 2H(9) N(-2) 15N(2)
View | 940 |  | Cys->ethylaminoAla | Carbamidomethylated Cys that undergoes beta-elimination and Michael addition of ethylamine | 11.0701 | 11.0028 | H(5) C(2) N S(-1)
View | 1009 |  | Thiazolidine | Addition of Carbon to cysteine | 12 | 12.0107 | C
View | 1056 |  | Cys->Asp | Cys->Asp substitution | 12.0178 | 11.9445 | C O(2) S(-1)
View | 664 |  | Thr->Xle | Thr->Leu/Ile substitution | 12.0364 | 12.0538 | H(4) C(2) O(-1)
View | 1202 |  | Ser->Val | Ser->Val substitution | 12.0364 | 12.0538 | H(4) C(2) O(-1)
View | 904 |  | Phe->CamCys | Phe->Cys substitution and carbamidomethylation | 12.9622 | 13.0204 | H(-1) C(-4) N O S
View | 660 |  | Thr->Asn | Thr->Asn substitution | 12.9952 | 12.9988 | H(-1) N
View | 337 | Methylamine | MethylamineST | Michael addition with methylamine | 13.0316 | 13.0418 | H(3) C N O(-1)
View | 1073 |  | Asp->Gln | Asp->Gln substitution | 13.0316 | 13.0418 | H(3) C N O(-1)
View | 1070 |  | Asp->Lys | Asp->Lys substitution | 13.068 | 13.0849 | H(7) C(2) N O(-2)
View | 359 | Pro->pyro-Glu | Pyroglutamic | proline oxidation to pyroglutamic acid | 13.9793 | 13.9835 | H(-2) O
View | 288 | Trp->Oxolactone | oxolactone | Tryptophan oxidation to oxolactone | 13.9793 | 13.9835 | H(-2) O
View | 1204 |  | Thr->Asp | Thr->Asp substitution | 13.9793 | 13.9835 | H(-2) O
View | 34 | Methyl | Methyl | Methylation | 14.0157 | 14.0266 | H(2) C
View | 558 |  | Asp->Glu | Asp->Glu substitution | 14.0157 | 14.0266 | H(2) C
View | 571 |  | Gly->Ala | Gly->Ala substitution | 14.0157 | 14.0266 | H(2) C
View | 650 |  | Ser->Thr | Ser->Thr substitution | 14.0157 | 14.0266 | H(2) C
View | 671 |  | Val->Xle | Val->Leu/Ile substitution | 14.0157 | 14.0266 | H(2) C
View | 1162 |  | Asn->Gln | Asn->Gln substitution | 14.0157 | 14.0266 | H(2) C
View | 618 |  | Asn->Lys | Asn->Lys substitution | 14.052 | 14.0696 | H(6) C(2) O(-1)
View | 381 | Lys->AminoadipicAcid | aminoadipic | alpha-amino adipic acid | 14.9633 | 14.9683 | H(-3) N(-1) O(2)
View | 607 |  | Xle->Gln | Leu/Ile->Gln substitution | 14.9745 | 14.9716 | H(-3) C(-1) N O
View | 1216 |  | Val->Asn | Val->Asn substitution | 14.9745 | 14.9716 | H(-3) C(-1) N O
View | 528 | Methyl+Deamidated | Methyl+Deamidated | Deamidation followed by a methylation | 14.9997 | 15.0113 | H C N(-1) O
View | 1157 |  | Asn->Glu | Asn->Glu substitution | 14.9997 | 15.0113 | H C N(-1) O
View | 342 | Amino | aminotyrosine | Tyrosine oxidation to 2-aminotyrosine | 15.0109 | 15.0146 | H N
View | 590 |  | Xle->Lys | Leu/Ile->Lys substitution | 15.0109 | 15.0146 | H N
View | 670 |  | Val->Asp | Val->Asp substitution | 15.9585 | 15.9563 | H(-4) C(-1) O(2)
View | 1128 |  | Xle->Glu | Leu/Ile->Glu substitution | 15.9585 | 15.9563 | H(-4) C(-1) O(2)
View | 420 | Carboxy->Thiocarboxy | thiocarboxy | thiocarboxylic acid | 15.9772 | 16.0656 | O(-1) S
View | 654 |  | Ser->Cys | Ser->Cys substitution | 15.9772 | 16.0656 | O(-1) S
View | 35 | Oxidation | Hydroxylation | Oxidation or Hydroxylation | 15.9949 | 15.9994 | O
View | 540 |  | Ala->Ser | Ala->Ser substitution | 15.9949 | 15.9994 | O
View | 569 |  | Phe->Tyr | Phe->Tyr substitution | 15.9949 | 15.9994 | O
View | 1071 |  | Asp->Met | Asp->Met substitution | 16.0135 | 16.1087 | H(4) C O(-2) S
View | 1146 |  | Met->Phe | Met->Phe substitution | 16.0279 | 15.9778 | C(4) S(-1)
View | 284 | Methyl:2H(2) | DeMet | Deuterium Methylation of Lysine | 16.0282 | 16.0389 | 2H(2) C
View | 629 |  | Pro->Xle | Pro->Leu/Ile substitution | 16.0313 | 16.0425 | H(4) C
View | 1173 |  | Pro->Asn | Pro->Asn substitution | 16.9902 | 16.9875 | H(-1) C(-1) N O
View | 1160 |  | Asn->Met | Asn->Met substitution | 16.9976 | 17.0934 | H(3) C N(-1) O(-1) S
View | 989 |  | Ammonium | replacement of proton with ammonium ion | 17.0265 | 17.0305 | H(3) N
View | 298 | Methyl:2H(3) | D3-Me-ester | deuterated methyl ester | 17.0345 | 17.0451 | H(-1) 2H(3) C
View | 696 | Label:2H(9)13C(6)15N(2) | 13C6-15N2-D9 | 13C(6) 15N(2) (D)9 SILAC label | 17.0707 | 16.9982 | H(-9) 2H(9) C(-6) 13C(6) N(-2) 15N(2)
View | 608 |  | Xle->Met | Leu/Ile->Met substitution | 17.9564 | 18.0384 | H(-2) C(-1) S
View | 1167 |  | Pro->Asp | Pro->Asp substitution | 17.9742 | 17.9722 | H(-2) C(-1) O(2)
View | 127 | Fluoro | Fphe | fluorophenylalanine replacement of phenylalanine | 17.9906 | 17.9905 | H(-1) F
View | 1079 |  | Glu->Phe | Glu->Phe substitution | 18.0258 | 18.0599 | H(2) C(4) O(-2)
View | 329 | Methyl:2H(3)13C(1) | Methyl-Heavy | monomethylated arginine | 18.0378 | 18.0377 | H(-1) 2H(3) 13C
View | 906 |  | Lys->MetOx | Lys->Met substitution and sulfoxidation | 18.9404 | 19.0232 | H(-3) C(-1) N(-1) O S
View | 1134 |  | Lys->Phe | Lys->Phe substitution | 18.9735 | 19.0016 | H(-3) C(3) N(-1)
View | 1180 |  | Gln->Phe | Gln->Phe substitution | 19.0098 | 19.0446 | H C(4) N(-1) O(-1)
View | 584 |  | His->Arg | His->Arg substitution | 19.0422 | 19.0464 | H(5) N
View | 350 | Trp->Hydroxykynurenin | hydroxykynurenin | tryptophan oxidation to hydroxykynurenin | 19.9898 | 19.9881 | C(-1) O(2)
View | 885 |  | Label:13C(1)2H(3)+Oxidation | Oxidised methionine 13C(1)2H(3) SILAC label | 20.0171 | 20.0105 | H(-3) 2H(3) C(-1) 13C O
View | 956 |  | Cation:Mg[II] | Replacement of 2 protons by magnesium | 21.9694 | 22.2891 | H(-2) Mg
View | 30 | Cation:Na | Sodiated | Sodium adduct | 21.9819 | 21.9818 | H(-1) Na
View | 554 |  | Asp->His | Asp->His substitution | 22.032 | 22.0519 | H(2) C(2) N(2) O(-2)
View | 620 |  | Asn->His | Asn->His substitution | 23.016 | 23.0366 | H C(2) N O(-1)
View | 1247 |  | Tyr->Trp | Tyr->Trp substitution | 23.016 | 23.0366 | H C(2) N O(-1)
View | 606 |  | Xle->His | Leu/Ile->His substitution | 23.9748 | 23.9816 | H(-4) N(2)
View | 1004 |  | Label:13C(6)15N(4)+Methyl | Monomethylated Arg13C(6) 15N(4) | 24.0239 | 23.9561 | H(2) C(-5) 13C(6) N(-4) 15N(4)
View | 438 | Cyano | cyano | cyano | 24.9952 | 25.0095 | H(-1) C N
View | 1064 |  | Cys->Gln | Cys->Gln substitution | 25.0494 | 24.9863 | H(3) C(2) N O S(-1)
View | 611 |  | Met->Arg | Met->Arg substitution | 25.0606 | 24.9896 | H(3) C N(3) S(-1)
View | 1060 |  | Cys->Lys | Cys->Lys substitution | 25.0858 | 25.0294 | H(7) C(3) N S(-1)
View | 581 |  | His->Tyr | His->Tyr substitution | 26.0044 | 26.034 | H(2) C(3) N(-2) O
View | 254 | Delta:H(2)C(2) | Acetald+26 | Acetaldehyde +26 | 26.0157 | 26.0373 | H(2) C(2)
View | 543 |  | Ala->Pro | Ala->Pro substitution | 26.0157 | 26.0373 | H(2) C(2)
View | 1057 |  | Cys->Glu | Cys->Glu substitution | 26.0334 | 25.9711 | H(2) C(2) O(2) S(-1)
View | 656 |  | Ser->Xle | Ser->Leu/Ile substitution | 26.052 | 26.0803 | H(6) C(3) O(-1)
View | 651 |  | Ser->Asn | Ser->Asn substitution | 27.0109 | 27.0253 | H C N
View | 1209 |  | Thr->Gln | Thr->Gln substitution | 27.0109 | 27.0253 | H C N
View | 661 |  | Thr->Lys | Thr->Lys substitution | 27.0473 | 27.0684 | H(5) C(2) N O(-1)
View | 926 |  | ethylamino | ethyl amino | 27.0473 | 27.0684 | H(5) C(2) N O(-1)
View | 1085 |  | Glu->Arg | Glu->Arg substitution | 27.0585 | 27.0717 | H(5) C N(3) O(-2)
View | 122 | Formyl | Formyl | Formylation | 27.9949 | 28.0101 | C O
View | 1196 |  | Ser->Asp | Ser->Asp substitution | 27.9949 | 28.0101 | C O
View | 1205 |  | Thr->Glu | Thr->Glu substitution | 27.9949 | 28.0101 | C O
View | 599 |  | Lys->Arg | Lys->Arg substitution | 28.0061 | 28.0134 | N(2)
View | 36 | Dimethyl | di-Methylation | di-Methylation | 28.0313 | 28.0532 | H(4) C(2)
View | 255 | Delta:H(4)C(2) | Acetald+28 | Acetaldehyde +28 | 28.0313 | 28.0532 | H(4) C(2)
View | 280 | Ethyl | Ethyl | Ethylation | 28.0313 | 28.0532 | H(4) C(2)
View | 546 |  | Ala->Val | Ala->Val substitution | 28.0313 | 28.0532 | H(4) C(2)
View | 1061 |  | Cys->Met | Cys->Met substitution | 28.0313 | 28.0532 | H(4) C(2)
View | 634 |  | Gln->Arg | Gln->Arg substitution | 28.0425 | 28.0565 | H(4) C N(2) O(-1)
View | 1006 |  | Label:13C(6)15N(4)+Methyl:2H(3)13C(1) | 2H(3) 13C(1) monomethylated Arg13C(6) 15N(4) | 28.0461 | 27.9673 | H(-1) 2H(3) C(-6) 13C(7) N(-4) 15N(4)
View | 275 | Nitrosyl | SNO | S-nitrosylation | 28.9902 | 28.9982 | H(-1) N O
View | 1218 |  | Val->Gln | Val->Gln substitution | 28.9902 | 28.9982 | H(-1) N O
View | 1215 |  | Val->Lys | Val->Lys substitution | 29.0265 | 29.0412 | H(3) C N
View | 529 | Delta:H(5)C(2) | Delta:H(5)C(2) | Dimethylation of proline residue | 29.0391 | 29.0611 | H(5) C(2)
View | 931 |  | EthylAmide | solvolysis of amide group on Asn or Gln by ethanol | 29.0391 | 29.0611 | H(5) C(2)
View | 392 | Quinone | quinone | quinone | 29.9742 | 29.9829 | H(-2) O(2)
View | 668 |  | Val->Glu | Val->Glu substitution | 29.9742 | 29.9829 | H(-2) O(2)
View | 637 |  | Arg->Trp | Arg->Trp substitution | 29.9782 | 30.0242 | H(-2) C(5) N(-2)
View | 663 |  | Thr->Met | Thr->Met substitution | 29.9928 | 30.0922 | H(2) C O(-1) S
View | 414 | Hydroxymethyl | hydroxymethyl | hydroxymethyl | 30.0106 | 30.026 | H(2) C O
View | 541 |  | Ala->Thr | Ala->Thr substitution | 30.0106 | 30.026 | H(2) C O
View | 572 |  | Gly->Ser | Gly->Ser substitution | 30.0106 | 30.026 | H(2) C O
View | 626 |  | Pro->Gln | Pro->Gln substitution | 31.0058 | 31.014 | H N O
View | 1171 |  | Pro->Lys | Pro->Lys substitution | 31.0422 | 31.0571 | H(5) C N
View | 903 |  | Lys->CamCys | Lys->Cys substitution and carbamidomethylation | 31.9357 | 32.0219 | H(-4) C(-1) O S
View | 421 | Sulfide | persulfide | persulfide | 31.9721 | 32.065 | S
View | 669 |  | Val->Met | Val->Met substitution | 31.9721 | 32.065 | S
View | 1044 |  | Ala->Cys | Ala->Cys substitution | 31.9721 | 32.065 | S
View | 425 | Dioxidation | dihydroxy | dihydroxy | 31.9898 | 31.9988 | O(2)
View | 1168 |  | Pro->Glu | Pro->Glu substitution | 31.9898 | 31.9988 | O(2)
View | 1154 |  | Met->Tyr | Met->Tyr substitution | 32.0228 | 31.9772 | C(4) O S(-1)
View | 1068 |  | Asp->Phe | Asp->Phe substitution | 32.0415 | 32.0865 | H(4) C(5) O(-2)
View | 199 | Dimethyl:2H(4) | CHD2 | DiMethyl-CHD2 | 32.0564 | 32.0778 | 2H(4) C(2)
View | 1158 |  | Asn->Phe | Asn->Phe substitution | 33.0255 | 33.0712 | H(3) C(5) N(-1) O(-1)
View | 905 |  | Leu->MetOx | Leu->Met substitution and sulfoxidation | 33.9513 | 34.0378 | H(-2) C(-1) O S
View | 602 |  | Xle->Phe | Leu/Ile->Phe substitution | 33.9843 | 34.0162 | H(-2) C(3)
View | 1172 |  | Pro->Met | Pro->Met substitution | 33.9877 | 34.0809 | H(2) S
View | 1089 |  | Glu->Tyr | Glu->Tyr substitution | 34.0207 | 34.0593 | H(2) C(4) O(-1)
View | 1058 |  | Cys->His | Cys->His substitution | 34.0497 | 33.9964 | H(2) C(3) N(2) S(-1)
View | 986 |  | Label:13C(6)+Dimethyl | Dimethyl 13C(6) Silac label | 34.0514 | 34.0091 | H(4) C(-4) 13C(6)
View | 510 | Dimethyl:2H(4)13C(2) | C13HD2 | DiMethyl-C13HD2 | 34.0631 | 34.0631 | 2H(4) 13C(2)
View | 1141 |  | Lys->Tyr | Lys->Tyr substitution | 34.9684 | 35.001 | H(-3) C(3) N(-1) O
View | 936 |  | Chlorination | Chlorination of tyrosine residues | 34.9689 | 35.453 | Cl
View | 1188 |  | Gln->Tyr | Gln->Tyr substitution | 35.0048 | 35.044 | H C(4) N(-1)
View | 1208 |  | Thr->His | Thr->His substitution | 36.0112 | 36.0354 | C(2) N(2) O(-1)
View | 987 |  | Label:13C(6)15N(2)+Dimethyl | Dimethyl 13C(6)15N(2) Silac label | 36.0455 | 35.9959 | H(4) C(-4) 13C(6) N(-2) 15N(2)
View | 330 | Dimethyl:2H(6)13C(2) | Dimethyl-Heavy | dimethylated arginine | 36.0757 | 36.0754 | H(-2) 2H(6) 13C(2)
View | 958 |  | Propargylamine | propargylamine | 37.0316 | 37.0632 | H(3) C(3) N O(-1)
View | 951 |  | Cation:Ca[II] | Replacement of 2 protons by calcium | 37.9469 | 38.0621 | H(-2) Ca
View | 530 | Cation:K | Cation:K | Replacement of proton by potassium | 37.9559 | 38.0904 | H(-1) K
View | 1214 |  | Val->His | Val->His substitution | 37.9905 | 38.0082 | H(-2) C N(2)
View | 207 | Delta:H(2)C(3) | Acrolein38 | Acrolein addition +38 | 38.0157 | 38.048 | H(2) C(3)
View | 1005 |  | Label:13C(6)15N(4)+Dimethyl | Dimethylated Arg13C(6) 15N(4) | 38.0396 | 37.9827 | H(4) C(-4) 13C(6) N(-4) 15N(4)
View | 1102 |  | Phe->Trp | Phe->Trp substitution | 39.0109 | 39.036 | H C(2) N
View | 26 | Pyro-carbamidomethyl | Pyro-cmC | S-carbamoylmethylcysteine cyclization (N-terminus) | 39.9949 | 40.0208 | C(2) O
View | 860 |  | G-H1 | Glyoxal-derived hydroimiadazolone | 39.9949 | 40.0208 | C(2) O
View | 625 |  | Pro->His | Pro->His substitution | 40.0061 | 40.0241 | C N(2)
View | 256 | Delta:H(4)C(3) | Propionald+40 | Propionaldehyde +40 | 40.0313 | 40.0639 | H(4) C(3)
View | 1109 |  | Gly->Pro | Gly->Pro substitution | 40.0313 | 40.0639 | H(4) C(3)
View | 141 | Amidine | amidine | amidination of lysines or N-terminal amines with methyl acetimidate | 41.0265 | 41.0519 | H(3) C(2) N
View | 1201 |  | Ser->Gln | Ser->Gln substitution | 41.0265 | 41.0519 | H(3) C(2) N
View | 1199 |  | Ser->Lys | Ser->Lys substitution | 41.0629 | 41.095 | H(7) C(3) N O(-1)
View | 1074 |  | Asp->Arg | Asp->Arg substitution | 41.0742 | 41.0983 | H(7) C(2) N(3) O(-2)
View | 1 | Acetyl | Acetyl | Acetylation | 42.0106 | 42.0367 | H(2) C(2) O
View | 1197 |  | Ser->Glu | Ser->Glu substitution | 42.0106 | 42.0367 | H(2) C(2) O
View | 52 | Guanidinyl | Guanidination | Guanidination | 42.0218 | 42.04 | H(2) C N(2)
View | 440 | Amidino | amidino | amidino | 42.0218 | 42.04 | H(2) C N(2)
View | 37 | Trimethyl | tri-Methylation | tri-Methylation | 42.047 | 42.0797 | H(6) C(3)
View | 575 |  | Gly->Val | Gly->Val substitution | 42.047 | 42.0797 | H(6) C(3)
View | 1047 |  | Ala->Xle | Ala->Leu/Ile substitution | 42.047 | 42.0797 | H(6) C(3)
View | 1163 |  | Asn->Arg | Asn->Arg substitution | 42.0582 | 42.083 | H(6) C(2) N(2) O(-1)
View | 5 | Carbamyl | Carbamyl | Carbamylation | 43.0058 | 43.0247 | H C N O
View | 1050 |  | Ala->Asn | Ala->Asn substitution | 43.0058 | 43.0247 | H C N O
View | 609 |  | Xle->Arg | Leu/Ile->Arg substitution | 43.017 | 43.028 | H N(3)
View | 734 |  | Ethanolamine | Carboxyl modification with ethanolamine | 43.0422 | 43.0678 | H(5) C(2) N
View | 299 | Carboxy | carboxyl | Carboxylation | 43.9898 | 44.0095 | C O(2)
View | 542 |  | Ala->Asp | Ala->Asp substitution | 43.9898 | 44.0095 | C O(2)
View | 327 | Delta:H(4)C(2)O(-1)S(1) | S-Eth | S-Ethylcystine from Serine | 44.0085 | 44.1188 | H(4) C(2) O(-1) S
View | 1200 |  | Ser->Met | Ser->Met substitution | 44.0085 | 44.1188 | H(4) C(2) O(-1) S
View | 278 | Ethanolyl | EtOH | Ethanolation | 44.0262 | 44.0526 | H(4) C(2) O
View | 1111 |  | Gly->Thr | Gly->Thr substitution | 44.0262 | 44.0526 | H(4) C(2) O
View | 547 |  | Cys->Phe | Cys->Phe substitution | 44.0592 | 44.031 | H(4) C(6) S(-1)
View | 354 | Nitro | Nitro | Oxidation to nitro | 44.9851 | 44.9976 | H(-1) N O(2)
View | 56 | Acetyl:2H(3) | Acetyl_heavy | Acetate labeling reagent (N-term & K) (heavy form, +3amu) | 45.0294 | 45.0552 | H(-1) 2H(3) C(2) O
View | 39 | Methylthio | b-methylthiol | Beta-methylthiolation | 45.9877 | 46.0916 | H(2) C S
View | 577 |  | Gly->Cys | Gly->Cys substitution | 45.9877 | 46.0916 | H(2) C S
View | 1206 |  | Thr->Phe | Thr->Phe substitution | 46.0207 | 46.07 | H(2) C(5) O(-1)
View | 834 |  | Label:2H(4)+Acetyl | Acetyl 4,4,5,5-D4 Lysine | 46.0357 | 46.0613 | H(-2) 2H(4) C(2) O
View | 1007 |  | Label:13C(6)15N(4)+Dimethyl:2H(6)13C(2) | 2H(6) 13C(2) Dimethylated Arg13C(6) 15N(4) | 46.0839 | 46.005 | H(-2) 2H(6) C(-6) 13C(8) N(-4) 15N(4)
View | 162 | Delta:S(-1)Se(1) | SeMet | Selenium replaces sulphur | 47.9444 | 46.895 | S(-1) Se
View | 345 | Trioxidation | Cysteic_acid | cysteine oxidation to cysteic acid | 47.9847 | 47.9982 | O(3)
View | 666 |  | Val->Phe | Val->Phe substitution | 48 | 48.0428 | C(4)
View | 835 |  | Label:13C(6)+Acetyl | Acetyl 13C(6) Silac label | 48.0307 | 47.9926 | H(2) C(-4) 13C(6) O
View | 557 |  | Asp->Tyr | Asp->Tyr substitution | 48.0364 | 48.0859 | H(4) C(5) O(-1)
View | 619 |  | Asn->Tyr | Asn->Tyr substitution | 49.0204 | 49.0706 | H(3) C(5) N(-1)
View | 1124 |  | His->Trp | His->Trp substitution | 49.0204 | 49.0706 | H(3) C(5) N(-1)
View | 1130 |  | Xle->Tyr | Leu/Ile->Tyr substitution | 49.9793 | 50.0156 | H(-2) C(3) O
View | 1169 |  | Pro->Phe | Pro->Phe substitution | 50.0157 | 50.0587 | H(2) C(4)
View | 836 |  | Label:13C(6)15N(2)+Acetyl | Acetyl_13C(6) 15N(2) Silac label | 50.0248 | 49.9794 | H(2) C(-4) 13C(6) N(-2) 15N(2) O
View | 1198 |  | Ser->His | Ser->His substitution | 50.0269 | 50.062 | H(2) C(3) N(2) O(-1)
View | 551 |  | Cys->Arg | Cys->Arg substitution | 53.0919 | 53.0428 | H(7) C(3) N(3) S(-1)
View | 952 |  | Cation:Fe[II] | Replacement of 2 protons by iron | 53.9193 | 53.8291 | H(-2) Fe
View | 750 |  | trifluoro | trifluoroleucine replacement of leucine | 53.9717 | 53.9714 | H(-3) F(3)
View | 319 | Delta:H(2)C(3)O(1) | MDA54 | MDA adduct +54 | 54.0106 | 54.0474 | H(2) C(3) O
View | 859 |  | MG-H1 | Methylglyoxal-derived hydroimidazolone | 54.0106 | 54.0474 | H(2) C(3) O
View | 1153 |  | Met->Trp | Met->Trp substitution | 55.0388 | 55.0138 | H C(6) N S(-1)
View | 665 |  | Thr->Arg | Thr->Arg substitution | 55.0534 | 55.0818 | H(5) C(2) N(3) O(-1)
View | 953 |  | Cation:Ni[II] | Replacement of 2 protons by nickel | 55.9197 | 56.6775 | H(-2) Ni
View | 58 | Propionyl | Propionyl_light | Propionate labeling reagent light form (N-term & K) | 56.0262 | 56.0633 | H(4) C(3) O
View | 206 | Delta:H(4)C(3)O(1) | Acrolein56 | Acrolein addition +56 | 56.0262 | 56.0633 | H(4) C(3) O
View | 518 | Diethyl | Diethylation | Diethylation, analogous to Dimethylation | 56.0626 | 56.1063 | H(8) C(4)
View | 1105 |  | Gly->Xle | Gly->Leu/Ile substitution | 56.0626 | 56.1063 | H(8) C(4)
View | 4 | Carbamidomethyl | Carbamidomethyl | Iodoacetamide derivative | 57.0215 | 57.0513 | H(3) C(2) N O
View | 1051 |  | Ala->Gln | Ala->Gln substitution | 57.0215 | 57.0513 | H(3) C(2) N O
View | 1108 |  | Gly->Asn | Gly->Asn substitution | 57.0215 | 57.0513 | H(3) C(2) N O
View | 1263 |  | Gly | Addition of Glycine | 57.0215 | 57.0513 | H(3) C(2) N O
View | 1219 |  | Val->Arg | Val->Arg substitution | 57.0327 | 57.0546 | H(3) C N(3)
View | 1088 |  | Glu->Trp | Glu->Trp substitution | 57.0367 | 57.0959 | H(3) C(6) N O(-2)
View | 1048 |  | Ala->Lys | Ala->Lys substitution | 57.0578 | 57.0944 | H(7) C(3) N
View | 1140 |  | Lys->Trp | Lys->Trp substitution | 57.9843 | 58.0376 | H(-2) C(5)
View | 6 | Carboxymethyl | Carboxymethyl | Iodoacetic acid derivative | 58.0055 | 58.0361 | H(2) C(2) O(2)
View | 545 |  | Ala->Glu | Ala->Glu substitution | 58.0055 | 58.0361 | H(2) C(2) O(2)
View | 576 |  | Gly->Asp | Gly->Asp substitution | 58.0055 | 58.0361 | H(2) C(2) O(2)
View | 1187 |  | Gln->Trp | Gln->Trp substitution | 58.0207 | 58.0807 | H(2) C(6) O(-1)
View | 977 |  | Carbofuran | 2,3-dihydro-2,2-dimethyl-7-benzofuranol N-methyl carbamate | 58.0293 | 58.0593 | H(4) C(2) N O
View | 472 | AEC-MAEC | AEC-MAEC | aminoethylcysteine | 59.0194 | 59.1334 | H(5) C(2) N O(-1) S
View | 59 | Propionyl:13C(3) | Propionyl_heavy | Propionate labeling reagent heavy form (+3amu), N-term & K | 59.0363 | 59.0412 | H(4) 13C(3) O
View | 768 |  | Methyl+Acetyl:2H(3) | Mono-methylated lysine labelled with Acetyl_heavy | 59.045 | 59.0817 | H 2H(3) C(3) O
View | 628 |  | Pro->Arg | Pro->Arg substitution | 59.0483 | 59.0705 | H(5) C N(3)
View | 445 | Hydroxytrimethyl | trimethyl-OH | 5-hydroxy-N6,N6,N6-trimethyl | 59.0497 | 59.0871 | H(7) C(3) O
View | 928 |  | MercaptoEthanol | 2-OH-ethyl thio-Ser | 60.0034 | 60.1182 | H(4) C(2) S
View | 1049 |  | Ala->Met | Ala->Met substitution | 60.0034 | 60.1182 | H(4) C(2) S
View | 775 |  | Carboxymethyl:13C(2) | Iodoacetic acid derivative w/ 13C label | 60.0122 | 60.0214 | H(2) 13C(2) O(2)
View | 647 |  | Ser->Phe | Ser->Phe substitution | 60.0364 | 60.0966 | H(4) C(6) O(-1)
View | 550 |  | Cys->Tyr | Cys->Tyr substitution | 60.0541 | 60.0304 | H(4) C(6) O S(-1)
View | 954 |  | Cation:Zn[II] | Replacement of 2 protons by zinc | 61.9135 | 63.3931 | H(-2) Zn
View | 531 | Cation:Cu[I] | Cation:Cu | Replacement of proton by copper | 61.9218 | 62.5381 | H(-1) Cu
View | 318 | Delta:H(2)C(5) | MDA62 | MDA adduct +62 | 62.0157 | 62.0694 | H(2) C(5)
View | 1212 |  | Thr->Tyr | Thr->Tyr substitution | 62.0157 | 62.0694 | H(2) C(5)
View | 792 |  | AEC-MAEC:2H(4) | deuterium cysteamine modification to S or T | 63.0445 | 63.158 | H 2H(4) C(2) N O(-1) S
View | 1223 |  | Val->Tyr | Val->Tyr substitution | 63.9949 | 64.0422 | C(4) O
View | 1176 |  | Pro->Tyr | Pro->Tyr substitution | 66.0106 | 66.0581 | H(2) C(4) O
View | 1046 |  | Ala->His | Ala->His substitution | 66.0218 | 66.0614 | H(2) C(3) N(2)
View | 94 | IMID | IMID_light | IMID d0 | 68.0374 | 68.0773 | H(4) C(3) N(2)
View | 520 | Piperidine | Piperidination | Piperidination | 68.0626 | 68.117 | H(8) C(5)
View | 655 |  | Ser->Arg | Ser->Arg substitution | 69.0691 | 69.1084 | H(7) C(3) N(3) O(-1)
View | 937 |  | dichlorination | Dichlorination of tyrosine residues | 69.9377 | 70.906 | Cl(2)
View | 422 | PyruvicAcidIminyl | pyruv-iminyl | N-pyruvic acid 2-iminyl | 70.0055 | 70.0468 | H(2) C(3) O(2)
View | 253 | Crotonaldehyde | Croton | Crotonaldehyde | 70.0419 | 70.0898 | H(6) C(4) O
View | 24 | Propionamide | Propionamide | Acrylamide adduct | 71.0371 | 71.0779 | H(5) C(3) N O
View | 1110 |  | Gly->Gln | Gly->Gln substitution | 71.0371 | 71.0779 | H(5) C(3) N O
View | 1077 |  | Asp->Trp | Asp->Trp substitution | 71.0524 | 71.1225 | H(5) C(7) N O(-2)
View | 1041 |  | Deoxyhypusine | Deoxyhypusine | 71.0735 | 71.121 | H(9) C(4) N
View | 1106 |  | Gly->Lys | Gly->Lys substitution | 71.0735 | 71.121 | H(9) C(4) N
View | 378 | Carboxyethyl | carboxyethyl | carboxyethyl | 72.0211 | 72.0627 | H(4) C(3) O(2)
View | 574 |  | Gly->Glu | Gly->Glu substitution | 72.0211 | 72.0627 | H(4) C(3) O(2)
View | 830 |  | Dihydroxyimidazolidine | Dihydroxy methylglyoxal adduct | 72.0211 | 72.0627 | H(4) C(3) O(2)
View | 1165 |  | Asn->Trp | Asn->Trp substitution | 72.0364 | 72.1073 | H(4) C(7) O(-1)
View | 95 | IMID:2H(4) | IMID_heavy | IMID d4 | 72.0626 | 72.1019 | 2H(4) C(3) N(2)
View | 603 |  | Xle->Trp | Leu/Ile->Trp substitution | 72.9952 | 73.0523 | H(-1) C(5) N
View | 915 |  | Ethoxyformyl | Ethoxyformylation | 73.029 | 73.0706 | H(5) C(3) O(2)
View | 1107 |  | Gly->Met | Gly->Met substitution | 74.019 | 74.1447 | H(6) C(3) S
View | 97 | Propionamide:2H(3) | Acrylamide_heavy | Acrylamide d3 | 74.0559 | 74.0964 | H(2) 2H(3) C(3) N O
View | 200 | Ethanedithiol | EDT | EDT | 75.9805 | 76.1838 | H(4) C(2) O(-1) S(2)
View | 303 | DeStreak | DeStreak | Cysteine mercaptoethanol | 75.9983 | 76.1176 | H(4) C(2) O S
View | 208 | Delta:H(4)C(6) | Acrolein76 | Acrolein addition +76 | 76.0313 | 76.096 | H(4) C(6)
View | 653 |  | Ser->Tyr | Ser->Tyr substitution | 76.0313 | 76.096 | H(4) C(6)
View | 1045 |  | Ala->Phe | Ala->Phe substitution | 76.0313 | 76.096 | H(4) C(6)
View | 340 | Bromo | bromo | bromination | 77.9105 | 78.8961 | H(-1) Br
View | 728 |  | Methylphosphonate | Methylphosphonylation | 77.9871 | 78.0071 | H(3) C O(2) P
View | 316 | DimethylpyrroleAdduct | pyrrole | 2,5-dimethypyrrole | 78.0469 | 78.1118 | H(6) C(6)
View | 423 | Delta:Se(1) | selenyl | selenyl | 79.9165 | 78.96 | Se
View | 40 | Sulfo | Sulfation | O-Sulfonation | 79.9568 | 80.0632 | O(3) S
View | 21 | Phospho | Phospho | Phosphorylation | 79.9663 | 79.9799 | H O(3) P
View | 1104 |  | Gly->His | Gly->His substitution | 80.0374 | 80.088 | H(4) C(4) N(2)
View | 837 |  | Arg->Npo | Arginine replacement by Nitropyrimidyl ornithine | 80.9851 | 81.0297 | H(-1) C(3) N O(2)
View | 549 |  | Cys->Trp | Cys->Trp substitution | 83.0701 | 83.067 | H(5) C(8) N S(-1)
View | 1211 |  | Thr->Trp | Thr->Trp substitution | 85.0316 | 85.106 | H(3) C(7) N O(-1)
View | 211 | NEIAA | NEIAA-d0 | N-ethyl iodoacetamide-d0 | 85.0528 | 85.1045 | H(7) C(4) N O
View | 1052 |  | Ala->Arg | Ala->Arg substitution | 85.064 | 85.1078 | H(7) C(3) N(3)
View | 801 |  | Pentylamine | Labeling transglutaminase substrate on glutamine side chain | 85.0891 | 85.1475 | H(11) C(5) N
View | 747 |  | Malonyl | Malonylation of C and S residues | 86.0004 | 86.0462 | H(2) C(3) O(3)
View | 371 | HMVK | HMVK86 | Michael addition of hydroxymethylvinyl ketone to cysteine | 86.0368 | 86.0892 | H(6) C(4) O(2)
View | 1222 |  | Val->Trp | Val->Trp substitution | 87.0109 | 87.0788 | H C(6) N
View | 324 | DTBP | DTBP | dimethyl 3,3\'-dithiobispropionimidate | 87.0143 | 87.1435 | H(5) C(3) N S
View | 1014 |  | glycidamide | glycidamide adduct | 87.032 | 87.0773 | H(5) C(3) N O(2)
View | 178 | DAET | ser_thr_DAET | phosphorylation to amine thiol | 87.0507 | 87.1866 | H(9) C(4) N O(-1) S
View | 379 | Hypusine | hypusine | hypusine | 87.0684 | 87.1204 | H(9) C(4) N O
View | 126 | Thioacyl | DSP | 3-sulfanylpropanoyl | 87.9983 | 88.1283 | H(4) C(3) O S
View | 185 | Label:13C(9)+Phospho | 13C9_Phospho_Tyr | C13 label (Phosphotyrosine) | 88.9965 | 88.9138 | H C(-9) 13C(9) O(3) P
View | 1175 |  | Pro->Trp | Pro->Trp substitution | 89.0265 | 89.0947 | H(3) C(6) N
View | 1103 |  | Gly->Phe | Gly->Phe substitution | 90.0469 | 90.1225 | H(6) C(7)
View | 212 | NEIAA:2H(5) | NEIAA-d5 | N-ethyl iodoacetamide-d5 | 90.0841 | 90.1353 | H(2) 2H(5) C(4) N O
View | 1054 |  | Ala->Tyr | Ala->Tyr substitution | 92.0262 | 92.0954 | H(4) C(6) O
View | 724 |  | O-Methylphosphate | O-Methylphosphorylation | 93.982 | 94.0065 | H(3) C O(3) P
View | 205 | Delta:H(6)C(6)O(1) | Acrolein94 | Acrolein addition +94 | 94.0419 | 94.1112 | H(6) C(6) O
View | 997 |  | sulfo+amino | aminotyrosine with sulfation | 94.9677 | 95.0778 | H N O(3) S
View | 260 | Thiophospho | Thio-phospho | Thiophosphorylation | 95.9435 | 96.0455 | H O(2) P S
View | 773 |  | maleimide | maleimide | 97.0164 | 97.0721 | H(3) C(4) N O(2)
View | 1042 |  | Acetyldeoxyhypusine | Acetyldeoxyhypusine | 97.0891 | 97.1582 | H(11) C(6) N
View | 1028 |  | Xlink:EGScleaved | EGS crosslinker to Lys or N-terminus following hydroxylamine cleavage | 99.032 | 99.088 | H(5) C(4) N O(2)
View | 649 |  | Ser->Trp | Ser->Trp substitution | 99.0473 | 99.1326 | H(5) C(8) N O(-1)
View | 17 | NIPCAM | NIPCAM | N-isopropylcarboxamidomethyl | 99.0684 | 99.1311 | H(9) C(5) N O
View | 578 |  | Gly->Arg | Gly->Arg substitution | 99.0796 | 99.1344 | H(9) C(4) N(3)
View | 64 | Succinyl | Suc_anh_light | Succinic anhydride labeling reagent light form (N-term & K) | 100.016 | 100.0728 | H(4) C(4) O(3)
View | 914 |  | Methylmalonylation | Methylmalonylation on Serine | 100.016 | 100.0728 | H(4) C(4) O(3)
View | 1257 |  | Ub-Br2 | Ub Bromide probe addition | 100.0637 | 100.1191 | H(8) C(4) N(2) O
View | 902 |  | DimethylArsino | Reaction with dimethylarsinous (AsIII) acid | 103.9607 | 103.9827 | H(5) C(2) As
View | 136 | Benzoyl | benzoyl | labeling reagent light form (N-term & K) | 104.0262 | 104.1061 | H(4) C(7) O
View | 66 | Succinyl:13C(4) | Suc_anh+4C13 | Succinic anhydride labeling reagent, heavy form (+4amu, 4C13), N-term & K | 104.0295 | 104.0434 | H(4) 13C(4) O(3)
View | 65 | Succinyl:2H(4) | Suc_anh+4H2 | Succinic anhydride labeling reagent, heavy form (+4amu, 4H2), N-term & K | 104.0412 | 104.0974 | 2H(4) C(4) O(3)
View | 1008 |  | SecCarbamidomethyl | Sec Iodoacetamide derivative | 104.9659 | 103.9463 | H(3) C(2) N O S(-1) Se
View | 365 | ICPL | ICPL_light | Bruker Daltonics SERVA-ICPL(TM) quantification chemistry, light form | 105.0215 | 105.0941 | H(3) C(6) N O
View | 697 |  | NIC | Nicotinic Acid | 105.0215 | 105.0941 | H(3) C(6) N O
View | 31 | Pyridylethyl | S-pyridylethyl | S-pyridylethylation | 105.0578 | 105.1372 | H(7) C(7) N
View | 955 |  | Cation:Ag | Replacement of proton by silver | 105.8973 | 106.8603 | H(-1) Ag
View | 1112 |  | Gly->Tyr | Gly->Tyr substitution | 106.0419 | 106.1219 | H(6) C(7) O
View | 1254 |  | Saligenin | o-toluene | 106.0419 | 106.1219 | H(6) C(7) O
View | 1000 |  | AHA-Alkyne | Azidohomoalanine (AHA) bound to propargylglycine-NH2 (alkyne) | 107.0773 | 107.0504 | H(5) C(4) N(5) O S(-1)
View | 723 |  | O-Dimethylphosphate | O-Dimethylphosphorylation | 107.9976 | 108.0331 | H(5) C(2) O(3) P
View | 726 |  | O-Ethylphosphate | O-Ethylphosphorylation | 107.9976 | 108.0331 | H(5) C(2) O(3) P
View | 886 |  | HydroxymethylOP | 2-ammonio-6-[4-(hydroxymethyl)-3-oxidopyridinium-1-yl]- hexanoate | 108.0211 | 108.0948 | H(4) C(6) O(2)
View | 687 | ICPL:2H(4) | ICPL_medium | Bruker Daltonics SERVA-ICPL(TM) quantification chemistry, medium form | 109.0466 | 109.1188 | H(-1) 2H(4) C(6) N O
View | 698 |  | dNIC | deuterated Nicotinic Acid | 109.0481 | 109.1205 | H 2H(3) C(6) N O
View | 435 | Methylpyrroline | me-pyrroline | 4-methyl-delta-1-pyrroline-5-carboxyl | 109.0528 | 109.1259 | H(7) C(6) N O
View | 314 | Nmethylmaleimide | NMM | Nmethylmaleimide | 111.032 | 111.0987 | H(5) C(5) N O(2)
View | 364 | ICPL:13C(6) | ICPL_heavy | Bruker Daltonics SERVA-ICPL(TM) quantification chemistry, heavy form | 111.0416 | 111.05 | H(3) 13C(6) N O
View | 1017 |  | DMPO | DMPO spin-trap nitrone adduct | 111.0684 | 111.1418 | H(9) C(6) N O
View | 209 | Delta:H(8)C(6)O(2) | Acrolein112 | Acrolein addition +112 | 112.0524 | 112.1265 | H(8) C(6) O(2)
View | 1043 |  | Acetylhypusine | Acetylhypusine | 113.0841 | 113.1576 | H(11) C(6) N O
View | 121 | GlyGly | GlyGly | ubiquitinylation residue | 114.0429 | 114.1026 | H(6) C(4) N(2) O(2)
View | 1053 |  | Ala->Trp | Ala->Trp substitution | 115.0422 | 115.132 | H(5) C(8) N
View | 866 |  | ICPL:13C(6)2H(4) | Bruker Daltonics SERVA-ICPL(TM) quantification chemistry, +10 Da form | 115.0667 | 115.0747 | H(-1) 2H(4) 13C(6) N O
View | 957 |  | 2-succinyl | S-(2-succinyl) cysteine | 116.011 | 116.0722 | H(4) C(4) O(4)
View | 959 |  | Phosphopropargyl | phospho-propargylamine | 116.998 | 117.0431 | H(4) C(3) N O(2) P
View | 488 | DHP | DHP | Dehydropyrrolizidine alkaloid (dehydroretronecine) on cysteines | 118.0657 | 118.1558 | H(8) C(8) N
View | 853 |  | Label:2H(4)+GlyGly | Ubiquitination 2H4 lysine | 118.068 | 118.1273 | H(2) 2H(4) C(4) N(2) O(2)
View | 312 | Cysteinyl | Cysteinyl | Cysteinylation | 119.0041 | 119.1423 | H(5) C(3) N O(2) S
View | 25 | Pyridylacetyl | pyridylacetyl | pyridylacetyl | 119.0371 | 119.1207 | H(5) C(7) N O
View | 411 | Phenylisocyanate | PIC | phenyl isocyanate | 119.0371 | 119.1207 | H(5) C(7) N O
View | 736 |  | DTT_C | Dithiothreitol (DTT) on Cys | 120.0245 | 120.1701 | H(8) C(4) O(2) S
View | 729 |  | O-Isopropylmethylphosphonate | O-Isopropylmethylphosphonylation | 120.034 | 120.0868 | H(9) C(4) O(2) P
View | 923 |  | Label:13C(4)15N(2)+GlyGly | 13C(4) 15N(2) Lysine glygly | 120.0504 | 120.0601 | H(6) 13C(4) 15N(2) O(2)
View | 799 |  | Label:13C(6)+GlyGly | 13C6 labeled ubiquitinylation residue | 120.0631 | 120.0586 | H(6) C(-2) 13C(6) N(2) O(2)
View | 264 | PET | PET | phosphorylation to pyridyl thiol | 121.035 | 121.2028 | H(7) C(7) N O(-1) S
View | 363 | Isopropylphospho | O-Isopropylphosphate | O-Isopropylphosphorylation | 122.0133 | 122.0596 | H(7) C(3) O(3) P
View | 864 |  | Label:13C(6)15N(2)+GlyGly | 13C(6) 15N(2) Lysine glygly | 122.0571 | 122.0454 | H(6) C(-2) 13C(6) 15N(2) O(2)
View | 456 | Xlink:DMP | DMP | Both ends of crosslink attached to same peptide | 122.0844 | 122.1677 | H(10) C(7) N(2)
View | 394 | GPIanchor | GPI-anchor | glycosylphosphatidylinositol | 123.0085 | 123.0477 | H(6) C(2) N O(3) P
View | 412 | Phenylisocyanate:2H(5) | d5-PIC | d5-phenyl isocyanate | 124.0685 | 124.1515 | 2H(5) C(7) N O
View | 108 | Nethylmaleimide | NEM | N-ethylmaleimide on cysteines | 125.0477 | 125.1253 | H(7) C(6) N O(2)
View | 129 | Iodo | Iodination | Iodination | 125.8966 | 125.8965 | H(-1) I
View | 764 |  | DTT_C:2H(6) | Isotopically labeled Dithiothreitol (DTT) modification of cysteines | 126.0622 | 126.2071 | H(2) 2H(6) C(4) O(2) S
View | 426 | Octanoyl | octanoyl | octanoyl | 126.1045 | 126.1962 | H(14) C(8) O
View | 29 | SMA | SMA | N-Succinimidyl-2-morpholine acetate | 127.0633 | 127.1412 | H(9) C(6) N O(2)
View | 60 | GIST-Quat | Quat_0 | Quaternary amine labeling reagent light form (N-term & K) | 127.0997 | 127.1842 | H(13) C(7) N O
View | 476 | TMAB | TMAB-d0 | 4-trimethyllammoniumbutyryl- | 128.1075 | 128.1922 | H(14) C(7) N O
View | 450 | Glu | -Glu- | monoglutamyl | 129.0426 | 129.114 | H(7) C(5) N O(3)
View | 500 | Nmethylmaleimide+water | NMMhyd | Nmethylmaleimidehydrolysis | 129.0426 | 129.114 | H(7) C(5) N O(3)
View | 573 |  | Gly->Trp | Gly->Trp substitution | 129.0578 | 129.1586 | H(7) C(9) N
View | 1253 |  | 2-monomethylsuccinyl | S-(2-monomethylsuccinyl) cysteine | 130.0266 | 130.0987 | H(6) C(5) O(4)
View | 776 |  | NEM:2H(5) | D5 N-ethylmaleimide on cysteines | 130.0791 | 130.1561 | H(2) 2H(5) C(6) N O(2)
View | 61 | GIST-Quat:2H(3) | Quat_3 | Quaternary amine labeling reagent heavy (+3amu) form, N-term & K | 130.1185 | 130.2027 | H(10) 2H(3) C(7) N O
View | 1022 |  | Xlink:DST | Monolink of DST crosslinker to Lys or N-terminus | 132.0059 | 132.0716 | H(4) C(4) O(5)
View | 186 | HPG | Arg1HPG | Hydroxyphenylglyoxal arginine | 132.0211 | 132.1162 | H(4) C(8) O(2)
View | 62 | GIST-Quat:2H(6) | Quat_6 | Quaternary amine labeling reagent heavy form (+6amu), N-term & K | 133.1374 | 133.2212 | H(7) 2H(6) C(7) N O
View | 501 | PyMIC | PyMIC | 3-methyl-2-pyridyl isocyanate | 134.048 | 134.1353 | H(6) C(7) N(2) O
View | 1032 |  | 2-nitrobenzyl | Tyrosine caged with 2-nitrobenzyl (ONB) | 135.032 | 135.1201 | H(5) C(7) N O(2)
View | 272 | CAF | CAF | sulfonation of N-terminus | 135.983 | 136.1265 | H(4) C(3) O(4) S
View | 735 |  | DTT_ST | Dithiothreitol (DTT) | 136.0017 | 136.2357 | H(8) C(4) O S(2)
View | 725 |  | O-Diethylphosphate | O-Diethylphosphorylation | 136.0289 | 136.0862 | H(9) C(4) O(3) P
View | 743 |  | 4-ONE+Delta:H(-2)O(-1) | Dehydrated 4-Oxononenal Michael adduct | 136.0888 | 136.191 | H(12) C(9) O
View | 63 | GIST-Quat:2H(9) | Quat_9 | Quaternary amine labeling reagent heavy form (+9amu), N-term & K | 136.1562 | 136.2397 | H(4) 2H(9) C(7) N O
View | 477 | TMAB:2H(9) | TMAB-d9 | d9-4-trimethyllammoniumbutyryl- | 137.164 | 137.2476 | H(5) 2H(9) C(7) N O
View | 1018 |  | ICDID | Isotope-Coded Dimedone light form | 138.0681 | 138.1638 | H(10) C(8) O(2)
View | 720 |  | HNE-Delta:H(2)O | Dehydrated 4-hydroxynonenal | 138.1045 | 138.2069 | H(14) C(9) O
View | 888 |  | mTRAQ | mTRAQ light | 140.095 | 140.183 | H(12) C(7) N(2) O
View | 1027 |  | Xlink:DMP-de | Monolink of DMP crosslinker to Lys or N-terminus | 140.095 | 140.183 | H(12) C(7) N(2) O
View | 525 |  | CLIP_TRAQ_2 | CLIP_TRAQ_2 | 141.0983 | 141.1756 | H(12) C(6) 13C N(2) O
View | 763 |  | DTT_ST:2H(6) | Isotopically labeled Dithiothreitol (DTT) modification of serines or threonines | 142.0393 | 142.2727 | H(2) 2H(6) C(4) O S(2)
View | 320 | Nethylmaleimide+water | NEMhyd | Nethylmaleimidehydrolysis | 143.0582 | 143.1406 | H(9) C(6) N O(3)
View | 375 | Diphthamide | Diphthamide | Diphthamide | 143.1184 | 143.2068 | H(15) C(7) N(2) O
View | 949 |  | 3-deoxyglucosone | Condensation product of 3-deoxyglucosone | 144.0423 | 144.1253 | H(8) C(6) O(4)
View | 1262 |  | 2-dimethylsuccinyl | S-(2-dimethylsuccinyl) cysteine | 144.0423 | 144.1253 | H(8) C(6) O(4)
View | 533 | iTRAQ4plex115 | iTRAQ115 | Accurate mass for 115 | 144.0996 | 144.1688 | H(12) C(6) 13C N 15N 18O
View | 214 | iTRAQ4plex | iTRAQ | Representative mass and accurate mass for 116 & 117 | 144.1021 | 144.1544 | H(12) C(4) 13C(3) N 15N O
View | 889 |  | mTRAQ:13C(3)15N(1) | mTRAQ heavy | 144.1021 | 144.1544 | H(12) C(4) 13C(3) N 15N O
View | 1019 |  | ICDID:2H(6) | Isotope-Coded Dimedone heavy form | 144.1057 | 144.2008 | H(4) 2H(6) C(8) O(2)
View | 532 | iTRAQ4plex114 | iTRAQ114 | Accurate mass for 114 | 144.1059 | 144.168 | H(12) C(5) 13C(2) N(2) 18O
View | 293 | CAMthiopropanoyl | CAMthiopropanoyl | 3-(carbamidomethylthio)propanoyl | 145.0197 | 145.1796 | H(7) C(5) N O(2) S
View | 407 | Hydroxycinnamyl | hydroxycinnamyl | hydroxycinnamyl | 146.0368 | 146.1427 | H(6) C(9) O(2)
View | 295 | dHex | Fuc | Fucose | 146.0579 | 146.1412 | dHex
View | 408 | Glycosyl | glycosyl | glycosyl-L-hydroxyproline | 148.0372 | 148.114 | H(8) C(5) O(5)
View | 978 |  | BITC | Benzyl isothiocyanate | 149.0299 | 149.2129 | H(7) C(8) N S
View | 851 |  | cGMP+RMP-loss | S-guanylation-2 | 150.0416 | 150.1182 | H(4) C(5) N(5) O
View | 172 | NBS | NBS-12C | Shimadzu NBS-12C | 152.9884 | 153.1585 | H(3) C(6) N O(2) S
View | 975 |  | Chlorpyrifos | O,O-diethyl o-3,5,6-trichloro-2-pyridyl phosphorothioate | 153.0139 | 153.1598 | H(10) C(4) O(2) P S
View | 419 | Glycerophospho | glycerophospho | glycerophospho | 154.0031 | 154.0584 | H(7) C(3) O(5) P
View | 721 |  | 4-ONE | 4-Oxononenal (ONE) | 154.0994 | 154.2063 | H(14) C(9) O(2)
View | 455 | Xlink:DMP-s | DMP-s | One end of crosslink attached, one end free | 154.1106 | 154.2096 | H(14) C(8) N(2) O
View | 449 | Decanoyl | decanoyl | lipid | 154.1358 | 154.2493 | H(18) C(10) O
View | 285 | SulfanilicAcid | SA_light | Light Sulfanilic Acid (SA) C12 | 155.0041 | 155.1744 | H(5) C(6) N O(2) S
View | 534 | Dibromo | Dibromo | Dibromo | 155.821 | 157.7921 | H(-2) Br(2)
View | 1020 |  | Xlink:DSS | Monolink of DSS/BS3 crosslinker to Lys or N-terminus | 156.0786 | 156.1791 | H(12) C(8) O(3)
View | 53 | HNE | HNE | 4-hydroxynonenal (HNE) | 156.115 | 156.2221 | H(16) C(9) O(2)
View | 335 | HNE+Delta:H(2) | redHNE | reduced 4-Hydroxynonenal | 158.1307 | 158.238 | H(18) C(9) O(2)
View | 171 | NBS:13C(6) | NBS-13C | Shimadzu NBS-13C | 159.0086 | 159.1144 | H(3) 13C(6) N O(2) S
View | 107 | FormylMet | +N-formyl-met | Addition of N-formyl met | 159.0354 | 159.2062 | H(9) C(6) N O(2) S
View | 967 |  | thioacylPA | membrane protein extraction | 159.0354 | 159.2062 | H(9) C(6) N O(2) S
View | 898 |  | pyrophospho | pyrophosphorylation of Ser/Thr | 159.9327 | 159.9598 | H(2) O(6) P(2)
View | 981 |  | glucosone | Condensation product of glucosone | 160.0372 | 160.1247 | H(8) C(6) O(5)
View | 286 | SulfanilicAcid:13C(6) | SA_heavy | Heavy Sulfanilic Acid (SA) C13 | 161.0242 | 161.1303 | H(5) 13C(6) N O(2) S
View | 454 | HexN | HexN | Hexosamine | 161.0688 | 161.1558 | H(11) C(6) N O(4)
View | 41 | Hex | Hex | Hexose | 162.0528 | 162.1406 | Hex
View | 727 |  | O-pinacolylmethylphosphonate | O-pinacolylmethylphosphonylation | 162.081 | 162.1666 | H(15) C(7) O(2) P
View | 979 |  | PEITC | Phenethyl isothiocyanate | 163.0456 | 163.2395 | H(9) C(9) N S
View | 362 | Diisopropylphosphate | O-Diisopropylphosphate | O-Diisopropylphosphorylation | 164.0602 | 164.1394 | H(13) C(6) O(3) P
View | 1255 |  | Cresylphosphate | o-toluyl-phosphorylation | 170.0133 | 170.1024 | H(7) C(7) O(3) P
View | 302 | Menadione | Menadione-Q | Menadione quinone derivative | 170.0368 | 170.1641 | H(6) C(11) O(2)
View | 194 | AccQTag | AccQTag | 6-aminoquinolyl-N-hydroxysuccinimidyl carbamate | 170.048 | 170.1674 | H(6) C(10) N(2) O
View | 195 | QAT | QAT | APTA-d0 | 171.1497 | 171.26 | H(19) C(9) N(2) O
View | 767 |  | Menadione-HQ | Menadione hydroquinone derivative | 172.0524 | 172.18 | H(8) C(11) O(2)
View | 1033 |  | SecNEM | N-ethylmaleimide on selenocysteines | 172.9921 | 172.0203 | H(7) C(6) N O(2) S(-1) Se
View | 1258 |  | Ub-VME | Ubiquitin vinylmethylester | 173.0926 | 173.1897 | H(13) C(7) N(2) O(3)
View | 1035 |  | Thiadiazole | Thiadiazolydation of Cys | 174.0252 | 174.2223 | H(6) C(9) N(2) S
View | 196 | QAT:2H(3) | QATd3 | APTA d3 | 174.1686 | 174.2784 | H(16) 2H(3) C(9) N(2) O
View | 457 | NDA | NDA | naphthalene-2,3-dicarboxaldehyde | 175.0422 | 175.1855 | H(5) C(13) N
View | 54 | Glucuronyl | N-glucuronyl | N-glucuronylation | 176.0321 | 176.1241 | H(8) C(6) O(6)
View | 1034 |  | SecNEM:2H(5) | D5 N-ethylmaleimide on selenocysteines | 178.0235 | 177.0511 | H(2) 2H(5) C(6) N O(2) S(-1) Se
View | 907 |  | Galactosyl | Galactosyl hydroxylysine | 178.0477 | 178.14 | O Hex
View | 276 | AEBS | AEBS | Aminoethylbenzenesulfonylation | 183.0354 | 183.2276 | H(9) C(8) N O(2) S
View | 748 |  | 3sulfo | derivatization by N-term modification using 3-Sulfobenzoic succinimidyl ester | 183.983 | 184.1693 | H(4) C(7) O(4) S
View | 911 |  | MTSL | Cys modification by (1-oxyl-2,2,5,5-tetramethyl-3-pyrroline-3-methyl)methanesulfonate (MTSL) | 184.0796 | 184.2786 | H(14) C(9) N O S
View | 197 | EQAT | EQAT | EAPTA d0 | 184.1576 | 184.2786 | H(20) C(10) N(2) O
View | 42 | Lipoyl | Lipoyl | Lipoyl | 188.033 | 188.3103 | H(12) C(8) O S(2)
View | 198 | EQAT:2H(5) | EQATd5 | EAPTA d5 | 189.1889 | 189.3094 | H(15) 2H(5) C(10) N(2) O
View | 301 | Bromobimane | Bromobimane | Monobromobimane derivative | 190.0742 | 190.1986 | H(10) C(10) N(2) O(2)
View | 1023 |  | Xlink:DTSSP | Monolink of DSP/DTSSP crosslinker to Lys or N-terminus | 191.9915 | 192.2559 | H(8) C(6) O(3) S(2)
View | 490 | Hep | Hep | Heptose | 192.0634 | 192.1666 | Hep
View | 1249 |  | AHA-SS | Azidohomoalanine coupled to reductively cleaved tag | 195.0756 | 195.1787 | H(9) C(7) N(5) O(2)
View | 1260 |  | Ub-amide | Ub amide probe addition | 196.1086 | 196.2264 | H(14) C(9) N(3) O(2)
View | 1031 |  | Biotin:Thermo-88310 | desthiobiotin modification of lysine | 196.1212 | 196.2462 | H(16) C(10) N(2) O(2)
View | 396 | GlycerylPE | glycerylPE | glycerylphosphorylethanolamine | 197.0453 | 197.1262 | H(12) C(5) N O(5) P
View | 941 |  | DNPS | 2,4-Dinitrobenzenesulfenyl | 198.9814 | 199.164 | H(3) C(6) N(2) O(4) S
View | 343 |  | Argbiotinhydrazide | oxidized Arginine biotinylated with biotin hydrazide | 199.0667 | 199.27 | H(13) C(9) N O(2) S
View | 291 | Delta:Hg(1) | Hg | Mercury Mercaptan | 201.9706 | 200.59 | Hg
View | 43 | HexNAc | HexNAc | N-Acetylhexosamine | 203.0794 | 203.1925 | HexNAc
View | 825 |  | DFDNB | Addition of DFDNB crosslinker | 203.9983 | 204.0879 | H(2) C(6) N(2) O(4) F(2)
View | 44 | Farnesyl | Farnesylation | Farnesylation | 204.1878 | 204.3511 | H(24) C(15)
View | 135 | Myristoyl+Delta:H(-4) | myristoyl-4H | (cis,cis-delta 5, delta 8)-tetradecadienoyl | 206.1671 | 206.3239 | H(22) C(14) O
View | 134 | Myristoleyl | myristoleylation | (cis-delta 5)-tetradecaenoyl | 208.1827 | 208.3398 | H(24) C(14) O
View | 893 |  | CarbamidomethylDTT | Carbamidomethylated DTT modification of cysteine | 209.018 | 209.2864 | H(11) C(6) N O(3) S(2)
View | 894 |  | CarboxymethylDTT | Carboxymethylated DTT modification of cysteine | 210.002 | 210.2712 | H(10) C(6) O(4) S(2)
View | 45 | Myristoyl | Myristoylation | Myristoylation | 210.1984 | 210.3556 | H(26) C(14) O
View | 751 |  | TNBS | tri nitro benzene | 210.9865 | 211.0886 | H C(6) N(3) O(6)
View | 261 | SPITC | SPITC | 4-sulfophenyl isothiocyanate | 214.9711 | 215.2495 | H(5) C(7) N O(3) S(2)
View | 762 |  | IDEnT | Isotope Distribution Encoded Tag | 214.9905 | 216.064 | H(7) C(9) N O Cl(2)
View | 176 | BHT | BHT | Michael addition of BHT quinone methide to Cysteine and Lysine | 218.1671 | 218.3346 | H(22) C(15) O
View | 824 |  | BMOE | Addition of BMOE crosslinker | 220.0484 | 220.1815 | H(8) C(10) N(2) O(4)
View | 376 | Hydroxyfarnesyl | hydroxyfarnesyl | hydroxyfarnesyl | 220.1827 | 220.3505 | H(24) C(15) O
View | 464 | SPITC:13C(6) | SPITC_Heavy | 4-sulfophenyl isothiocyanate (Heavy C13) | 220.9912 | 221.2054 | H(5) C 13C(6) N O(3) S(2)
View | 739 |  | TMT | Native Tandem Mass Tag® | 224.1525 | 224.2994 | H(20) C(12) N(2) O(2)
View | 89 | Iminobiotin | Im_biotin | Iminobiotinylation | 225.0936 | 225.3106 | H(15) C(10) N(3) O S
View | 738 |  | TMT2plex | Duplex Tandem Mass Tag® | 225.1558 | 225.2921 | H(20) C(11) 13C N(2) O(2)
View | 3 | Biotin | Biotin | Biotinylation | 226.0776 | 226.2954 | H(14) C(10) N(2) O(2) S
View | 105 | ICAT-C | ICAT_light | Applied Biosystems cleavable ICAT(TM) light | 227.127 | 227.2603 | H(17) C(10) N(3) O(3)
View | 513 | C8-QAT | C8-QAT | [3-(2,5)-Dioxopyrrolidin-1-yloxycarbonyl)-propyl]dimethyloctylammonium | 227.2249 | 227.3862 | H(29) C(14) N O
View | 910 |  | Bacillosamine | 2,4-diacetamido-2,4,6-trideoxyglucopyranose | 228.111 | 228.245 | H(16) C(10) N(2) O(4)
View | 46 | PyridoxalPhosphate | Pyridoxal-phos | Pyridoxal phosphate | 229.014 | 229.1266 | H(8) C(8) N O(5) P
View | 737 |  | TMT6plex | Sixplex Tandem Mass Tag® | 229.1629 | 229.2634 | H(20) C(8) 13C(4) N 15N O(2)
View | 514 | PropylNAGthiazoline | propyl-NAG-thiazoline | propyl-1,2-dideoxy-2\'-methyl-alpha-D-glucopyranoso-[2,1-d]-Delta2\'-thiazoline | 232.0644 | 232.2768 | H(14) C(9) N O(4) S
View | 139 | Dansyl | dansyl | 5-dimethylaminonaphthalene-1-sulfonyl | 233.051 | 233.2862 | H(11) C(12) N O(2) S
View | 498 | BHTOH | BHTOH | Michael addition of t-butyl hydroxylated BHT (BHTOH) to C, H or K | 234.162 | 234.334 | H(22) C(15) O(2)
View | 106 | ICAT-C:13C(9) | ICAT_heavy | Applied Biosystems cleavable ICAT(TM) heavy | 236.1572 | 236.1942 | H(17) C 13C(9) N(3) O(3)
View | 431 | Palmitoleyl | palmitoleyl | palmitoleyl | 236.214 | 236.3929 | H(28) C(16) O
View | 1024 |  | Xlink:SMCC | Monolink of sulfoSMCC/SMCC crosslinker to Cys | 237.1001 | 237.2518 | H(15) C(12) N O(4)
View | 47 | Palmitoyl | Palmitoylation | Palmitoylation | 238.2297 | 238.4088 | H(30) C(16) O
View | 361 |  | Thrbiotinhydrazide | oxidized Threonine biotinylated with biotin hydrazide | 240.1045 | 240.3252 | H(16) C(10) N(4) O S
View | 771 |  | lapachenole | lapachenole photochemically added to cysteine | 240.115 | 240.297 | H(16) C(16) O(2)
View | 353 |  | Lysbiotinhydrazide | oxidized Lysine biotinylated with biotin hydrazide | 241.0885 | 241.31 | H(15) C(10) N(3) O(2) S
View | 429 | PhosphoHex | p-Man | phosphoglycosyl-D-mannose-1-phosphoryl | 242.0192 | 242.1205 | H O(3) P Hex
View | 1264 |  | pupylation | addition of GGE | 243.0855 | 243.2166 | H(13) C(9) N(3) O(5)
View | 1021 |  | Xlink:EGS | Monolink of EGS crosslinker to Lys or N-terminus | 244.0583 | 244.1981 | H(12) C(10) O(7)
View | 537 |  | CLIP_TRAQ_4 | CLIP_TRAQ_4 | 244.1015 | 244.2292 | H(15) C(9) 13C N(2) O(5)
View | 130 | Diiodo | di-Iodination | di-Iodination | 251.7933 | 251.7931 | H(-2) I(2)
View | 744 |  | NO_SMX_SEMD | Nitroso Sulfamethoxazole Sulphenamide thiol adduct | 252.0443 | 252.2697 | H(10) C(10) N(3) O(3) S
View | 1250 |  | AHA-SS_CAM | carbamidomethylated form of reductively cleaved tag coupled to azidohomoalanine | 252.0971 | 252.23 | H(12) C(9) N(6) O(3)
View | 684 |  | BDMAPP | Mass Defect Tag on lysine e-amino | 253.0102 | 254.1231 | H(12) C(11) N O Br
View | 273 | Xlink:SSD | SSD_K | covalent modification of lysine by cross-linking reagent | 253.095 | 253.2512 | H(15) C(12) N O(5)
View | 451 | GluGlu | -Glu-Glu- | diglutamyl | 258.0852 | 258.228 | H(14) C(10) N(2) O(6)
View | 357 |  | probiotinhydrazide | oxidized proline biotinylated with biotin hydrazide | 258.115 | 258.3405 | H(18) C(10) N(4) O(2) S
View | 433 | Didehydroretinylidene | didehydroretinyl | 3,4-didehydroretinylidene | 264.1878 | 264.4046 | H(24) C(20)
View | 848 |  | Arg2PG | Adduct of phenylglyoxal with Arg | 266.0579 | 266.2482 | H(10) C(16) O(4)
View | 380 | Retinylidene | retinal | retinal | 266.2035 | 266.4204 | H(26) C(20)
View | 746 |  | NO_SMX_SIMD | Nitroso Sulfamethoxazole Sulfinamide thiol adduct | 267.0314 | 267.2612 | H(9) C(10) N(3) O(4) S
View | 943 |  | DimethylamineGMBS | Modified GMBS X linker for proteins | 267.1583 | 267.3241 | H(21) C(13) N(3) O(3)
View | 745 |  | NO_SMX_SMCT | Nitroso Sulfamethoxazole semimercaptal thiol adduct | 268.0392 | 268.2691 | H(10) C(10) N(3) O(4) S
View | 536 |  | CLIP_TRAQ_3 | CLIP_TRAQ_3 | 271.1487 | 271.2976 | H(20) C(11) 13C N(3) O(4)
View | 48 | GeranylGeranyl | Geranyl-geranyl | Geranyl-geranyl | 272.2504 | 272.4681 | H(32) C(20)
View | 1256 |  | CresylSaligeninPhosphate | Cresyl-Saligenin-phosphorylation | 276.0551 | 276.2244 | H(13) C(14) O(4) P
View | 187 | 2HPG | Arg2HPG | bis(hydroxphenylglyoxal) arginine | 282.0528 | 282.2476 | H(10) C(16) O(5)
View | 428 | PhosphoHexNAc | p-GlcNAc | N-acetylglucosamine-1-phosphoryl | 283.0457 | 283.1724 | H O(3) P HexNAc
View | 1003 |  | EGCG2 | (-)-dehydroepigallocatechin | 287.0556 | 287.2442 | H(11) C(15) O(6)
View | 505 |  | LG-lactam-R | Levuglandinyl - arginine lactam adduct | 290.177 | 290.3939 | H(26) C(19) N(-2) O(4)
View | 434 | CHDH | LTP1-lipid | cis-14-hydroxy-10,13-dioxo-7-heptadecenoic ester | 294.1831 | 294.3859 | H(26) C(17) O(4)
View | 243 | IGBP | IGBP | Light IDBEST tag for quantitation | 296.016 | 297.1478 | H(13) C(12) N(2) O(2) Br
View | 499 | IGBP:13C(2) | IGBP:13C(2) | Heavy IDBEST tag for quantitation | 298.0227 | 299.1331 | H(13) C(10) 13C(2) N(2) O(2) Br
View | 948 |  | LG-anhyropyrrole | Levuglandinyl-lysine anhyropyrrole adduct | 298.1933 | 298.4192 | H(26) C(20) O(2)
View | 984 |  | cysTMT | Native cysteine-reactive Tandem Mass Tag® | 299.1667 | 299.4322 | H(25) C(14) N(3) O(2) S
View | 985 |  | cysTMT6plex | cysteine-reactive Sixplex Tandem Mass Tag® | 304.1772 | 304.3962 | H(25) C(10) 13C(4) N(2) 15N O(2) S
View | 731 |  | iTRAQ8plex:13C(6)15N(2) | Accurate mass for 115, 118, 119 & 121 | 304.199 | 304.3081 | H(24) C(8) 13C(6) N(2) 15N(2) O(3)
View | 730 |  | iTRAQ8plex | Representative mass and accurate mass for 113, 114, 116 & 117 | 304.2054 | 304.3074 | H(24) C(7) 13C(7) N(3) 15N O(3)
View | 55 | Glutathione | Glutathione | glutathione disulfide | 305.0682 | 305.3076 | H(15) C(10) N(3) O(6) S
View | 417 | PhosphoUridine | p-uridine | uridine phosphodiester | 306.0253 | 306.166 | H(11) C(9) N(2) O(8) P
View | 506 |  | LG-Hlactam-R | Levuglandinyl - arginine hydroxylactam adduct | 306.1719 | 306.3933 | H(26) C(19) N(-2) O(5)
View | 1015 |  | Ahx2+Hsl | C-terminal homoserine lactone and two aminohexanoic acids | 309.2052 | 309.4039 | H(27) C(16) N(3) O(3)
View | 116 |  | OxArgBiotin | Oxidized arginine biotinylated with biotin-LC-hydrazide | 310.1351 | 310.4118 | H(22) C(15) N(2) O(3) S
View | 800 |  | Biotin:Thermo-21345 | was PentylamineBiotin | 311.1667 | 311.4429 | H(25) C(15) N(3) O(2) S
View | 117 |  | OxArgBiotinRed | Oxidized arginine biotinylated with biotin-LC-hydrazide, reduced | 312.1508 | 312.4277 | H(24) C(15) N(2) O(3) S
View | 946 |  | LG-anhydrolactam | Levuglandinyl-lysine anhydrolactam adduct | 314.1882 | 314.4186 | H(26) C(20) O(3)
View | 119 | IBTP | IBTP | Thio Ether Formation - BTP Adduct | 316.1381 | 316.3759 | H(21) C(22) P
View | 947 |  | LG-pyrrole | Levuglandinyl-lysine pyrrole adduct | 316.2038 | 316.4345 | H(28) C(20) O(3)
View | 908 |  | SMCC-maleimide | Modified SMCC maleimide with 3-(dimethylamino)-1-propylamine | 321.2052 | 321.4146 | H(27) C(17) N(3) O(3)
View | 292 | IodoU-AMP | IodoU-AMP | Cross-link of (Iodo)-uracil MP with W,F,Y | 322.0202 | 322.1654 | H(11) C(9) N(2) O(9) P
View | 512 | Hex(2) | Gal-Glu | Lactosylation | 324.1056 | 324.2812 | Hex(2)
View | 685 |  | NA-LNO2 | Nitroalkylation by Nitro Linoleic Acid | 325.2253 | 325.443 | H(31) C(18) N O(4)
View | 294 | IED-Biotin | ied-biotin | biotinoyl-iodoacetyl-ethylenediamine | 326.1413 | 326.4145 | H(22) C(14) N(4) O(3) S
View | 686 |  | NA-OA-NO2 | Nitroalkylation by Nitro Oleic Acid | 327.241 | 327.4589 | H(33) C(18) N O(4)
View | 405 | Phosphoadenosine | p-adenosine | AMP binding site | 329.0525 | 329.2059 | H(12) C(10) N(5) O(6) P
View | 503 |  | LG-lactam-K | Levuglandinyl - lysine lactam adduct | 332.1988 | 332.4339 | H(28) C(20) O(4)
View | 90 | ESP | ESP-Tag_light | ESP-Tag light d0 | 338.1776 | 338.4682 | H(26) C(16) N(4) O(2) S
View | 92 | NHS-LC-Biotin | Biot_LC | NHS-LC-Biotin | 339.1617 | 339.453 | H(25) C(16) N(3) O(3) S
View | 998 |  | BHAc | N-biotinyl-6-aminohexanoyl | 339.1617 | 339.453 | H(25) C(16) N(3) O(3) S
View | 49 | Phosphopantetheine | p-pantetheine | Phosphopantetheine | 340.0858 | 340.333 | H(21) C(11) N(2) O(6) P S
View | 393 | Glucosylgalactosyl | glucosylgalactosyl | glucosylgalactosyl hydroxylysine | 340.1006 | 340.2806 | O Hex(2)
View | 493 | BADGE | BADGE | Bisphenol A diglycidyl ether derivative | 340.1675 | 340.4129 | H(24) C(21) O(4)
View | 439 | Diironsubcluster | Fe-cluster | hydrogenase diiron subcluster | 342.7869 | 342.876 | H(-1) C(5) N(2) O(5) S(2) Fe(2)
View | 849 |  | cGMP | S-guanylation | 344.0396 | 344.1974 | H(11) C(10) N(5) O(7) P
View | 413 | Phosphoguanosine | p-guanosine | phospho-guanosine | 345.0474 | 345.2053 | H(12) C(10) N(5) O(7) P
View | 123 | ICAT-H | Harbury_ICAT_C12 | N-iodoacetyl, p-chlorobenzyl-12C6-glucamine | 345.0979 | 345.7754 | H(20) C(15) N O(6) Cl
View | 504 |  | LG-Hlactam-K | Levuglandinyl - lysine hydroxylactam adduct | 348.1937 | 348.4333 | H(28) C(20) O(5)
View | 91 | ESP:2H(10) | ESP-Tag_heavy | ESP-Tag heavy d10 | 348.2404 | 348.5299 | H(16) 2H(10) C(16) N(4) O(2) S
View | 142 | HexNAc(1)dHex(1) | HexNAc1dHex1 | HexNAc1dHex1 | 349.1373 | 349.3337 | dHex HexNAc
View | 124 | ICAT-H:13C(6) | Harbury_ICAT_C13 | N-iodoacetyl, p-chlorobenzyl-13C6-glucamine | 351.118 | 351.7313 | H(20) C(9) 13C(6) N O(6) Cl
View | 113 |  | OxLysBiotin | Oxidized lysine biotinylated with biotin-LC-hydrazide | 352.1569 | 352.4518 | H(24) C(16) N(4) O(3) S
View | 112 |  | OxLysBiotinRed | Oxidized lysine biotinylated with biotin-LC-hydrazide, reduced | 354.1726 | 354.4676 | H(26) C(16) N(4) O(3) S
View | 289 | Biotin-PEO-Amine | BiotinPEO-amine | Biotin polyethyleneoxide amine | 356.1882 | 356.4835 | H(28) C(16) N(4) O(3) S
View | 270 | Cytopiloyne | cytopiloyne | nucleophilic addtion to cytopiloyne | 362.1366 | 362.3738 | H(22) C(19) O(7)
View | 793 |  | Hex1HexNAc1 | Hex1HexNAc1 | 365.1322 | 365.3331 | Hex HexNAc
View | 432 | Cholesterol | C-cholesterol | cholesterol ester | 368.3443 | 368.6383 | H(44) C(27)
View | 115 |  | OxProBiotin | Oxidized Proline biotinylated with biotin-LC-hydrazide | 369.1835 | 369.4823 | H(27) C(16) N(5) O(3) S
View | 114 |  | OxProBiotinRed | Oxidized proline biotinylated with biotin-LC-hydrazide, reduced | 371.1991 | 371.4982 | H(29) C(16) N(5) O(3) S
View | 901 |  | 4AcAllylGal | 2,3,4,6-tetra-O-Acetyl-1-allyl-α-D-galactopyranoside modification of cysteine | 372.142 | 372.3671 | H(24) C(17) O(9)
View | 131 | Triiodo | tri-Iodination | tri-Iodination | 377.6899 | 377.6896 | H(-3) I(3)
View | 271 | Cytopiloyne+water | cytopiloyne+H2O | nucleophilic addition to cytopiloyne+H2O | 380.1471 | 380.3891 | H(24) C(19) O(8)
View | 887 |  | MDCC | covalent linkage of maleimidyl coumarin probe (Molecular Probes D-10253) | 383.1481 | 383.3978 | H(21) C(20) N(3) O(5)
View | 535 | LeuArgGlyGly | LeuArgGlyGly | Ubiquitination | 383.2281 | 383.446 | H(29) C(16) N(7) O(4)
View | 437 | MicrocinC7 | C-Asn-deriv | (3-aminopropyl)(L-aspartyl-1-amino)phosphoryl-5-adenosine | 386.1104 | 386.3003 | H(19) C(13) N(6) O(6) P
View | 452 | GluGluGlu | -Glu-Glu-Glu- | triglutamyl | 387.1278 | 387.3419 | H(21) C(15) N(3) O(9)
View | 128 | Fluorescein | Fluorescein | 5-Iodoacetamidofluorescein (Molecular Probe, Eugene, OR) | 388.0821 | 388.3497 | H(14) C(22) N O(6)
View | 143 | HexNAc(2) | HexNAc2 | HexNAc2 | 406.1587 | 406.385 | HexNAc(2)
View | 878 |  | Bodipy | Bodipy modifications onto cysteine | 414.1675 | 414.2135 | H(21) B C(20) N(4) O(3) F(2)
View | 20 | PEO-Iodoacetyl-LC-Biotin | PEO-Biotin | Biotinyl-iodoacetamidyl-3,6-dioxaoctanediamine | 414.1937 | 414.5196 | H(30) C(18) N(4) O(5) S
View | 416 | Dipyrrolylmethanemethyl | dipyrrole | dipyrrolylmethanemethyl | 418.1376 | 418.3973 | H(22) C(20) N(2) O(8)
View | 478 | FTC | FTC | fluorescein-5-thiosemicarbazide | 421.0732 | 421.4259 | H(15) C(21) N(3) O(5) S
View | 515 | FNEM | F-NEM | fluorescein-5-maleimide | 427.0692 | 427.3625 | H(13) C(24) N O(7)
View | 290 | Biotin-HPDP | Biotin:Thermo-21341 | Pierce EZ-Link Biotin-HPDP | 428.1916 | 428.6124 | H(32) C(19) N(4) O(3) S(2)
View | 442 | FMN | FMN2 | O3-(riboflavin phosphoryl) | 438.0941 | 438.3285 | H(19) C(17) N(4) O(8) P
View | 13 | ICAT-D | AB_old_ICATd0 | Applied Biosystems original ICAT(TM) d0 | 442.225 | 442.5728 | H(34) C(20) N(4) O(5) S
View | 1037 |  | Biotin:Thermo-88317 | desthiobiotin fluorophosphonate | 443.2913 | 443.5603 | H(42) C(22) N(3) O(4) P
View | 333 | Can-FP-biotin | Biotin:TRC-B394885 | 6-N-biotinylaminohexyl isopropyl phosphate | 447.1957 | 447.5291 | H(34) C(19) N(3) O(5) P S
View | 993 |  | Biotin:Sigma-B1267 | was Biotin-maleimide | 449.1733 | 449.5239 | H(27) C(20) N(5) O(5) S
View | 12 | ICAT-D:2H(8) | AB_old_ICATd8 | Applied Biosystems original ICAT(TM) d8 | 450.2752 | 450.6221 | H(26) 2H(8) C(20) N(4) O(5) S
View | 523 | Sulfo-NHS-LC-LC-Biotin | Biotin:Thermo-21338 | Biot_LC_LC | 452.2457 | 452.6106 | H(36) C(22) N(4) O(4) S
View | 973 |  | Puromycin | Puromycin | 453.2125 | 453.4943 | H(27) C(22) N(7) O(4)
View | 409 | FMNH | FMN | flavin mononucleotide | 454.089 | 454.3279 | H(19) C(17) N(4) O(9) P
View | 1002 |  | EGCG1 | (-)-epigallocatechin-3-gallate | 456.0693 | 456.3558 | H(16) C(22) O(11)
View | 443 | FMNC | FMN3 | S-(4a-FMN) | 456.1046 | 456.3438 | H(21) C(17) N(4) O(9) P
View | 942 |  | SulfoGMBS | High molecular absorption label for proteins | 458.1624 | 458.5306 | H(26) C(22) N(4) O(5) S
View | 397 | Triiodothyronine | triiodo | triiodo | 469.7162 | 469.785 | H C(6) O I(3)
View | 1036 |  | Withaferin | Modification of cystein by withaferin | 470.2668 | 470.5977 | H(38) C(28) O(6)
View | 935 |  | Atto495Maleimide | High molecular absorption maleimide label for proteins | 474.2505 | 474.5747 | H(32) C(27) N(5) O(3)
View | 846 |  | EQIGG | Sumo mutant Smt3-WT tail following trypsin digestion | 484.2282 | 484.5035 | H(32) C(20) N(6) O(8)
View | 144 | Hex(3) | Hex3 | Hex3 | 486.1585 | 486.4218 | Hex(3)
View | 8 | ICAT-G | Gygi_ICATd0 | Gygi ICAT(TM) d0 | 486.2512 | 486.6253 | H(38) C(22) N(4) O(6) S
View | 811 |  | Biotin:Thermo-21360 | was Biotin-PEO4-hydrazide | 487.2465 | 487.6134 | H(37) C(21) N(5) O(6) S
View | 118 | EDT-iodoacetyl-PEO-biotin | EDT-i-biotin | EDT-iodo-PEO-biotin | 490.1742 | 490.7034 | H(34) C(20) N(4) O(4) S(3)
View | 9 | ICAT-G:2H(8) | Gygi_ICATd8 | Gygi ICAT(TM) d8 | 494.3014 | 494.6746 | H(30) 2H(8) C(22) N(4) O(6) S
View | 145 | HexNAc(1)dHex(2) | HexNAc1dHex2 | HexNAc1dHex2 | 495.1952 | 495.4749 | dHex(2) HexNAc
View | 934 |  | AMTzHexNAc2 | Photocleavable Biotin + GalNAz on O-GlcNAc | 502.2023 | 502.4757 | H(30) C(19) N(6) O(10)
View | 146 | Hex(1)HexNAc(1)dHex(1) | Hex1HexNAc1dHex1 | Hex1HexNAc1dHex1 | 511.1901 | 511.4743 | dHex Hex HexNAc
View | 912 |  | HNE-BAHAH | 4-hydroxy-2-nonenal and biotinamidohexanoic acid hydrazide, reduced | 511.3192 | 511.7209 | H(45) C(25) N(5) O(4) S
View | 453 | GluGluGluGlu | -Glu-Glu-Glu-Glu- | tetraglutamyl | 516.1704 | 516.4559 | H(28) C(20) N(4) O(12)
View | 391 | Molybdopterin | molybdopterin | molybdopterin | 521.8841 | 520.2668 | H(11) C(10) N(5) O(8) P S(2) Mo
View | 1012 |  | Biotin:Invitrogen-M1602 | Nalpha-(3-maleimidylpropionyl)biocytin | 523.2101 | 523.6024 | H(33) C(23) N(5) O(7) S
View | 332 | Thiophos-S-S-biotin | Thiophos-S-S-biotin | thiophosphate labeled with biotin-HPDP | 525.1429 | 525.6658 | H(34) C(19) N(4) O(5) P S(3)
View | 522 | Maleimide-PEO2-Biotin | Biotin:Thermo-21901 | Maleimide-Biotin | 525.2257 | 525.6183 | H(35) C(23) N(5) O(7) S
View | 213 | ADP-Ribosyl | ADP-Ribose | ADP Ribose addition | 541.0611 | 541.3005 | H(21) C(15) N(5) O(13) P(2)
View | 1039 |  | Biotin:Thermo-21901+H2O | Maleimide-Biotin + Water | 543.2363 | 543.6336 | H(37) C(23) N(5) O(8) S
View | 1252 |  | Biotin:Thermo-33033-H | Sulfo-SBED Label Photoreactive Biotin Crosslinker minus Hydrogen | 546.2083 | 546.7053 | H(34) C(25) N(6) O(4) S(2)
View | 1251 |  | Biotin:Thermo-33033 | Sulfo-SBED Label Photoreactive Biotin Crosslinker | 547.2161 | 547.7132 | H(35) C(25) N(6) O(4) S(2)
View | 147 | HexNAc(2)dHex(1) | HexNAc2dHex1 | HexNAc2dHex1 | 552.2167 | 552.5262 | dHex HexNAc(2)
View | 148 | Hex(1)HexNAc(2) | Hex1HexNAc2 | Hex1HexNAc2 | 568.2116 | 568.5256 | Hex HexNAc(2)
View | 827 |  | TMPP-Ac | tris(2,4,6-trimethoxyphenyl)phosphonium acetic acid N-hydroxysuccinimide ester derivative | 572.1811 | 572.5401 | H(33) C(29) O(10) P
View | 325 | FP-Biotin | FP-Biotin | 10-ethoxyphosphinyl-N-(biotinamidopentyl)decanamide | 572.3161 | 572.7405 | H(49) C(27) N(4) O(5) P S
View | 377 | Diacylglycerol | diacylglycerol | diacylglycerol | 576.5118 | 576.9334 | H(68) C(37) O(4)
View | 895 |  | Biotin-PEG-PRA | Biotin polyethyleneoxide (n=3) alkyne | 578.3176 | 578.6611 | H(42) C(26) N(8) O(7)
View | 389 | Phytochromobilin | phytochromobilin | phytochromobilin | 584.2635 | 584.6621 | H(36) C(33) N(4) O(6)
View | 387 | Phycocyanobilin | phycocyanobilin | phycocyanobilin | 586.2791 | 586.678 | H(38) C(33) N(4) O(6)
View | 388 | Phycoerythrobilin | phycoerythrobilin | phycoerythrobilin | 588.2948 | 588.6939 | H(40) C(33) N(4) O(6)
View | 519 | BisANS | BisANS | 4,4\'-dianilino-1,1\'-binaphthyl-5,5\'-disulfonic acid | 592.0763 | 592.641 | H(20) C(32) N(2) O(6) S(2)
View | 398 | Thyroxine | tetraiodo | tetraiodo | 595.6128 | 595.6815 | C(6) O I(4)
View | 1261 |  | Ub-fluorescein | Ub Fluorescein probe addition | 597.2098 | 597.598 | H(29) C(31) N(6) O(7)
View | 877 |  | QQQTGG | SUMOylation by SUMO-2/3 | 599.2663 | 599.5942 | H(37) C(23) N(9) O(10)
View | 876 |  | QEQTGG | SUMOylation by SUMO-1 | 600.2504 | 600.5789 | H(36) C(23) N(8) O(11)
View | 93 | EDT-maleimide-PEO-biotin | EDT-m-biotin | EDT-maleimide-PEO-biotin | 601.2062 | 601.8021 | H(39) C(25) N(5) O(6) S(3)
View | 436 | Hydroxyheme | hydroxyheme | hydroxyheme | 614.1616 | 614.4714 | H(30) C(34) N(4) O(4) Fe
View | 390 | Heme | heme | heme | 616.1773 | 616.4873 | H(32) C(34) N(4) O(4) Fe
View | 774 |  | Biotin-phenacyl | Alkylation by biotinylated form of phenacyl bromide | 626.2635 | 626.727 | H(38) C(29) N(8) O(6) S
View | 538 |  | Biotin:Cayman-10141 | was 15dB-biotin | 626.3866 | 626.8927 | H(54) C(35) N(4) O(4) S
View | 410 | Archaeol | S-archaeol | S-diphytanylglycerol diether | 634.6628 | 635.1417 | H(86) C(43) O(2)
View | 149 | Hex(1)HexNAc(1)NeuAc(1) | Hex1HexNAc1NeuAc1 | Hex1HexNAc1NeuAc1 | 656.2276 | 656.5877 | Hex HexNAc NeuAc
View | 1038 |  | TAMRA-FP | TAMRA fluorophosphonate modification of serine | 659.3124 | 659.7514 | H(46) C(37) N(3) O(6) P
View | 539 |  | Biotin:Cayman-10013 | was PGA1-biotin | 660.4284 | 660.9504 | H(60) C(36) N(4) O(5) S
View | 494 | CyDye-Cy3 | CyDye-Cy3 | Cy3 CyDye DIGE Fluor saturation dye | 672.2982 | 672.8335 | H(44) C(37) N(4) O(6) S
View | 495 | CyDye-Cy5 | CyDye-Cy5 | Cy5 CyDye DIGE Fluor saturation dye | 684.2982 | 684.8442 | H(44) C(38) N(4) O(6) S
View | 1001 |  | AHA-Alkyne-KDDDD | Azidohomoalanine (AHA) bound to DDDDK-propargylglycine-NH2 (alkyne) | 695.2801 | 695.5723 | H(37) C(26) N(11) O(14) S(-1)
View | 884 |  | Biotin:Thermo-21325 | was ChromoBiotin | 695.3101 | 695.8288 | H(45) C(34) N(7) O(7) S
View | 150 | HexNAc(2)dHex(2) | HexNAc2dHex2 | HexNAc2dHex2 | 698.2746 | 698.6674 | dHex(2) HexNAc(2)
View | 151 | Hex(1)HexNAc(2)Pent(1) | Hex1HexNAc2Pent1 | Hex1HexNAc2Pent1 | 700.2538 | 700.6403 | Pent Hex HexNAc(2)
View | 891 |  | Methyl-PEO12-Maleimide | Methyl-PEO12-Maleimide | 710.3837 | 710.8073 | H(58) C(32) N(2) O(15)
View | 323 | Xlink:B10621 | bNis | bis-N-I-sulfonerahodamine | 713.0931 | 713.5626 | H(30) C(31) N(4) O(6) S I
View | 152 | Hex(1)HexNAc(2)dHex(1) | Hex1HexNAc2dHex1 | Hex1HexNAc2dHex1 | 714.2695 | 714.6668 | dHex Hex HexNAc(2)
View | 153 | Hex(2)HexNAc(2) | Hex2HexNAc2 | Hex2HexNAc2 | 730.2644 | 730.6662 | Hex(2) HexNAc(2)
View | 861 |  | ZGB | NHS ester linked Green Fluorescent Bodipy Dye | 758.3808 | 758.7261 | H(53) B C(37) N(6) O(6) F(2) S
View | 281 | CoenzymeA | CysCoA | Cysteine modified Coenzyme A | 765.0996 | 765.5182 | H(34) C(21) N(7) O(16) P(3) S
View | 50 | FAD | FAD | Flavin adenine dinucleotide | 783.1415 | 783.5339 | H(31) C(27) N(9) O(15) P(2)
View | 51 | Tripalmitate | Tripalmitate | N-acyl diglyceride cysteine | 788.7258 | 789.3049 | H(96) C(51) O(5)
View | 821 |  | Cy3b-maleimide | fluorescent dye that labels cysteines | 796.239 | 796.8086 | H(39) C(39) N(4) O(9) F(3) S
View | 938 |  | AROD | Cysteine modifier | 820.336 | 820.979 | H(52) C(35) N(10) O(9) S(2)
View | 154 | Hex(3)HexNAc(1)Pent(1) | Hex3HexNAc1Pent1 | Hex3HexNAc1Pent1 | 821.2801 | 821.7289 | Pent Hex(3) HexNAc
View | 448 | Microcin | microcin | microcin E492 siderophore ester from serine | 831.197 | 831.6871 | H(37) C(36) N(3) O(20)
View | 932 |  | VFQQQTGG | SUMOylation by SUMO-2/3 (formic acid cleavage) | 845.4032 | 845.8991 | H(55) C(37) N(11) O(12)
View | 155 | Hex(1)HexNAc(2)dHex(1)Pent(1) | Hex1HexNAc2dHex1Pent1 | Hex1HexNAc2dHex1Pent1 | 846.3117 | 846.7815 | Pent dHex Hex HexNAc(2)
View | 156 | Hex(1)HexNAc(2)dHex(2) | Hex1HexNAc2dHex2 | Hex1HexNAc2dHex2 | 860.3274 | 860.808 | dHex(2) Hex HexNAc(2)
View | 157 | Hex(2)HexNAc(2)Pent(1) | Hex2HexNAc2Pent1 | Hex2HexNAc2Pent1 | 862.3067 | 862.7809 | Pent Hex(2) HexNAc(2)
View | 158 | Hex(2)HexNAc(2)dHex(1) | Hex2HexNAc2dHex1 | Hex2HexNAc2dHex1 | 876.3223 | 876.8074 | dHex Hex(2) HexNAc(2)
View | 395 | PhosphoribosyldephosphoCoA | pRibodePcoA | phosphoribosyl dephospho-coenzyme A | 881.1469 | 881.6335 | H(42) C(26) N(7) O(19) P(3) S
View | 159 | Hex(3)HexNAc(2) | Hex3HexNAc2 | Hex3HexNAc2 | 892.3172 | 892.8068 | Hex(3) HexNAc(2)
View | 444 | CuSMo | CuSMo | copper sulfido molybdopterin cytosine dinuncleotide | 922.8349 | 922.067 | H(24) C(19) N(8) O(15) P(2) S(3) Cu Mo
View | 161 | Hex(3)HexNAc(2)P(1) | Hex3HexNAc2P1 | Hex3HexNAc2P1 | 923.291 | 923.7806 | P Hex(3) HexNAc(2)
View | 890 |  | DyLight-maleimide | Thiol-reactive dye for fluorescence labelling of proteins | 940.1999 | 941.0762 | H(48) C(39) N(4) O(15) S(4)
View | 160 | Hex(1)HexNAc(1)NeuAc(2) | Hex1HexNAc1NeuAc2 | Hex1HexNAc1NeuAc2 | 947.323 | 947.8423 | Hex HexNAc NeuAc(2)
View | 971 |  | maleimide3 | maleimide-3-saccharide | 969.3662 | 969.8975 | H(59) C(37) N(7) O(23)
View | 740 |  | ExacTagThiol | ExacTag Thiol label mass for 2-4-7-10 plex | 972.3652 | 972.7268 | H(50) C(23) 13C(12) N(8) 15N(6) O(18)
View | 741 |  | ExacTagAmine | ExacTag Amine label mass for 2-4-7-10 plex | 1046.3479 | 1046.8285 | H(52) C(25) 13C(12) N(8) 15N(6) O(19) S
View | 933 |  | VIEVYQEQTGG | SUMOylation by SUMO-1 (formic acid cleavage) | 1203.5772 | 1204.2859 | H(81) C(53) N(13) O(19)
View | 137 | Hex(5)HexNAc(2) | N-glycan | N-linked glycan core | 1216.4229 | 1217.088 | Hex(5) HexNAc(2)
View | 972 |  | maleimide5 | maleimide-5-saccharide | 1293.4719 | 1294.1787 | H(79) C(49) N(7) O(33)
View | 309 | Hex(3)HexNAc(4) | G0-Fuc | Biantennary (-2 galactose) | 1298.476 | 1299.1919 | Hex(3) HexNAc(4)
View | 305 | dHex(1)Hex(3)HexNAc(4) | G0 | Fucosylated biantennary (-2 galactose) | 1444.5339 | 1445.3331 | dHex Hex(3) HexNAc(4)
View | 310 | Hex(4)HexNAc(4) | G1-Fuc | Biantennary (-1 galactose) | 1460.5288 | 1461.3325 | Hex(4) HexNAc(4)
View | 424 | MolybdopterinGD | molybdopterin-gd | molybdenum bis(molybdopterin guanine dinucleotide) | 1572.9858 | 1572.0146 | H(47) C(40) N(20) O(26) P(4) S(4) Mo
View | 307 | dHex(1)Hex(4)HexNAc(4) | G1 | Fucosylated biantennary (-1 galactose) | 1606.5867 | 1607.4737 | dHex Hex(4) HexNAc(4)
View | 415 | MolybdopterinGD+Delta:S(-1)Se(1) | molybdopterin-se | L-selenocysteinyl molybdenum bis(molybdopterin guanine dinucleotide) | 1620.9302 | 1618.9096 | H(47) C(40) N(20) O(26) P(4) S(3) Se Mo
View | 311 | Hex(5)HexNAc(4) | G2-Fuc | Biantennary | 1622.5816 | 1623.4731 | Hex(5) HexNAc(4)
View | 308 | dHex(1)Hex(5)HexNAc(4) | G2 | Fucosylated biantennary | 1768.6395 | 1769.6143 | dHex Hex(5) HexNAc(4)
View | 1010 |  | DEDGFLYMVYASQETFG | Addition of DEDGFLYMVYASQETFG | 1970.8244 | 1972.088 | H(122) C(89) N(18) O(31) S
View | 960 |  | SUMO2135 | SUMOylation by SUMO-1 after tryptic cleavage | 2135.9205 | 2137.2343 | H(137) C(90) N(21) O(37) S
View | 961 |  | SUMO3549 | SUMOylation by SUMO-2/3 after tryptic cleavage | 3549.5366 | 3551.6672 | H(224) C(150) N(38) O(60) S

## Agilent
Agilent Common LC/MS Contaminants |  | Common LC/MS Contaminants - CIGS
This is a list of contaminants that have been reported with LC/MS Quad and Trap systems. It is not exhaustive and the specific natures of most of the compounds have not been identified.
Mass | Ionization Mode | Compound(s) | Source(s) | Remedy
Multiple fragments with +44 mass interval | ESI Pos | PEG | Cap Pump EMPV (normal mode) | Run Cap pump in micro mode with MS.
102 | ESI Pos | Triethylamine (TEA) | Common LC buffer, very persistent. Tubing, column, etc. | Passivate LC with 5-10% acetic acid in mobile phase. Note: If TEA is used regularly, it is better to have PEEK tubing & columns dedicated for TEA use.
106 (M-H)-, 108 (M+H)+ | APCI Pos/Neg | Silicone based | Possibly o-rings in source | Possible change in o-ring material under investigation
113 | ESI Neg | Trifluoroacetic acid (TFA) | Common LC buffer used in ESI Pos mode. Will see acetate anion as a background contaminate in ESI Neg mode. | Rinse LC system. If need to continue to use TFA as buffer, don’t scan down to 113.
114 | ESI Pos | ?? | Plastic N2 gas tubing. | Use the teflon N2 gas tubing supplied with the system.
116 | ESI Pos | Detergent | CDS bottles. | Rinse CDS bottles with solvent before adding calibrant
| ESI Neg | Acetate adducts of a fluorinated polymer (PTFE). | From lithium grease used on 1100 LC pump modules.
141.0197 |  | C4H5O2F3
223.0278 |  | C6H6O2F6
305.0345 |  | C8H7O2F9
387.0421 |  | C10H8O2F12
469.043 |  | C12H9O2F15
551.034 |  | C14H10O2F18
| ESI Pos | Sodium adducts of a fluorinated polymer (PTFE). | From lithium grease used on 1100 LC pump modules.
104.9919 |  | C2HF3Na
186.9941 |  | C4H2F6Na
268.9968 |  | C6H3F9Na
350.9996 |  | C8H4F12Na
433.0012 |  | C10H5F15Na
515.0013 |  | C12H6F18Na
596.9869 |  | C14H7F21Na
678.9799 |  | C16H8F24Na
760.9734 |  | C18H9F27Na
842.9773 |  | C20H10F30Na
924.9744 |  | C22H11F33Na
149* | ESI Pos | Red Bloom Algae | Water purification system; water solvent bottle | Run water purification system for 30 minutes before taking water; boil water; bake solvent bottle in muffle furnace at 400C.
APCI Pos
149, 209, 223, 237, 255, 391 | APCI Pos | Phthalates | Plastics. | Rinse plasticware with isopropal alcohol. Minimize use of low-density polyethylene plasticware. Use Nalgene (high-density polyethylene or polypropylene) or teflon plasticware.
| ESI Pos |  | (Nalgene bottles, tubing)
Rinse LCMS system with G1969-85025 Flushing Solvent (25/25/50 IPA/Cyclohexane/ACN). Place carbon traps in gas lines (HT3-4 For N2 – upstream from standard zeolyte trap; HT3-2 for helium)
256, 284 | ESI Pos | Nylon 66 | Altech nylon filters. | Don’t use nylon filters to filter solvents.
311, 352, 640 | ESI Pos | ?? | Column bleed | From non-Agilent ODS columns under acidic conditions.
331+ peaks 300 to 700 amu | ESI Pos | Ethylene propylene | Seals in CDS valves. You should only see this contaminate ion during tuning. | Replace calibrant in CDS bottles with acetonitrile or IPA solvent, then select Flush Calibrant System from Instrument menu in MSD Tune view to flush the calibrant system overnight (min. 8 hours).
301, 341, 360, 689, 699 | ESI Pos | ?? | Binary pump | Pump was replaced to fix problem
Pump seals and/or contaminated Teflon frit
371, 445, 519, 593, 667, 741 | ESI Pos | Polysiloxane | C-18 column bleed & common air contaminant. May also come from spray chamber vent. | 1. Check pH compatibility of column and consider ZORBAX SB or XDB phases for extended pH.
| APCI Pos | [O-Si(CH3)2-]n |  | 2. Replace column if it is old.
3. Remove exhaust line from waste bottle & check background levels.
149, 391 | ESI Pos | Phthalates | Plasticizers – Tygon tubing, solvents like methanol, and nitrogen generators.
453, 472 | ESI Pos | ?? | Binary pump. | Pump was replaced to fix problem
(unknown contaminant)
431, 481, 531, 581, 631, 681 | ESI Pos | Fluoronated polymer | New style pump seals (p/n 5063-6589) + using ACN/water as mobile phase. See MS News #114. | If possible, switch to MeOH/water or IPA/water as mobile phase, or replace pump seals with yellow polypropylene seals (p/n 0905-1420)
122, 123.9, 132, 134, 144.9, 146.9, 258, 815 | ESI Pos | ?? | Water | Replace filters on MilliQ system. May also be leeching from tubing in MilliQ system. Apparently, "high-end" system uses cleaner (teflon) tubing.
536, 610 | APCI Pos | ?? | Silicone rubber sleeve on heater. | Try to bake vaporizer at 500C overnight. If that doesn’t help, replace vaporizer.
* Also attributed to phthalates.

## Leiden
Contaminants observed in LC-MS background |  |  |  |  | http://www.lc-ms.nl/contaminants.htm
| Accurate Mass (Da) | Mode | Compound | Ion
m/z
42 | 42.0338 | + | Acetonitril | (M+H)+
59 | 59.0604 | + | Acetonitril + Ammonia | (M+NH4)+
64 | 64.0158 | + | Acetonitril + Sodium | (M+Na)+
65 | 65.0597 | + | Methanol | (2M+H)+
79 | 79.0212 | + | DMSO | (M+H)+
83 | 83.0604 | + | Acetonitrile | (2M+H)+
101 | 101.0032 | + | DMSO + Sodium | (M+Na)+
102 | 102.1277 | + | Triethylamine (TEA) | (M+H)+
105 | 105.0423 | + | Acetonitrile + Sodium | (2M+Na)+
120 | 120.0478 | + | DMSO + Acetonitril | (M+H+CH3CN)+
142 | 142.0297 | + | DMSO + Acetonitril | (M+Na+CH3CN)+
122 | 122.0812 | + | Tris (HOCH2)3CNH2 | (M+H)+
123 | 123.0917 | + | Dimethylaminopyridine (DMAP) | (M+H)+
130 | 130.159 | + | Diisopropylethylamine (DIPEA) | (M+H)+
144 | 144.1747 | + | Tripropylamine (TPA) | (M+H)+
146 | 146.0689 | + | Acetonitrile + Sodium | (3M+Na)+
150 | 150.1277 | + | Pphenyldiethylamine | (M+H)+
153 | 153.1386 | + | 1,8-diazabicyclo[5.4.0]undec-7-ene (DBU) | (M+H)+
157 | 157.0351 | + | DMSO | (2M+H)+
159 | 158.964 | + | Sodium trifluoroacetate | (M+Na)+
179 | 179.0171 | + | DMSO + Sodium | (2M+Na)+
186 | 186.2216 | + | Tributylamine | (M+H)+
214 |  | + | unidentified contaminant, maybe related to methanol | (M+H)+
225 | 225.1961 | + | Dicyclohexyl urea (DCU) (C13H24N2O) | (M+H)+
239/241 |  | + | Triethylamine | [(M.HCl)2-Cl]+
242 | 242.2842 | + | Ttetrabutylammonium (C4H9)4N+ | M+
243 | 243.1168 | + | Trityl cation | M+
235 | 235.0491 | + | DMSO | (3M+H)+
257 | 257.031 | + | DMSO | (3M+Na)+
273 | 273.1274 | + | Monomethoxytrityl cation (MMT) | M+
279 | 279.1591 | + | Dibutylphthalate (plasticiser) | (M+H)+
282 | 282.2793 | + | Oleamide | (M+H)+
301 | 301.141 | + | Dibutylphthalate (plasticiser) + Sodium | (M+Na)+
317 | 317.115 | + | Dibutylphthalate (plasticiser) + Potassium | (M+K)+
338 | 338.3417 | + | Erucamide | (M+H)+
360 | 360.3237 | + | Erucamide + Sodium | (M+Na)+
391 | 391.2843 | + | Diisooctyl phthalate (plasticiser) | (M+H)+
413 | 413.2662 | + | Diisooctyl phthalate (plasticiser) + Sodium | (M+Na)+
425 |  | + | Unidentified contaminant (plasticiser) + Sodium | (M+Na)+
429 | 429.2402 | + | Diisooctyl phthalate (plasticiser) + Potassium | (M+K)+
449 | 449.385 | + | Dicyclohexyl urea (DCU) | (2M+H)+
454 | 454.2928 | + | Diisooctyl phthalate (plasticiser) + Sodium + Acetonitril | (M+Na+CH3CN)+
798 | 798.5878 | + | Diisooctyl phthalate (plasticiser) + Ammonia | (2M+NH4)+
803 | 803.5432 | + | Diisooctyl phthalate (plasticiser) + Sodium | (2M+Na)+
231, 522, 550 |  | + | From rubber tip of disposable syringe plunger
297, 371, 445, 519 (+74n) |  | + | Polysiloxanes, from air, silicone rubber
355, 429, 503, 593, 667, 741, 815 (+74n) |  | + | Polysiloxanes, from air, silicone rubber
538 , 556, 574 |  | + | Iron carboxylate complexes: 538 = Fe3O(AcO)6, 556 = Fe3O(AcO)6(H2O), 574 = Fe3O(AcO)6(H2O)2
Repeating units observed in LC-MS background
| Accurate Mass (Da) | Mode | Compound
Ion Mass (Da)
+14n | 14.0157 | + | [CH2], Alkane chains
+15n | 15.9949 | + | O, Oxidation
+18n | 18.0106 | +/- | Water clusters
+28n | 28.0313 | + | [C2H4], natural alkene chains
+32n | 32.0262 | + | Methanol clusters
+41n | 41.0266 | + | Acetonitrile clusters
+42n | 42.047 | + | [C3H6], propyl repeating chains
+44n | 44.0262 | + | Polyethylene glycol related (-CH2CH2O-)n, PEG
+50n | 49.9968 | + | [CF2], perfluoro compounds
+53n | 53.0032 | + | Ammonium chloride adducts (NH4Cl)
+56n | 56.0626 | + | [C4H8], Butyl repeating units
+58n | 58.0419 | + | Polypropylene glycol related (CH2CH2CH2O)n
+58n | 57.9586 | + | NaCl adducts (distinguishable from PPG by Cl Isotope pattern)
+63n | 63.032 | + | Ammonium formate adducts (HCOONH4)
+68n | 67.9874 | +/- | Sodium formate adducts (HCOO-Na)
+72n | 72.0395 | + | Replacement of –OH by –OSi(CH3)3
+74n | 74.0188 | +/- | Polysiloxanes (Si(CH3)2O)
+74n | 73.9326 | + | KCl adducts (distinguishable by Cl isotope pattern)
+77n | 77.0477 | + | Ammonium acetate salts (CH3COONH4)
+78n | 78.0139 | + | DMSO adducts ((CH3)2SO)
+82n | 82.0031 | +/- | Sodium acetate adducts (CH3COONa)
+114n | 113.9929 | - | TFA (trifluoroacetic acid) adducts (CF3COOH)
+122n | 121.9383 | - | Sodium perchlorate adducts (NaClO4)
+136n | 135.9748 | +/- | Sodium TFA (trifluoroacetic acid) adducts (CF3COONa)
+162n | 162.0528 | + | Polysaccharides (C6H10O5)
+226n | 226.1681 | + | [C12H22N2O2] cyclic oligomers from polyamide
+288n | 288.1371 | - | SDS (sodium dodecylsulfate) adducts (CH3(CH2)11OSO3Na)
Mass difference ± | Origin
0.98402 | OH ↔ NH2, e.g. de-amidiation, CHNO compounds
1.9957 | F ↔ OH, halogen exchange with hydroxy group (typically -F + OH)
2.0156 | ± 2H, opening or forming of double bond
4.9554 | Na+↔ NH4+, salt adduct
7.0047 | F ↔ CN, halogen exchange with cyano group
8.9658 | Cl ↔ CN, halogen exchange with cyano group
13.9793 | O ↔ 2H, e.g. Oxidation follwed by H2O elimination
14.0157 | ± CH2, alkane chains, waxes, fatty acids, methylation
14.0157 | -[C3H6ON] ↔ -[C2H4ON], acrylamide versus iodoacetamide in cysteine alkylation (gels)
15.9739 | Na+↔ K+, salt adduct
15.9772 | S ↔ O, sulfur compounds
15.9949 | ± O, e.g. oxidation/reduction
17.0265 | ± NH3, ammonium adduct/neutral ammonium loss
17.0265 | NH4+ ↔ H+, salt adduct
17.9661 | Cl ↔ OH, halogen exchange with hydroxy group (typically -Cl + OH)
17.9906 | F ↔ H, halogen exchange
18.0106 | ± H2O, water addition/loss
18.9946 | CN ↔ COOH, nitrile compounds
20.9293 | K+↔ NH4+, salt adduct
21.9819 | Na+↔ H+, salt adduct
24.9952 | CN ↔ H, nitrile compounds
27.0109 | ± HCN, nitrile compounds
27.9949 | ± CO
28.0061 | - 2N, nitrogen loss, e.g. azido compounds (N2)
28.0313 | ± C2H4, natural alkane chains such as fatty acids
29.9742 | NO2 ↔ NH2, nitro compounds
29.998 | -NO, nitroso compounds
31.9721 | ± S, sulfur compounds
31.9898 | ± 2O, oxygen loss
33.0215 | - NH2OH, loss from hydroxamic acids
33.961 | Cl ↔ H, halogen exchange
33.9877 | ± H2S, sulfur compounds
37.9559 | K+↔ H+, salt adduct
37.9892 | 2CN ↔ 2COOH, nitrile compounds
40.0313 | +(C3H6O - H2O), acetone condensation after dehydration
42.0106 | ± COCH2
42.047 | ± C3H6, propylation
43.0058 | ± CONH (wrong calc. in ref.)
43.9495 | 79Br ↔ Cl, halogen exchange (typically -Br +Cl)
43.9898 | ± CO2
45.9474 | 81Br ↔ Cl, halogen exchange (typically -Br +Cl)
47.967 | ± SO, sulfur compounds
52.9153 | 79Br ↔ CN, halogen exchange with cyano group
54.9132 | 81Br ↔ CN, halogen exchange with cyano group
56.0626 | ± C4H8, butylation
58.0055 | ± CO2CH2
58.0419 | +C3H6O, acetone condensation
61.9156 | 79Br ↔ OH, halogen exchange with hydroxy group (typically -Br + OH)
63.9136 | 81Br ↔ OH, halogen exchange with hydroxy group (typically -Br + OH)
63.9619 | ± SO2, sulfur compounds
63.9983 | -CH3SOH, specific loss from singly oxidized methionine residues
77.9105 | 79Br ↔ H, halogen exchange
79.9085 | 81Br ↔ H, halogen exchange
79.9568 | ± SO3, sulfur compounds
91.9356 | I ↔ Cl, halogen exchange (typically -I +Cl)
97.9674 | ± H2SO4, sulfur compounds
97.9769 | ± H3PO4, phosphorous compounds
100.9014 | I ↔ CN, halogen exchange with cyano group
109.9017 | I ↔ OH, halogen exchange with hydroxy group (typically -I + OH)
125.8966 | I ↔ H, halogen exchange
146.0579 | ± [Deoxy-Hexose-H2O, C6O4H10], e.g. Fucose
162.0528 | ± [Hexose-H2O, C6O5H10], e.g. Glucose, Galactose, Mannose, Fructose
164.0685 | ± [Deoxy-Hexose-H2O, C6O5H12], e.g. Fucose
176.0321 | ± [Glucuronic acid-H2O, C6O6H8]
180.0634 | ± [Hexose, C6O6H12], e.g. Glucose, Galactose, Mannose, Fructose
194.0427 | ± [Glucuronic acid, C6O7H10]
203.0794 | ± [HexNAc-H2O, C8O5NH13], N-acetylhexoseamine
206.0579 | + sinapic acid - H2O, (C11H12O5-H2O), MALDI-matrix for proteins, possible photo- or condensation adduct
221.0899 | ± [HexNAc, C8O6NH15], N-acetylhexoseamine
289.0732 | ± [Glutathione-H2O,C10O5N3SH15]
291.0954 | ± [Neu5Ac-H2O, C11O8NH17], N-acetylneuraminic acid, sialic acid, NANA
305.0682 | ± [Glutathione+O-H2O, C10O6N3SH15]
307.0838 | ± [Glutathione, C10O6N3SH17]
309.106 | ± [Neu5Ac, C11O9NH19], N-acetylneuraminic acid, sialic acid, NANA
324.1057 | ± [Sucrose-H2O, C12O10H20]
342.1162 | ± [Sucrose, C12O11H22]
484.171 | +Tetraphenyl-tetramethyl-trisiloxane, (C28H32O2Si3), ( from silicon-based diffusion pump oil (DC7040, Dow corning)
546.1867 | +Pentaphenyl-trimethyl-trisiloxane, (C33H34O2Si3), ( from silicon-based diffusion pump oil (DC7050, Dow corning)
Table of the most common phthalates
Name | Acronym | Structural formula | CAS No. | Exact mass | (M+H)
Dimethyl phthalate | DMP | C6H4(COOCH3)2 | 131-11-3 | 194.0574 | 195.0652
Diethyl phthalate | DEP | C6H4(COOC2H5)2 | 84-66-2 | 222.0887 | 223.0965
Diallyl phthalate | DAP | C6H4(COOCH2CH=CH2)2 | 131-17-9 | 246.0887 | 247.0965
Di-n-propyl phthalate | DPP | C6H4(COO(CH2)2CH3)2 | 131-16-8 | 250.12 | 251.1278
Di-n-butyl phthalate | DBP | C6H4(COO(CH2)3CH3)2 | 84-74-2 | 278.1513 | 279.1591
Diisobutyl phthalate | DIBP | C6H4(COOCH2CH(CH3)2)2 | 84-69-5 | 278.1513 | 279.1591
Butyl cyclohexyl phthalate | BCP | CH3(CH2)3OOCC6H4COOC6H11 | 84-64-0 | 304.1669 | 305.1747
Di-n-pentyl phthalate | DNPP | C6H4(COO(CH2)4CH3)2 | 131-18-0 | 306.1826 | 307.1904
Dicyclohexyl phthalate | DCP | C6H4(COOC6H11)2 | 84-61-7 | 330.1826 | 331.1904
Butyl benzyl phthalate | BBP | CH3(CH2)3OOCC6H4COOCH2C6H5 | 85-68-7 | 312.1356 | 313.1434
Di-n-hexyl phthalate | DNHP | C6H4(COO(CH2)5CH3)2 | 84-75-3 | 334.2139 | 335.2217
Diisohexyl phthalate | DIHxP | C6H4(COO(CH2)3CH(CH3)2)2 | 146-50-9 | 334.2139 | 335.2217
Diisoheptyl phthalate | DIHpP | C6H4(COO(CH2)4CH(CH3)2)2 | 41451-28-9 | 362.2452 | 363.253
Butyl decyl phthalate | BDP | CH3(CH2)3OOCC6H4COO(CH2)9CH3 | 89-19-0 | 362.2452 | 363.253
Di(2-ethylhexyl) phthalate | DEHP, DOP | C6H4(COOCH2CH(C2H5)(CH2)3CH3)2 | 117-81-7 | 390.2765 | 391.2843
Di(n-octyl) phthalate | DNOP | C6H4(COO(CH2)7CH3)2 | 117-84-0 | 390.2765 | 391.2843
Diisooctyl phthalate | DIOP | C6H4(COO(CH2)5CH(CH3)2)2 | 27554-26-3 | 390.2765 | 391.2843
n-Octyl n-decyl phthalate | ODP | CH3(CH2)7OOCC6H4COO(CH2)9CH3 | 119-07-3 | 418.3078 | 419.3156
Diisononyl phthalate | DINP | C6H4(COO(CH2)6CH(CH3)2)2 | 28553-12-0 | 418.3078 | 419.3156
Diisodecyl phthalate | DIDP | C6H4(COO(CH2)7CH(CH3)2)2 | 26761-40-0 | 446.3391 | 447.3469
Diundecyl phthalate | DUP | C6H4(COO(CH2)10CH3)2 | 3648-20-2 | 474.3704 | 475.3782
Diisoundecyl phthalate | DIUP | C6H4(COO(CH2)8CH(CH3)2)2 | 85507-79-5 | 474.3704 | 475.3782
Ditridecyl phthalate | DTDP | C6H4(COO(CH2)12CH3)2 | 119-06-2 | 530.433 | 531.4408
Diisotridecyl phthalate | DIUP | C6H4(COO(CH2)10CH(CH3)2)2 | 68515-47-9 | 530.433 | 531.4408
http://en.wikipedia.org/wiki/Plasticizer
http://en.wikipedia.org/wiki/Phthalate

## MaConDa
MaConDa is offered to the public as a freely available resource. If you use MaConDa in your work, please cite:
Ralf J. M. Weber, Eva Li, Jonathan Bruty, Shan He & Mark R. Viant (2012). MaConDa: a publicly accessible Mass spectrometry Contaminants Database. Bioinformatics, doi:10.1093/bioinformatics/bts527.
name | formula | exact_mass | std_inchi | std_inchi_key | id | type_of_contaminant | pubchem_cid | reference | ion_mode | mz | exact_adduct_mass | ion_form | instrument_type | instrument | chromatography | ion_source_type
Acetic Acid | C2H4O2 | 60.0211 | InChI=1S/C2H4O2/c1-2(3)4/h1H3,(H,3,4) | QTBSBXVTEAMEQO-UHFFFAOYSA-N | CON00001 | Solvent | 176 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | NEG | 59 | 59.0139 | [M-H]- | Ion trap | Micromass Platform II | LC | ESI
Acetic Acid | C2H4O2 | 60.0211 | InChI=1S/C2H4O2/c1-2(3)4/h1H3,(H,3,4) | QTBSBXVTEAMEQO-UHFFFAOYSA-N | CON00001 | Solvent | 176 | Ijames, C. F., Dutky, R. C., & Fales, H. M. (1995). Iron carboxylate oxygen-centered-triangle complexes detected during electrospray use of organic acid modifiers with a comment on the Finnigan TSQ-700 electrospray inlet system. Journal of the American Society for Mass Spectrometry, 6(12), 1226-1231, doi:10.1016/1044-0305(95)00579-x. | POS | 537.88 | 537.879 | [M6-H6+Fe3+O]+ | Triple quadrupole | Finnigan TSQ-700 |  | ESI
Acetic Acid | C2H4O2 | 60.0211 | InChI=1S/C2H4O2/c1-2(3)4/h1H3,(H,3,4) | QTBSBXVTEAMEQO-UHFFFAOYSA-N | CON00001 | Solvent | 176 | Ijames, C. F., Dutky, R. C., & Fales, H. M. (1995). Iron carboxylate oxygen-centered-triangle complexes detected during electrospray use of organic acid modifiers with a comment on the Finnigan TSQ-700 electrospray inlet system. Journal of the American Society for Mass Spectrometry, 6(12), 1226-1231, doi:10.1016/1044-0305(95)00579-x. | POS | 555 | 555.8896 | [M6-H6+H2O+Fe3+O]+ | Triple quadrupole | Finnigan TSQ-700 |  | ESI
Acetic Acid | C2H4O2 | 60.0211 | InChI=1S/C2H4O2/c1-2(3)4/h1H3,(H,3,4) | QTBSBXVTEAMEQO-UHFFFAOYSA-N | CON00001 | Solvent | 176 | Ijames, C. F., Dutky, R. C., & Fales, H. M. (1995). Iron carboxylate oxygen-centered-triangle complexes detected during electrospray use of organic acid modifiers with a comment on the Finnigan TSQ-700 electrospray inlet system. Journal of the American Society for Mass Spectrometry, 6(12), 1226-1231, doi:10.1016/1044-0305(95)00579-x. | POS | 598 | 597.9001 | [M7-H6+Fe3+O]+ | Triple quadrupole | Finnigan TSQ-700 |  | ESI
Acetic Acid | C2H4O2 | 60.0211 | InChI=1S/C2H4O2/c1-2(3)4/h1H3,(H,3,4) | QTBSBXVTEAMEQO-UHFFFAOYSA-N | CON00001 | Solvent | 176 | Ijames, C. F., Dutky, R. C., & Fales, H. M. (1995). Iron carboxylate oxygen-centered-triangle complexes detected during electrospray use of organic acid modifiers with a comment on the Finnigan TSQ-700 electrospray inlet system. Journal of the American Society for Mass Spectrometry, 6(12), 1226-1231, doi:10.1016/1044-0305(95)00579-x. | POS | 658 | 657.9213 | [M8-H6+Fe3+O]+ | Triple quadrupole | Finnigan TSQ-700 |  | ESI
Acetonitrile (fragment) | CN | 26.0031 |  |  | CON00002 |  |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | NEG | 26.0036 | 26.0036 | [M]-
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 42 | 42.0338 | [M+H]+
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 59 | 59.0604 | [M+NH4]+ | Ion trap | Micromass Platform II | LC | ESI
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 64 | 64.0158 | [M+Na]+ | Ion trap | Micromass Platform II | LC | ESI
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 83 | 83.0604 | [M2+H]+ | Ion trap | Micromass Platform II | LC | ESI
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 88 | 88.0393 | [M+H+CH2O2]+
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 104 | 103.9556 | [M+63Cu(I)]+ | Ion trap | Micromass Platform II | LC | ESI
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 105 | 105.0423 | [M2+Na]+ | Ion trap | Micromass Platform II | LC | ESI
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 106 | 105.9538 | [M+65Cu(I)]+ | Ion trap | Micromass Platform II | LC | ESI
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 115 | 115.0866 | [M+H+C3H7NO]+
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 145 | 144.9821 | [M2+63Cu(I)]+
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 146 | 146.0689 | [M3+Na]+ | Ion trap | Micromass Platform II | LC | ESI
Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | CON00003 | Solvent | 6342 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 147 | 146.9803 | [M2+65Cu(I)]+
Acetonitrile.1.Acetic acid.1 | C4H7NO2 | 101.0477 |  |  | CON00004 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 102.055 | 102.055 | [M+H]+
Acetonitrile.1.Acetic acid.1 | C4H7NO2 | 101.0477 |  |  | CON00004 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 124.0369 | 124.0369 | [M+Na]+
Acetonitrile.1.Methanol.1 | C3H7NO | 73.0528 |  |  | CON00005 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 74.06 | 74.06 | [M+H]+
Acetonitrile.1.Methanol.1 | C3H7NO | 73.0528 |  |  | CON00005 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 96.042 | 96.042 | [M+Na]+
Acetonitrile.2.Methanol.2 | C6H14N2O2 | 146.1055 |  |  | CON00006 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 147.1128 | 147.1128 | [M+H]+
Acetonitrile.2.Methanol.2 | C6H14N2O2 | 146.1055 |  |  | CON00006 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 169.0947 | 169.0947 | [M+Na]+
Behentrimonium chloride | C25H54NCl | 403.3945 | InChI=1S/C25H54N.ClH/c1-5-6-7-8-9-10-11-12-13-14-15-16-17-18-19-20-21-22-23-24-25-26(2,3)4;/h5-25H2,1-4H3;1H/q+1;/p-1 | YSJGOMATDFSEED-UHFFFAOYSA-M | CON00007 | Personal care products |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 368.4251 | 368.4251 | [M-Cl]+
Bis(2-ethylhexyl) Adipate | C22H42O4 | 370.3083 | InChI=1S/C22H42O4/c1-5-9-13-19(7-3)17-25-21(23)15-11-12-16-22(24)26-18-20(8-4)14-10-6-2/h19-20H,5-18H2,1-4H3 | SAOKZLXYCUGLFA-UHFFFAOYSA-N | CON00008 | Plasticiser | 7641 | Verge, K. M., & Agnes, G. R. (2002). Plasticizer contamination from vacuum system o-rings in a quadrupole ion trap mass spectrometer. Journal of the American Society for Mass Spectrometry, 13(8), 901-905, doi:10.1016/s1044-0305(02)00386-0. | POS | 370.57 | 371.3156 | [M+H]+ | Ion trap | Model 3DQ | GC | ESI
Bis(2-ethylhexyl) Phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-5-9-13-19(7-3)17-27-23(25)21-15-11-12-16-22(21)24(26)28-18-20(8-4)14-10-6-2/h11-12,15-16,19-20H,5-10,13-14,17-18H2,1-4H3 | BJQHLKABXJIVAM-UHFFFAOYSA-N | CON00009 | Plasticiser | 8343 | Verge, K. M., & Agnes, G. R. (2002). Plasticizer contamination from vacuum system o-rings in a quadrupole ion trap mass spectrometer. Journal of the American Society for Mass Spectrometry, 13(8), 901-905, doi:10.1016/s1044-0305(02)00386-0. | POS | 390.56 | 391.2843 | [M+H]+ | Ion trap | Model 3DQ | GC | ESI
Bis(2-ethylhexyl) Phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-5-9-13-19(7-3)17-27-23(25)21-15-11-12-16-22(21)24(26)28-18-20(8-4)14-10-6-2/h11-12,15-16,19-20H,5-10,13-14,17-18H2,1-4H3 | BJQHLKABXJIVAM-UHFFFAOYSA-N | CON00009 | Plasticiser | 8343 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 391.2843 | 391.2843 | [M+H]+
Bis(2-ethylhexyl) Phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-5-9-13-19(7-3)17-27-23(25)21-15-11-12-16-22(21)24(26)28-18-20(8-4)14-10-6-2/h11-12,15-16,19-20H,5-10,13-14,17-18H2,1-4H3 | BJQHLKABXJIVAM-UHFFFAOYSA-N | CON00009 | Plasticiser | 8343 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 391.3 | 391.2843 | [M+H]+
Bis(2-ethylhexyl) Phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-5-9-13-19(7-3)17-27-23(25)21-15-11-12-16-22(21)24(26)28-18-20(8-4)14-10-6-2/h11-12,15-16,19-20H,5-10,13-14,17-18H2,1-4H3 | BJQHLKABXJIVAM-UHFFFAOYSA-N | CON00009 | Plasticiser | 8343 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 413.2662 | 413.2662 | [M+Na]+
Bis(2-ethylhexyl) Phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-5-9-13-19(7-3)17-27-23(25)21-15-11-12-16-22(21)24(26)28-18-20(8-4)14-10-6-2/h11-12,15-16,19-20H,5-10,13-14,17-18H2,1-4H3 | BJQHLKABXJIVAM-UHFFFAOYSA-N | CON00009 | Plasticiser | 8343 | Thermo Corporation, Notes on Troubleshooting LC/MS Contamination | POS | 798 | 798.5878 | [M2+NH4]+
Bis(2-ethylhexyl) Phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-5-9-13-19(7-3)17-27-23(25)21-15-11-12-16-22(21)24(26)28-18-20(8-4)14-10-6-2/h11-12,15-16,19-20H,5-10,13-14,17-18H2,1-4H3 | BJQHLKABXJIVAM-UHFFFAOYSA-N | CON00009 | Plasticiser | 8343 | Thermo Corporation, Notes on Troubleshooting LC/MS Contamination | POS | 803 | 803.5432 | [M2+Na]+
Butyl Carbitol | C8H18O3 | 162.1256 | InChI=1S/C8H18O3/c1-2-3-5-10-7-8-11-6-4-9/h9H,2-8H2,1H3 | OAYXUHPQHDHDDZ-UHFFFAOYSA-N | CON00010 | Scintillation cocktail | 8177 | Gibson, C. R., & Brown, C. M. (2003). Identification of diethylene glycol monobutyl ether as a source of contamination in an ion trap mass spectrometer. Journal of the American Society for Mass Spectrometry, 14(11), 1247-1249, doi:10.1016/s1044-0305(03)00534-8. | POS | 163 | 163.1329 | [M+H]+ | Ion trap | ThermoFinnigan LCQ Deca XP | LC | ESI
Butyl Carbitol | C8H18O3 | 162.1256 | InChI=1S/C8H18O3/c1-2-3-5-10-7-8-11-6-4-9/h9H,2-8H2,1H3 | OAYXUHPQHDHDDZ-UHFFFAOYSA-N | CON00010 | Scintillation cocktail | 8177 | Gibson, C. R., & Brown, C. M. (2003). Identification of diethylene glycol monobutyl ether as a source of contamination in an ion trap mass spectrometer. Journal of the American Society for Mass Spectrometry, 14(11), 1247-1249, doi:10.1016/s1044-0305(03)00534-8. | POS | 324.9 | 325.2585 | [M2+H]+ | Ion trap | ThermoFinnigan LCQ Deca XP | LC | ESI
Butylated Hydroxyanisole | C11H16O2 | 180.115 | InChI=1S/C11H16O2/c1-11(2,3)9-7-8(13-4)5-6-10(9)12/h5-7,12H,1-4H3 | MRBKEAMVRSLQPH-UHFFFAOYSA-N | CON00011 | Antioxidant | 11954184 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 181.1223 | 181.1223 | [M+H]+
Butylated Hydroxyanisole | C11H16O2 | 180.115 | InChI=1S/C11H16O2/c1-11(2,3)9-7-8(13-4)5-6-10(9)12/h5-7,12H,1-4H3 | MRBKEAMVRSLQPH-UHFFFAOYSA-N | CON00011 | Antioxidant | 11954184 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 203.1042 | 203.1043 | [M+Na]+
Butylated Hydroxytoluene | C15H24O | 220.1827 | InChI=1S/C15H24O/c1-10-8-11(14(2,3)4)13(16)12(9-10)15(5,6)7/h8-9,16H,1-7H3 | NLZUEZXRPGMBCV-UHFFFAOYSA-N | CON00012 | Antioxidant | 31404 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 221.19 | 221.19 | [M+H]+
Butylated Hydroxytoluene | C15H24O | 220.1827 | InChI=1S/C15H24O/c1-10-8-11(14(2,3)4)13(16)12(9-10)15(5,6)7/h8-9,16H,1-7H3 | NLZUEZXRPGMBCV-UHFFFAOYSA-N | CON00012 | Antioxidant | 31404 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 243.1719 | 243.1719 | [M+Na]+
Copper catalyst iodide adduct (63Cu(I)) | 63Cu(I) | 189.8341 |  |  | CON00013 |  |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | NEG | 316.7391 | 316.7391 | [M+I]-
Copper catalyst iodide adduct (65Cu(I)) | 65Cu(I) | 191.8323 |  |  | CON00014 |  |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | NEG | 318.7373 | 318.7373 | [M+I]-
Cs-133 | Cs | 132.9055 |  |  | CON00015 |  |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 132.9049 | 132.9049 | [M]+
d6-Dimethylsulfoxide | C2D6OS | 84.0516 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3/i1D3,2D3 | IAZDPXIOMUYVGZ-WFGJKAKNSA-N | CON00016 | Solvent | 75151 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 85 | 85.0589 | [M+H]+
d6-Dimethylsulfoxide | C2D6OS | 84.0516 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3/i1D3,2D3 | IAZDPXIOMUYVGZ-WFGJKAKNSA-N | CON00016 | Solvent | 75151 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 169 | 169.1105 | [M2+H]+
d6-Dimethylsulfoxide | C2D6OS | 84.0516 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3/i1D3,2D3 | IAZDPXIOMUYVGZ-WFGJKAKNSA-N | CON00016 | Solvent | 75151 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 85 | 85.0589 | [M+H]+
d6-Dimethylsulfoxide | C2D6OS | 84.0516 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3/i1D3,2D3 | IAZDPXIOMUYVGZ-WFGJKAKNSA-N | CON00016 | Solvent | 75151 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 169 | 169.1105 | [M2+H]+
DDTDP | C30H58O4S | 514.4056 |  |  | CON00017 |  |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 515.4129 | 515.4129 | [M+H]+
Decamethylcyclopentasiloxane | C10H30Si5O5 | 370.094 | InChI=1S/C10H30O5Si5/c1-16(2)11-17(3,4)13-19(7,8)15-20(9,10)14-18(5,6)12-16/h1-10H3 | XMSXQFUHVRWGNA-UHFFFAOYSA-N | CON00018 | Airborne Contaminant | 10193 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 355 | 355.0699 | [M+H-CH4]+
Decamethylcyclopentasiloxane | C10H30Si5O5 | 370.094 | InChI=1S/C10H30O5Si5/c1-16(2)11-17(3,4)13-19(7,8)15-20(9,10)14-18(5,6)12-16/h1-10H3 | XMSXQFUHVRWGNA-UHFFFAOYSA-N | CON00018 | Airborne Contaminant | 10193 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 355.092 | 355.0699 | [M+H-CH4]+ | ToF | Q-TOF Micro(Micromass) |  | nanoESI
Decamethylcyclopentasiloxane | C10H30Si5O5 | 370.094 | InChI=1S/C10H30O5Si5/c1-16(2)11-17(3,4)13-19(7,8)15-20(9,10)14-18(5,6)12-16/h1-10H3 | XMSXQFUHVRWGNA-UHFFFAOYSA-N | CON00018 | Airborne Contaminant | 10193 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 371 | 371.1012 | [M+H]+
Decamethylcyclopentasiloxane | C10H30Si5O5 | 370.094 | InChI=1S/C10H30O5Si5/c1-16(2)11-17(3,4)13-19(7,8)15-20(9,10)14-18(5,6)12-16/h1-10H3 | XMSXQFUHVRWGNA-UHFFFAOYSA-N | CON00018 | Airborne Contaminant | 10193 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 371.092 | 371.1012 | [M+H]+ | ToF | Q-TOF Micro(Micromass) |  | nanoESI
Decamethylcyclopentasiloxane | C10H30Si5O5 | 370.094 | InChI=1S/C10H30O5Si5/c1-16(2)11-17(3,4)13-19(7,8)15-20(9,10)14-18(5,6)12-16/h1-10H3 | XMSXQFUHVRWGNA-UHFFFAOYSA-N | CON00018 | Airborne Contaminant | 10193 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 388 | 388.1278 | [M+NH4]+
Decamethylcyclopentasiloxane | C10H30Si5O5 | 370.094 | InChI=1S/C10H30O5Si5/c1-16(2)11-17(3,4)13-19(7,8)15-20(9,10)14-18(5,6)12-16/h1-10H3 | XMSXQFUHVRWGNA-UHFFFAOYSA-N | CON00018 | Airborne Contaminant | 10193 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 388.092 | 388.1278 | [M+NH4]+ | ToF | Q-TOF Micro(Micromass) |  | nanoESI
Dibutyl phthalate | C16H22O4 | 278.1518 | InChI=1S/C16H22O4/c1-3-5-7-11-9-10-13(15(17)18)14(16(19)20)12(11)8-6-4-2/h9-10H,3-8H2,1-2H3,(H,17,18)(H,19,20)/p-2 | NUONTTMMRCYJMI-UHFFFAOYSA-L | CON00019 | Plasticiser | 3026 | NewObjective Common Background Ions for Electrospray (Positive Ion). http://www.newobjective.com/downloads/technotes/PV-3.pdf. | POS | 279 | 279.1591 | [M+H]+
Dibutyl phthalate | C16H22O4 | 278.1518 | InChI=1S/C16H22O4/c1-3-5-7-11-9-10-13(15(17)18)14(16(19)20)12(11)8-6-4-2/h9-10H,3-8H2,1-2H3,(H,17,18)(H,19,20)/p-2 | NUONTTMMRCYJMI-UHFFFAOYSA-L | CON00019 | Plasticiser | 3026 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 279.1591 | 279.1591 | [M+H]+
Dibutyl phthalate | C16H22O4 | 278.1518 | InChI=1S/C16H22O4/c1-3-5-7-11-9-10-13(15(17)18)14(16(19)20)12(11)8-6-4-2/h9-10H,3-8H2,1-2H3,(H,17,18)(H,19,20)/p-2 | NUONTTMMRCYJMI-UHFFFAOYSA-L | CON00019 | Plasticiser | 3026 | NewObjective Common Background Ions for Electrospray (Positive Ion). http://www.newobjective.com/downloads/technotes/PV-3.pdf. | POS | 301 | 301.141 | [M+Na]+
Dibutyl phthalate | C16H22O4 | 278.1518 | InChI=1S/C16H22O4/c1-3-5-7-11-9-10-13(15(17)18)14(16(19)20)12(11)8-6-4-2/h9-10H,3-8H2,1-2H3,(H,17,18)(H,19,20)/p-2 | NUONTTMMRCYJMI-UHFFFAOYSA-L | CON00019 | Plasticiser | 3026 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 301.141 | 301.141 | [M+Na]+
Dibutyl phthalate | C16H22O4 | 278.1518 | InChI=1S/C16H22O4/c1-3-5-7-11-9-10-13(15(17)18)14(16(19)20)12(11)8-6-4-2/h9-10H,3-8H2,1-2H3,(H,17,18)(H,19,20)/p-2 | NUONTTMMRCYJMI-UHFFFAOYSA-L | CON00019 | Plasticiser | 3026 | NewObjective Common Background Ions for Electrospray (Positive Ion). http://www.newobjective.com/downloads/technotes/PV-3.pdf. | POS | 317 | 317.115 | [M+K]+
Dibutyl Sebacate | C18H34O4 | 314.2457 | InChI=1S/C18H34O4/c1-3-5-15-21-17(19)13-11-9-7-8-10-12-14-18(20)22-16-6-4-2/h3-16H2,1-2H3 | PYGXAGIECVVIOZ-UHFFFAOYSA-N | CON00020 | Plasticiser | 7986 | Guo, X. H., Bruins, A. P., & Covey, T. R. (2006). Characterization of typical chemical background interferences in atmospheric pressure ionization liquid chromatography-mass spectrometry. Rapid Communications in Mass Spectrometry, 20(20), 3145-3150, doi:10.1002/rcm.2715. | POS | 315 | 315.253 | [M+H]+ | Triple quadrupole | API365 | LC
Didodecyl 3,3'-thiodipropanoate sulfone product | C30H58O6S | 546.3954 |  |  | CON00021 | Antioxidant |  | Xia, Y. Q., Patel, S., Bakhtiar, R., Franklin, R. B., & Doss, G. A. (2005). Identification of a new source of interference leached from polypropylene tubes in mass-selective analysis. Journal of the American Society for Mass Spectrometry, 16(3), 417-421, doi:10.1016/j.jasms.2004.11.020. | POS | 569.3829 | 569.3846 | [M+Na]+ | ToF | Micromass QTOF 2 | LC | ESI
Didodecyl 3,3'-thiodipropanoate sulfone product | C30H58O6S | 546.3954 |  |  | CON00021 | Antioxidant |  | Xia, Y. Q., Patel, S., Bakhtiar, R., Franklin, R. B., & Doss, G. A. (2005). Identification of a new source of interference leached from polypropylene tubes in mass-selective analysis. Journal of the American Society for Mass Spectrometry, 16(3), 417-421, doi:10.1016/j.jasms.2004.11.020. | POS | 547.4036 | 547.4027 | [M+H]+ | ToF | Micromass QTOF 2 | LC | ESI
Didodecyl 3,3'-thiodipropanoate sulfone product | C30H58O6S | 546.3954 |  |  | CON00021 | Antioxidant |  | Xia, Y. Q., Patel, S., Bakhtiar, R., Franklin, R. B., & Doss, G. A. (2005). Identification of a new source of interference leached from polypropylene tubes in mass-selective analysis. Journal of the American Society for Mass Spectrometry, 16(3), 417-421, doi:10.1016/j.jasms.2004.11.020. | POS | 585.3581 | 585.3586 | [M+K]+ | ToF | Micromass QTOF 2 | LC | ESI
Didodecyl 3,3'-thiodipropanoate sulfoxide product | C30H58O5S | 530.4005 |  |  | CON00022 | Antioxidant |  | Xia, Y. Q., Patel, S., Bakhtiar, R., Franklin, R. B., & Doss, G. A. (2005). Identification of a new source of interference leached from polypropylene tubes in mass-selective analysis. Journal of the American Society for Mass Spectrometry, 16(3), 417-421, doi:10.1016/j.jasms.2004.11.020. | POS | 531.4083 | 531.4078 | [M+H]+ | ToF | Micromass QTOF 2 | LC | ESI
Didodecyl 3,3'-thiodipropanoate sulfoxide product | C30H58O5S | 530.4005 |  |  | CON00022 | Antioxidant |  | Xia, Y. Q., Patel, S., Bakhtiar, R., Franklin, R. B., & Doss, G. A. (2005). Identification of a new source of interference leached from polypropylene tubes in mass-selective analysis. Journal of the American Society for Mass Spectrometry, 16(3), 417-421, doi:10.1016/j.jasms.2004.11.020. | POS | 553.3903 | 553.3897 | [M+Na]+ | ToF | Micromass QTOF 2 | LC | ESI
Didodecyl 3,3'-thiodipropanoate sulfoxide product | C30H58O5S | 530.4005 |  |  | CON00022 | Antioxidant |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 569 | 569.3637 | [M+K]+
Didodecyl 3,3'-thiodipropanoate sulfoxide product | C30H58O5S | 530.4005 |  |  | CON00022 | Antioxidant |  | Xia, Y. Q., Patel, S., Bakhtiar, R., Franklin, R. B., & Doss, G. A. (2005). Identification of a new source of interference leached from polypropylene tubes in mass-selective analysis. Journal of the American Society for Mass Spectrometry, 16(3), 417-421, doi:10.1016/j.jasms.2004.11.020. | POS | 569.3662 | 569.3637 | [M+K]+ | ToF | Micromass QTOF 2 | LC | ESI
Diethyl carbitol | C8H18O3 | 162.1256 | InChI=1S/C8H18O3/c1-3-9-5-7-11-8-6-10-4-2/h3-8H2,1-2H3 | RRQYJINTUHWNHW-UHFFFAOYSA-N | CON00023 | Scintillation cocktail | 8179 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 185.1148 | 185.1148 | [M+Na]+
Diisooctyl phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-19(2)13-7-5-11-17-27-23(25)21-15-9-10-16-22(21)24(26)28-18-12-6-8-14-20(3)4/h9-10,15-16,19-20H,5-8,11-14,17-18H2,1-4H3 | IJFPVINAQGWBRJ-UHFFFAOYSA-N | CON00024 | Plasticiser | 33934 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 391 | 391.2843 | [M+H]+
Diisooctyl phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-19(2)13-7-5-11-17-27-23(25)21-15-9-10-16-22(21)24(26)28-18-12-6-8-14-20(3)4/h9-10,15-16,19-20H,5-8,11-14,17-18H2,1-4H3 | IJFPVINAQGWBRJ-UHFFFAOYSA-N | CON00024 | Plasticiser | 33934 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 413 | 413.2662 | [M+Na]+
Diisooctyl phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-19(2)13-7-5-11-17-27-23(25)21-15-9-10-16-22(21)24(26)28-18-12-6-8-14-20(3)4/h9-10,15-16,19-20H,5-8,11-14,17-18H2,1-4H3 | IJFPVINAQGWBRJ-UHFFFAOYSA-N | CON00024 | Plasticiser | 33934 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 429 | 429.2402 | [M+K]+
Diisooctyl phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-19(2)13-7-5-11-17-27-23(25)21-15-9-10-16-22(21)24(26)28-18-12-6-8-14-20(3)4/h9-10,15-16,19-20H,5-8,11-14,17-18H2,1-4H3 | IJFPVINAQGWBRJ-UHFFFAOYSA-N | CON00024 | Plasticiser | 33934 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 454 | 454.2928 | [M+Na+CH3CN]+
Diisooctyl phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-19(2)13-7-5-11-17-27-23(25)21-15-9-10-16-22(21)24(26)28-18-12-6-8-14-20(3)4/h9-10,15-16,19-20H,5-8,11-14,17-18H2,1-4H3 | IJFPVINAQGWBRJ-UHFFFAOYSA-N | CON00024 | Plasticiser | 33934 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 798 | 798.5878 | [M2+NH4]+
Diisooctyl phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-19(2)13-7-5-11-17-27-23(25)21-15-9-10-16-22(21)24(26)28-18-12-6-8-14-20(3)4/h9-10,15-16,19-20H,5-8,11-14,17-18H2,1-4H3 | IJFPVINAQGWBRJ-UHFFFAOYSA-N | CON00024 | Plasticiser | 33934 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 803 | 803.5432 | [M2+Na]+
Diisooctyl phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-19(2)13-7-5-11-17-27-23(25)21-15-9-10-16-22(21)24(26)28-18-12-6-8-14-20(3)4/h9-10,15-16,19-20H,5-8,11-14,17-18H2,1-4H3 | IJFPVINAQGWBRJ-UHFFFAOYSA-N | CON00024 | Plasticiser | 33934 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 819.5172 | 819.5172 | [M2+K]+
Diisopropylethylamine | C8H19N | 129.1517 | InChI=1S/C8H19N/c1-6-9(7(2)3)8(4)5/h7-8H,6H2,1-5H3 |  | CON00025 | Solvent | 81531 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 130 | 130.159 | [M+H]+
Dimethyl Formamide | C3H7NO | 73.0528 | InChI=1S/C3H7NO/c1-4(2)3-5/h3H,1-2H3 | ZMXDDKWLCZADIW-UHFFFAOYSA-N | CON00026 | Solvent | 6228 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 74 | 74.06 | [M+H]+ | Ion Trap | Micromass Platform II | LC | ESI
Dimethyl Phthalate | C10H10O4 | 194.0579 | InChI=1S/C10H10O4/c1-13-9(11)7-5-3-4-6-8(7)10(12)14-2/h3-6H,1-2H3 |  | CON00027 | Plasticiser | 8554 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 163 | 163.039 | [M+H-CH3OH]+
Dimethyl Phthalate | C10H10O4 | 194.0579 | InChI=1S/C10H10O4/c1-13-9(11)7-5-3-4-6-8(7)10(12)14-2/h3-6H,1-2H3 |  | CON00027 | Plasticiser | 8554 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 195 | 195.0652 | [M+H]+
Dimethylaminopyridine | C7H10N2 | 122.0844 | InChI=1S/C7H10N2/c1-9(2)7-5-3-4-6-8-7/h3-6H,1-2H3 | PSHKMPUSSFXUIA-UHFFFAOYSA-N | CON00028 | Solvent | 21885 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 123 | 123.0917 | [M+H]+
Dimethylsulfoxide (DMSO) | C2H6OS | 78.0139 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3 | IAZDPXIOMUYVGZ-UHFFFAOYSA-N | CON00029 | Solvent | 679 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 79 | 79.0212 | [M+H]+ | Ion trap | Micromass Platform II | LC | ESI
Dimethylsulfoxide (DMSO) | C2H6OS | 78.0139 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3 | IAZDPXIOMUYVGZ-UHFFFAOYSA-N | CON00029 | Solvent | 679 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 101 | 101.0032 | [M+Na]+ | Ion trap | Micromass Platform II | LC | ESI
Dimethylsulfoxide (DMSO) | C2H6OS | 78.0139 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3 | IAZDPXIOMUYVGZ-UHFFFAOYSA-N | CON00029 | Solvent | 679 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 120 | 120.0478 | [M+H+CH3CN]+ | Ion trap | Micromass Platform II | LC | ESI
Dimethylsulfoxide (DMSO) | C2H6OS | 78.0139 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3 | IAZDPXIOMUYVGZ-UHFFFAOYSA-N | CON00029 | Solvent | 679 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 137 | 137.0743 | [M+NH4+CH3CN]+ | Ion trap | Micromass Platform II | LC | ESI
Dimethylsulfoxide (DMSO) | C2H6OS | 78.0139 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3 | IAZDPXIOMUYVGZ-UHFFFAOYSA-N | CON00029 | Solvent | 679 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 142.0297 | 142.0297 | [M+Na+CH3CN]+
Dimethylsulfoxide (DMSO) | C2H6OS | 78.0139 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3 | IAZDPXIOMUYVGZ-UHFFFAOYSA-N | CON00029 | Solvent | 679 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 157 | 157.0352 | [M2+H]+ | Ion trap | Micromass Platform II | LC | ESI
Dimethylsulfoxide (DMSO) | C2H6OS | 78.0139 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3 | IAZDPXIOMUYVGZ-UHFFFAOYSA-N | CON00029 | Solvent | 679 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | POS | 179 | 179.0171 | [M2+Na]+ | Ion trap | Micromass Platform II | LC | ESI
Dimethylsulfoxide (DMSO) | C2H6OS | 78.0139 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3 | IAZDPXIOMUYVGZ-UHFFFAOYSA-N | CON00029 | Solvent | 679 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 257 | 257.031 | [M3+Na]+
Dipalmityldimethylammonium chloride | C34H72NCl | 529.5353 |  |  | CON00030 | Personal care products |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 494.5659 | 494.5659 | [M-Cl]+
Diphenyl ketone | C13H10O | 182.0732 | InChI=1S/C13H10O/c14-13(11-7-3-1-4-8-11)12-9-5-2-6-10-12/h1-10H | RWCCWEUUXYIKHB-UHFFFAOYSA-N | CON00031 |  | 3102 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 183.0804 | 183.0804 | [M+H]+
Distearyldimethylammonium chloride | C38H80NCl | 585.5979 | InChI=1S/C38H80N.ClH/c1-5-7-9-11-13-15-17-19-21-23-25-27-29-31-33-35-37-39(3,4)38-36-34-32-30-28-26-24-22-20-18-16-14-12-10-8-6-2;/h5-38H2,1-4H3;1H/q+1;/p-1 | REZZEXDLIUJMMS-UHFFFAOYSA-M | CON00032 | Personal care products | 7879 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 550 | 550.6285 | [M-Cl]+
Dodecamethylcyclohexasiloxane | C12H36Si6O6 | 444.1128 | InChI=1S/C12H36O6Si6/c1-19(2)13-20(3,4)15-22(7,8)17-24(11,12)18-23(9,10)16-21(5,6)14-19/h1-12H3 | IUMSDRXLFWAGNT-UHFFFAOYSA-N | CON00033 | Airborne Contaminant | 109111 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 429.101 | 429.0887 | [M+H-CH4]+ | ToF | Q-TOF Micro(Micromass) |  | nanoESI
Dodecamethylcyclohexasiloxane | C12H36Si6O6 | 444.1128 | InChI=1S/C12H36O6Si6/c1-19(2)13-20(3,4)15-22(7,8)17-24(11,12)18-23(9,10)16-21(5,6)14-19/h1-12H3 | IUMSDRXLFWAGNT-UHFFFAOYSA-N | CON00033 | Airborne Contaminant | 109111 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 445 | 445.12 | [M+H]+
Dodecamethylcyclohexasiloxane | C12H36Si6O6 | 444.1128 | InChI=1S/C12H36O6Si6/c1-19(2)13-20(3,4)15-22(7,8)17-24(11,12)18-23(9,10)16-21(5,6)14-19/h1-12H3 | IUMSDRXLFWAGNT-UHFFFAOYSA-N | CON00033 | Airborne Contaminant | 109111 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 445.101 | 445.12 | [M+H]+ | ToF | Q-TOF Micro(Micromass) |  | nanoESI
Dodecamethylcyclohexasiloxane | C12H36Si6O6 | 444.1128 | InChI=1S/C12H36O6Si6/c1-19(2)13-20(3,4)15-22(7,8)17-24(11,12)18-23(9,10)16-21(5,6)14-19/h1-12H3 | IUMSDRXLFWAGNT-UHFFFAOYSA-N | CON00033 | Airborne Contaminant | 109111 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 462 | 462.1466 | [M+NH4]+
Dodecamethylcyclohexasiloxane | C12H36Si6O6 | 444.1128 | InChI=1S/C12H36O6Si6/c1-19(2)13-20(3,4)15-22(7,8)17-24(11,12)18-23(9,10)16-21(5,6)14-19/h1-12H3 | IUMSDRXLFWAGNT-UHFFFAOYSA-N | CON00033 | Airborne Contaminant | 109111 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 462.101 | 462.1466 | [M+NH4]+ | ToF | Q-TOF Micro(Micromass) |  | nanoESI
Eicosamethylcyclodecasiloxane | C20H60Si10O10 | 740.1879 | InChI=1S/C20H60O10Si10/c1-31(2)21-32(3,4)23-34(7,8)25-36(11,12)27-38(15,16)29-40(19,20)30-39(17,18)28-37(13,14)26-35(9,10)24-33(5,6)22-31/h1-20H3 | CDNNKGWZSNSADW-UHFFFAOYSA-N | CON00034 | Airborne Contaminant | 519601 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 725.189 | 725.1639 | [M+H-CH4]+
Eicosamethylcyclodecasiloxane | C20H60Si10O10 | 740.1879 | InChI=1S/C20H60O10Si10/c1-31(2)21-32(3,4)23-34(7,8)25-36(11,12)27-38(15,16)29-40(19,20)30-39(17,18)28-37(13,14)26-35(9,10)24-33(5,6)22-31/h1-20H3 | CDNNKGWZSNSADW-UHFFFAOYSA-N | CON00034 | Airborne Contaminant | 519601 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 741 | 741.1952 | [M+H]+
Eicosamethylcyclodecasiloxane | C20H60Si10O10 | 740.1879 | InChI=1S/C20H60O10Si10/c1-31(2)21-32(3,4)23-34(7,8)25-36(11,12)27-38(15,16)29-40(19,20)30-39(17,18)28-37(13,14)26-35(9,10)24-33(5,6)22-31/h1-20H3 | CDNNKGWZSNSADW-UHFFFAOYSA-N | CON00034 | Airborne Contaminant | 519601 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 741.189 | 741.1952 | [M+H]+
Eicosamethylcyclodecasiloxane | C20H60Si10O10 | 740.1879 | InChI=1S/C20H60O10Si10/c1-31(2)21-32(3,4)23-34(7,8)25-36(11,12)27-38(15,16)29-40(19,20)30-39(17,18)28-37(13,14)26-35(9,10)24-33(5,6)22-31/h1-20H3 | CDNNKGWZSNSADW-UHFFFAOYSA-N | CON00034 | Airborne Contaminant | 519601 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 758.189 | 758.2217 | [M+NH4]+
Erucamide | C22H43NO | 337.3345 | InChI=1S/C22H43NO/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18-19-20-21-22(23)24/h9-10H,2-8,11-21H2,1H3,(H2,23,24)/b10-9- | UAUDZVJPLUQNMU-KTKRTIGZSA-N | CON00035 | Anti-static agent | 5365371 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 338 | 338.3417 | [M+H]+
Erucamide | C22H43NO | 337.3345 | InChI=1S/C22H43NO/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18-19-20-21-22(23)24/h9-10H,2-8,11-21H2,1H3,(H2,23,24)/b10-9- | UAUDZVJPLUQNMU-KTKRTIGZSA-N | CON00035 | Anti-static agent | 5365371 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 360 | 360.3237 | [M+Na]+
Formic acid | H2CO2 | 46.0055 | InChI=1S/CH2O2/c2-1-3/h1H,(H,2,3) | BDAGIHXWWSANSR-UHFFFAOYSA-N | CON00036 |  | 284 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | NEG | 45 | 44.9982 | [M-H]-
Formic acid | H2CO2 | 46.0055 | InChI=1S/CH2O2/c2-1-3/h1H,(H,2,3) | BDAGIHXWWSANSR-UHFFFAOYSA-N | CON00036 |  |  | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | NEG | 113 | 112.9856 | [M2+Na-H2]-
Hexadecamethylcyclooctasiloxane | C16H48Si8O8 | 592.1503 | InChI=1S/C16H48O8Si8/c1-25(2)17-26(3,4)19-28(7,8)21-30(11,12)23-32(15,16)24-31(13,14)22-29(9,10)20-27(5,6)18-25/h1-16H3 | XKJMJYZFAWYREL-UHFFFAOYSA-N | CON00037 | Airborne Contaminant | 11170 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 577.131 | 577.1263 | [M+H-CH4]+
Hexadecamethylcyclooctasiloxane | C16H48Si8O8 | 592.1503 | InChI=1S/C16H48O8Si8/c1-25(2)17-26(3,4)19-28(7,8)21-30(11,12)23-32(15,16)24-31(13,14)22-29(9,10)20-27(5,6)18-25/h1-16H3 | XKJMJYZFAWYREL-UHFFFAOYSA-N | CON00037 | Airborne Contaminant | 11170 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 593 | 593.1576 | [M+H]+
Hexadecamethylcyclooctasiloxane | C16H48Si8O8 | 592.1503 | InChI=1S/C16H48O8Si8/c1-25(2)17-26(3,4)19-28(7,8)21-30(11,12)23-32(15,16)24-31(13,14)22-29(9,10)20-27(5,6)18-25/h1-16H3 | XKJMJYZFAWYREL-UHFFFAOYSA-N | CON00037 | Airborne Contaminant | 11170 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 593.131 | 593.1576 | [M+H]+
Hexadecamethylcyclooctasiloxane | C16H48Si8O8 | 592.1503 | InChI=1S/C16H48O8Si8/c1-25(2)17-26(3,4)19-28(7,8)21-30(11,12)23-32(15,16)24-31(13,14)22-29(9,10)20-27(5,6)18-25/h1-16H3 | XKJMJYZFAWYREL-UHFFFAOYSA-N | CON00037 | Airborne Contaminant | 11170 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 610.131 | 610.1842 | [M+NH4]+
Iodine anion (Iodide) | I | 126.9045 |  |  | CON00038 |  |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | NEG | 126.905 | 126.905 | [M]-
Irganox | C35H62O3 | 530.4699 | InChI=1S/C35H62O3/c1-8-9-10-11-12-13-14-15-16-17-18-19-20-21-22-23-26-38-32(36)25-24-29-27-30(34(2,3)4)33(37)31(28-29)35(5,6)7/h27-28,37H,8-26H2,1-7H3 | SSDSCDGVMJFTEQ-UHFFFAOYSA-N | CON00039 | Antioxidant | 16386 | Xia, Y. Q., Patel, S., Bakhtiar, R., Franklin, R. B., & Doss, G. A. (2005). Identification of a new source of interference leached from polypropylene tubes in mass-selective analysis. Journal of the American Society for Mass Spectrometry, 16(3), 417-421, doi:10.1016/j.jasms.2004.11.020. | POS | 531 | 531.4772 | [M+H]+ | ToF | Micromass QTOF 2 | LC | ESI
Irganox | C35H62O3 | 530.4699 | InChI=1S/C35H62O3/c1-8-9-10-11-12-13-14-15-16-17-18-19-20-21-22-23-26-38-32(36)25-24-29-27-30(34(2,3)4)33(37)31(28-29)35(5,6)7/h27-28,37H,8-26H2,1-7H3 | SSDSCDGVMJFTEQ-UHFFFAOYSA-N | CON00039 | Antioxidant | 16386 | Xia, Y. Q., Patel, S., Bakhtiar, R., Franklin, R. B., & Doss, G. A. (2005). Identification of a new source of interference leached from polypropylene tubes in mass-selective analysis. Journal of the American Society for Mass Spectrometry, 16(3), 417-421, doi:10.1016/j.jasms.2004.11.020. | POS | 553 | 553.4591 | [M+Na]+ | ToF | Micromass QTOF 2 | LC | ESI
Methanesulfonate | CH4O3S | 95.9881 | InChI=1S/CH4O3S/c1-5(2,3)4/h1H3,(H,2,3,4)/p-1 | AFVFQIVMOAPDHO-UHFFFAOYSA-M | CON00040 |  |  | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | NEG | 95 | 94.9808 | [M-H]-
Methanol | CH3OH | 32.0262 | InChI=1S/CH4O/c1-2/h2H,1H3 | OKKJLVBELUTLKV-UHFFFAOYSA-N | CON00041 | Solvent | 887 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 33 | 33.0335 | [M+H]+
Methanol | CH3OH | 32.0262 | InChI=1S/CH4O/c1-2/h2H,1H3 | OKKJLVBELUTLKV-UHFFFAOYSA-N | CON00041 | Solvent | 887 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 65 | 65.0597 | [M2+H]+
Methanol.2.Water.2 | C2H12O4 | 100.0736 |  |  | CON00042 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 101.0808 | 101.0808 | [M+H]+
Methanol.2.Water.2 | C2H12O4 | 100.0736 |  |  | CON00042 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 123.0628 | 123.0628 | [M+Na]+
Methanol.3.Water.2 | C3H16O5 | 132.0998 |  |  | CON00043 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 133.107 | 133.1071 | [M+H]+
Methanol.3.Water.2 | C3H16O5 | 132.0998 |  |  | CON00043 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 155.089 | 155.089 | [M+Na]+
Methanol.4.Water.3 | C4H22O7 | 182.1366 |  |  | CON00044 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 183.1438 | 183.1438 | [M+H]+
Methanol.4.Water.3 | C4H22O7 | 182.1366 |  |  | CON00044 | Solvent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 205.1258 | 205.1258 | [M+Na]+
Monomethoxytrityl | C20H17O | 273.1279 | InChI=1S/C20H17O/c1-21-19-14-12-18(13-15-19)20(16-8-4-2-5-9-16)17-10-6-3-7-11-17/h2-15H,1H3 | NTSFISBKIKSZLT-UHFFFAOYSA-N | CON00045 |  |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 273 | 273.1274 | [M]+
N,N-Bis(2-hydroxyethyl)dodecanamide | C16H33NO3 | 287.246 | InChI=1S/C16H33NO3/c1-2-3-4-5-6-7-8-9-10-11-16(20)17(12-14-18)13-15-19/h18-19H,2-15H2,1H3 | AOMUHOFOVNGZAN-UHFFFAOYSA-N | CON00046 | Anti-static agent |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 288 | 288.2533 | [M+H]+
N-Butylbenzenesulfonamide | C10H15NO2S | 213.0824 | InChI=1S/C10H15NO2S/c1-2-3-9-11-14(12,13)10-7-5-4-6-8-10/h4-8,11H,2-3,9H2,1H3 | IPRJXAGUEGOFGG-UHFFFAOYSA-N | CON00047 | Plasticiser | 19241 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | NEG | 212.0757 | 212.0751 | [M-H]-
N-Butylbenzenesulfonamide | C10H15NO2S | 213.0824 | InChI=1S/C10H15NO2S/c1-2-3-9-11-14(12,13)10-7-5-4-6-8-10/h4-8,11H,2-3,9H2,1H3 | IPRJXAGUEGOFGG-UHFFFAOYSA-N | CON00047 | Plasticiser | 19241 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 214 | 214.0896 | [M+H]+
N-Butylbenzenesulfonamide | C10H15NO2S | 213.0824 | InChI=1S/C10H15NO2S/c1-2-3-9-11-14(12,13)10-7-5-4-6-8-10/h4-8,11H,2-3,9H2,1H3 | IPRJXAGUEGOFGG-UHFFFAOYSA-N | CON00047 | Plasticiser | 19241 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 231 | 231.1162 | [M+NH4]+
N-Butylbenzenesulfonamide | C10H15NO2S | 213.0824 | InChI=1S/C10H15NO2S/c1-2-3-9-11-14(12,13)10-7-5-4-6-8-10/h4-8,11H,2-3,9H2,1H3 | IPRJXAGUEGOFGG-UHFFFAOYSA-N | CON00047 | Plasticiser | 19241 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 236 | 236.0716 | [M+Na]+
N-Methyl 2-pyrrolidone | C5H9NO | 99.0684 | InChI=1S/C5H9NO/c1-6-4-2-3-5(6)7/h2-4H2,1H3 | SECXISVLQFMRJM-UHFFFAOYSA-N | CON00048 | Solvent | 13387 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 100.0757 | 100.0757 | [M+H]+ | Ion trap | Thermofinnigan LCQ | LC | ESI
Nylon 66 polymer | C24H44N4O4 | 452.3363 |  |  | CON00051 | Filter |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. | POS | 453.4 | 453.3435 | [M+H]+ | Ion trap | Thermofinnigan LCQ | LC | ESI
Nylon 66 polymer | C24H44N4O4 | 452.3363 |  |  | CON00051 | Filter |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. | POS | 605.3 | 605.2923 | [M+K+C2HF3O2]+ | Ion trap | Thermofinnigan LCQ | LC | ESI
Nylon 66 polymer | C36H66N6O6 | 678.5044 |  |  | CON00051 | Filter |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 679 | 679.5117 | [M+H]+
Nylon 66 polymer | C36H66N6O6 | 678.5044 |  |  | CON00051 | Filter |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. | POS | 679.5 | 679.5117 | [M+H]+ | Ion trap | Thermofinnigan LCQ | LC | ESI
Nylon 66 polymer | C36H66N6O6 | 678.5044 |  |  | CON00051 | Filter |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. | POS | 701.8 | 701.4936 | [M+Na]+ | Ion trap | Thermofinnigan LCQ | LC | ESI
Nylon 66 polymer | C36H66N6O6 | 678.5044 |  |  | CON00051 | Filter |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. | POS | 717.4 | 717.4675 | [M+K]+ | Ion trap | Thermofinnigan LCQ | LC | ESI
Nylon 66 polymer | C36H66N6O6 | 678.5044 |  |  | CON00051 | Filter |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. | POS | 831.5 | 831.4604 | [M+K+C2HF3O2]+ | Ion trap | Thermofinnigan LCQ | LC | ESI
Nylon 66 polymer | C36H66N6O6 | 678.5044 |  |  | CON00051 | Filter |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. | POS | 1097.3 | 1098.4098 | [M+K2+C2HF3O2+C2HF3O2+C2HF3O2-H]+ | Ion trap | Thermofinnigan LCQ | LC | ESI
Nylon 66 polymer | C48H88N8O8 | 904.6725 |  |  | CON00051 | Filter |  | Tran, J. C., & Doucette, A. A. (2006). Cyclic polyamide oligomers extracted from nylon 66 membrane filter disks as a source of contamination in liquid chromatography/mass spectrometry. Journal of the American Society for Mass Spectrometry, 17(5), 652-656, doi:10.1016/j.jasms.2006.01.008. | POS | 905.7 | 905.6798 | [M+H]+ | Ion trap | Thermofinnigan LCQ | LC | ESI
Octadecamethylcyclononasiloxane | C18H54Si9O9 | 666.1691 | InChI=1S/C18H54O9Si9/c1-28(2)19-29(3,4)21-31(7,8)23-33(11,12)25-35(15,16)27-36(17,18)26-34(13,14)24-32(9,10)22-30(5,6)20-28/h1-18H3 | ISXOGOLHEGHGQF-UHFFFAOYSA-N | CON00052 | Airborne Contaminant | 11172 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 651.159 | 651.1451 | [M+H-CH4]+
Octadecamethylcyclononasiloxane | C18H54Si9O9 | 666.1691 | InChI=1S/C18H54O9Si9/c1-28(2)19-29(3,4)21-31(7,8)23-33(11,12)25-35(15,16)27-36(17,18)26-34(13,14)24-32(9,10)22-30(5,6)20-28/h1-18H3 | ISXOGOLHEGHGQF-UHFFFAOYSA-N | CON00052 | Airborne Contaminant | 11172 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 667 | 667.1764 | [M+H]+
Octadecamethylcyclononasiloxane | C18H54Si9O9 | 666.1691 | InChI=1S/C18H54O9Si9/c1-28(2)19-29(3,4)21-31(7,8)23-33(11,12)25-35(15,16)27-36(17,18)26-34(13,14)24-32(9,10)22-30(5,6)20-28/h1-18H3 | ISXOGOLHEGHGQF-UHFFFAOYSA-N | CON00052 | Airborne Contaminant | 11172 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 667.159 | 667.1764 | [M+H]+
Octadecamethylcyclononasiloxane | C18H54Si9O9 | 666.1691 | InChI=1S/C18H54O9Si9/c1-28(2)19-29(3,4)21-31(7,8)23-33(11,12)25-35(15,16)27-36(17,18)26-34(13,14)24-32(9,10)22-30(5,6)20-28/h1-18H3 | ISXOGOLHEGHGQF-UHFFFAOYSA-N | CON00052 | Airborne Contaminant | 11172 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 684.159 | 684.203 | [M+NH4]+
Oleamide | C18H35NO | 281.2719 | InChI=1S/C18H35NO/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18(19)20/h9-10H,2-8,11-17H2,1H3,(H2,19,20)/b10-9- | FATBGEAMYMYZAF-KTKRTIGZSA-N | CON00053 | Slip agent | 5283387 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 282 | 282.2791 | [M+H]+ |  |  |  | ESI
Oleamide | C18H35NO | 281.2719 | InChI=1S/C18H35NO/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18(19)20/h9-10H,2-8,11-17H2,1H3,(H2,19,20)/b10-9- | FATBGEAMYMYZAF-KTKRTIGZSA-N | CON00053 | Slip agent | 5283387 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 303 | 304.2611 | [M+Na]+ |  |  |  | ESI
Oleic acid | C18H34O2 | 282.2559 | InChI=1S/C18H34O2/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18(19)20/h9-10H,2-8,11-17H2,1H3,(H,19,20)/b10-9- | ZQPPMHVWECSIRJ-KTKRTIGZSA-N | CON00054 |  | 445639 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | NEG | 281.2486 | 281.2486 | [M-H]-
Palmitic acid | C16H32O2 | 256.2402 | InChI=1S/C16H32O2/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16(17)18/h2-15H2,1H3,(H,17,18) | IPCSVZSSVZVIGE-UHFFFAOYSA-N | CON00055 |  | 985 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | NEG | 255.2329 | 255.233 | [M-H]-
Palmitamidopropyl-trimonium chloride | C22H47N2OCl | 390.3377 |  |  | CON00056 | Personal care products |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 355.3683 | 355.3683 | [M-Cl]+
PEG.1 | C2H6O2 | 62.0368 | InChI=1S/C2H6O2/c3-1-2-4/h3-4H,1-2H2 | LYCAIKOWRPUZTN-UHFFFAOYSA-N | CON00057 | Polyether | 174 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 63.0441 | 63.0441 | [M+H]+
PEG.1 | C2H6O2 | 62.0368 | InChI=1S/C2H6O2/c3-1-2-4/h3-4H,1-2H2 | LYCAIKOWRPUZTN-UHFFFAOYSA-N | CON00057 | Polyether | 174 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 85.026 | 85.026 | [M+Na]+
PEG.1 | C2H6O2 | 62.0368 | InChI=1S/C2H6O2/c3-1-2-4/h3-4H,1-2H2 | LYCAIKOWRPUZTN-UHFFFAOYSA-N | CON00057 | Polyether | 174 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 100.9999 | 100.9999 | [M+K]+
PEG.10 | C20H42O11 | 458.2727 | InChI=1S/C20H42O11/c21-1-3-23-5-7-25-9-11-27-13-15-29-17-19-31-20-18-30-16-14-28-12-10-26-8-6-24-4-2-22/h21-22H,1-20H2 | DTPCFIHYWYONMD-UHFFFAOYSA-N | CON00058 | Polyether | 79689 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 459.28 | 459.28 | [M+H]+
PEG.10 | C20H42O11 | 458.2727 | InChI=1S/C20H42O11/c21-1-3-23-5-7-25-9-11-27-13-15-29-17-19-31-20-18-30-16-14-28-12-10-26-8-6-24-4-2-22/h21-22H,1-20H2 | DTPCFIHYWYONMD-UHFFFAOYSA-N | CON00058 | Polyether | 79689 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 481.2619 | 481.2619 | [M+Na]+
PEG.10 | C20H42O11 | 458.2727 | InChI=1S/C20H42O11/c21-1-3-23-5-7-25-9-11-27-13-15-29-17-19-31-20-18-30-16-14-28-12-10-26-8-6-24-4-2-22/h21-22H,1-20H2 | DTPCFIHYWYONMD-UHFFFAOYSA-N | CON00058 | Polyether | 79689 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 497.2359 | 497.2359 | [M+K]+
PEG.11 | C22H46O12 | 502.2989 | InChI=1S/C22H46O12/c23-1-3-25-5-7-27-9-11-29-13-15-31-17-19-33-21-22-34-20-18-32-16-14-30-12-10-28-8-6-26-4-2-24/h23-24H,1-22H2 | PSVXZQVXSXSQRO-UHFFFAOYSA-N | CON00059 | Polyether | 81260 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 503.3062 | 503.3062 | [M+H]+
PEG.11 | C22H46O12 | 502.2989 | InChI=1S/C22H46O12/c23-1-3-25-5-7-27-9-11-29-13-15-31-17-19-33-21-22-34-20-18-32-16-14-30-12-10-28-8-6-26-4-2-24/h23-24H,1-22H2 | PSVXZQVXSXSQRO-UHFFFAOYSA-N | CON00059 | Polyether | 81260 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 525.2881 | 525.2882 | [M+Na]+
PEG.11 | C22H46O12 | 502.2989 | InChI=1S/C22H46O12/c23-1-3-25-5-7-27-9-11-29-13-15-31-17-19-33-21-22-34-20-18-32-16-14-30-12-10-28-8-6-26-4-2-24/h23-24H,1-22H2 | PSVXZQVXSXSQRO-UHFFFAOYSA-N | CON00059 | Polyether | 81260 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 541.2621 | 541.2621 | [M+K]+
PEG.12 | C24H50O13 | 546.3251 | InChI=1S/C24H50O13/c25-1-3-27-5-7-29-9-11-31-13-15-33-17-19-35-21-23-37-24-22-36-20-18-34-16-14-32-12-10-30-8-6-28-4-2-26/h25-26H,1-24H2 | WRZXKWFJEFFURH-UHFFFAOYSA-N | CON00060 | Polyether | 81248 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 547.3324 | 547.3324 | [M+H]+
PEG.12 | C24H50O13 | 546.3251 | InChI=1S/C24H50O13/c25-1-3-27-5-7-29-9-11-31-13-15-33-17-19-35-21-23-37-24-22-36-20-18-34-16-14-32-12-10-30-8-6-28-4-2-26/h25-26H,1-24H2 | WRZXKWFJEFFURH-UHFFFAOYSA-N | CON00060 | Polyether | 81248 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 569.3144 | 569.3144 | [M+Na]+
PEG.12 | C24H50O13 | 546.3251 | InChI=1S/C24H50O13/c25-1-3-27-5-7-29-9-11-31-13-15-33-17-19-35-21-23-37-24-22-36-20-18-34-16-14-32-12-10-30-8-6-28-4-2-26/h25-26H,1-24H2 | WRZXKWFJEFFURH-UHFFFAOYSA-N | CON00060 | Polyether | 81248 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 585.2883 | 585.2883 | [M+K]+
PEG.13 | C26H54O14 | 590.3514 | InChI=1S/C26H54O14/c27-1-3-29-5-7-31-9-11-33-13-15-35-17-19-37-21-23-39-25-26-40-24-22-38-20-18-36-16-14-34-12-10-32-8-6-30-4-2-28/h27-28H,1-26H2 | AKWFJQNBHYVIPY-UHFFFAOYSA-N | CON00061 | Polyether | 87168 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 591.3586 | 591.3586 | [M+H]+
PEG.13 | C26H54O14 | 590.3514 | InChI=1S/C26H54O14/c27-1-3-29-5-7-31-9-11-33-13-15-35-17-19-37-21-23-39-25-26-40-24-22-38-20-18-36-16-14-34-12-10-32-8-6-30-4-2-28/h27-28H,1-26H2 | AKWFJQNBHYVIPY-UHFFFAOYSA-N | CON00061 | Polyether | 87168 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 613.3406 | 613.3406 | [M+Na]+
PEG.13 | C26H54O14 | 590.3514 | InChI=1S/C26H54O14/c27-1-3-29-5-7-31-9-11-33-13-15-35-17-19-37-21-23-39-25-26-40-24-22-38-20-18-36-16-14-34-12-10-32-8-6-30-4-2-28/h27-28H,1-26H2 | AKWFJQNBHYVIPY-UHFFFAOYSA-N | CON00061 | Polyether | 87168 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 629.3145 | 629.3145 | [M+K]+
PEG.14 | C28H58O15 | 634.3776 | InChI=1S/C28H58O15/c29-1-3-31-5-7-33-9-11-35-13-15-37-17-19-39-21-23-41-25-27-43-28-26-42-24-22-40-20-18-38-16-14-36-12-10-34-8-6-32-4-2-30/h29-30H,1-28H2 | ILLKMACMBHTSHP-UHFFFAOYSA-N | CON00062 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 635.3849 | 635.3849 | [M+H]+
PEG.14 | C28H58O15 | 634.3776 | InChI=1S/C28H58O15/c29-1-3-31-5-7-33-9-11-35-13-15-37-17-19-39-21-23-41-25-27-43-28-26-42-24-22-40-20-18-38-16-14-36-12-10-34-8-6-32-4-2-30/h29-30H,1-28H2 | ILLKMACMBHTSHP-UHFFFAOYSA-N | CON00062 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 657.3668 | 657.3668 | [M+Na]+
PEG.14 | C28H58O15 | 634.3776 | InChI=1S/C28H58O15/c29-1-3-31-5-7-33-9-11-35-13-15-37-17-19-39-21-23-41-25-27-43-28-26-42-24-22-40-20-18-38-16-14-36-12-10-34-8-6-32-4-2-30/h29-30H,1-28H2 | ILLKMACMBHTSHP-UHFFFAOYSA-N | CON00062 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 673.3407 | 673.3407 | [M+K]+
PEG.15 | C30H62O16 | 678.4038 | InChI=1S/C30H62O16/c31-1-3-33-5-7-35-9-11-37-13-15-39-17-19-41-21-23-43-25-27-45-29-30-46-28-26-44-24-22-42-20-18-40-16-14-38-12-10-36-8-6-34-4-2-32/h31-32H,1-30H2 | OWTQQPNDSWCHOV-UHFFFAOYSA-N | CON00063 | Polyether | 120062 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 679.4111 | 679.4111 | [M+H]+
PEG.15 | C30H62O16 | 678.4038 | InChI=1S/C30H62O16/c31-1-3-33-5-7-35-9-11-37-13-15-39-17-19-41-21-23-43-25-27-45-29-30-46-28-26-44-24-22-42-20-18-40-16-14-38-12-10-36-8-6-34-4-2-32/h31-32H,1-30H2 | OWTQQPNDSWCHOV-UHFFFAOYSA-N | CON00063 | Polyether | 120062 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 701.393 | 701.393 | [M+Na]+
PEG.15 | C30H62O16 | 678.4038 | InChI=1S/C30H62O16/c31-1-3-33-5-7-35-9-11-37-13-15-39-17-19-41-21-23-43-25-27-45-29-30-46-28-26-44-24-22-42-20-18-40-16-14-38-12-10-36-8-6-34-4-2-32/h31-32H,1-30H2 | OWTQQPNDSWCHOV-UHFFFAOYSA-N | CON00063 | Polyether | 120062 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 717.3669 | 717.3669 | [M+K]+
PEG.16 | C32H66O17 | 722.43 | InChI=1S/C32H66O17/c33-1-3-35-5-7-37-9-11-39-13-15-41-17-19-43-21-23-45-25-27-47-29-31-49-32-30-48-28-26-46-24-22-44-20-18-42-16-14-40-12-10-38-8-6-36-4-2-34/h33-34H,1-32H2 | DHORSBRLGKJPFC-UHFFFAOYSA-N | CON00064 | Polyether | 81262 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 723.4373 | 723.4373 | [M+H]+
PEG.16 | C32H66O17 | 722.43 | InChI=1S/C32H66O17/c33-1-3-35-5-7-37-9-11-39-13-15-41-17-19-43-21-23-45-25-27-47-29-31-49-32-30-48-28-26-46-24-22-44-20-18-42-16-14-40-12-10-38-8-6-36-4-2-34/h33-34H,1-32H2 | DHORSBRLGKJPFC-UHFFFAOYSA-N | CON00064 | Polyether | 81262 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 745.4192 | 745.4192 | [M+Na]+
PEG.16 | C32H66O17 | 722.43 | InChI=1S/C32H66O17/c33-1-3-35-5-7-37-9-11-39-13-15-41-17-19-43-21-23-45-25-27-47-29-31-49-32-30-48-28-26-46-24-22-44-20-18-42-16-14-40-12-10-38-8-6-36-4-2-34/h33-34H,1-32H2 | DHORSBRLGKJPFC-UHFFFAOYSA-N | CON00064 | Polyether | 81262 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 761.3932 | 761.3932 | [M+K]+
PEG.17 | C34H70O18 | 766.4562 |  |  | CON00065 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 767.4635 | 767.4635 | [M+H]+
PEG.17 | C34H70O18 | 766.4562 |  |  | CON00065 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 789.4454 | 789.4454 | [M+Na]+
PEG.17 | C34H70O18 | 766.4562 |  |  | CON00065 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 805.4194 | 805.4194 | [M+K]+
PEG.18 | C36H74O19 | 810.4824 |  |  | CON00066 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 811.4897 | 811.4897 | [M+H]+
PEG.18 | C36H74O19 | 810.4824 |  |  | CON00066 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 833.4718 | 833.4717 | [M+Na]+
PEG.18 | C36H74O19 | 810.4824 |  |  | CON00066 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 849.4456 | 849.4456 | [M+K]+
PEG.19 | C38H78O20 | 854.5086 |  |  | CON00067 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 855.5159 | 855.5159 | [M+H]+
PEG.19 | C38H78O20 | 854.5086 |  |  | CON00067 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 877.4979 | 877.4979 | [M+Na]+
PEG.19 | C38H78O20 | 854.5086 |  |  | CON00067 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 893.4718 | 893.4718 | [M+K]+
PEG.2 | C4H10O3 | 106.063 | InChI=1S/C4H10O3/c5-1-3-7-4-2-6/h5-6H,1-4H2 | MTHSVFCYNBDYFN-UHFFFAOYSA-N | CON00068 | Polyether | 8117 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 107.0703 | 107.0703 | [M+H]+
PEG.2 | C4H10O3 | 106.063 | InChI=1S/C4H10O3/c5-1-3-7-4-2-6/h5-6H,1-4H2 | MTHSVFCYNBDYFN-UHFFFAOYSA-N | CON00068 | Polyether | 8117 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 129.0522 | 129.0522 | [M+Na]+
PEG.2 | C4H10O3 | 106.063 | InChI=1S/C4H10O3/c5-1-3-7-4-2-6/h5-6H,1-4H2 | MTHSVFCYNBDYFN-UHFFFAOYSA-N | CON00068 | Polyether | 8117 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 145.0262 | 145.0262 | [M+K]+
PEG.20 | C40H82O21 | 898.5349 | InChI=1S/C40H82O21/c41-1-3-43-5-7-45-9-11-47-13-15-49-17-19-51-21-23-53-25-27-55-29-31-57-33-35-59-37-39-61-40-38-60-36-34-58-32-30-56-28-26-54-24-22-52-20-18-50-16-14-48-12-10-46-8-6-44-4-2-42/h41-42H,1-40H2 | ZHMUMDLGQBRJIW-UHFFFAOYSA-N | CON00069 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 899.5421 | 899.5421 | [M+H]+
PEG.20 | C40H82O21 | 898.5349 | InChI=1S/C40H82O21/c41-1-3-43-5-7-45-9-11-47-13-15-49-17-19-51-21-23-53-25-27-55-29-31-57-33-35-59-37-39-61-40-38-60-36-34-58-32-30-56-28-26-54-24-22-52-20-18-50-16-14-48-12-10-46-8-6-44-4-2-42/h41-42H,1-40H2 | ZHMUMDLGQBRJIW-UHFFFAOYSA-N | CON00069 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 921.5241 | 921.5241 | [M+Na]+
PEG.20 | C40H82O21 | 898.5349 | InChI=1S/C40H82O21/c41-1-3-43-5-7-45-9-11-47-13-15-49-17-19-51-21-23-53-25-27-55-29-31-57-33-35-59-37-39-61-40-38-60-36-34-58-32-30-56-28-26-54-24-22-52-20-18-50-16-14-48-12-10-46-8-6-44-4-2-42/h41-42H,1-40H2 | ZHMUMDLGQBRJIW-UHFFFAOYSA-N | CON00069 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 937.498 | 937.498 | [M+K]+
PEG.3 | C6H14O4 | 150.0892 | InChI=1S/C6H14O4/c7-1-3-9-5-6-10-4-2-8/h7-8H,1-6H2 | ZIBGPFATKBEMQZ-UHFFFAOYSA-N | CON00070 | Polyether | 8172 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 151.0965 | 151.0965 | [M+H]+
PEG.3 | C6H14O4 | 150.0892 | InChI=1S/C6H14O4/c7-1-3-9-5-6-10-4-2-8/h7-8H,1-6H2 | ZIBGPFATKBEMQZ-UHFFFAOYSA-N | CON00070 | Polyether | 8172 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 173.0784 | 173.0784 | [M+Na]+
PEG.3 | C6H14O4 | 150.0892 | InChI=1S/C6H14O4/c7-1-3-9-5-6-10-4-2-8/h7-8H,1-6H2 | ZIBGPFATKBEMQZ-UHFFFAOYSA-N | CON00070 | Polyether | 8172 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 189.0524 | 189.0524 | [M+K]+
PEG.4 | C8H18O5 | 194.1154 | InChI=1S/C8H18O5/c9-1-3-11-5-7-13-8-6-12-4-2-10/h9-10H,1-8H2 | UWHCKJMYHZGTIT-UHFFFAOYSA-N | CON00071 | Polyether | 8200 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 195.1127 | 195.1227 | [M+H]+
PEG.4 | C8H18O5 | 194.1154 | InChI=1S/C8H18O5/c9-1-3-11-5-7-13-8-6-12-4-2-10/h9-10H,1-8H2 | UWHCKJMYHZGTIT-UHFFFAOYSA-N | CON00071 | Polyether | 8200 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 217.1046 | 217.1046 | [M+Na]+
PEG.4 | C8H18O5 | 194.1154 | InChI=1S/C8H18O5/c9-1-3-11-5-7-13-8-6-12-4-2-10/h9-10H,1-8H2 | UWHCKJMYHZGTIT-UHFFFAOYSA-N | CON00071 | Polyether | 8200 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 233.0786 | 233.0786 | [M+K]+
PEG.5 | C10H22O6 | 238.1416 | InChI=1S/C10H22O6/c11-1-3-13-5-7-15-9-10-16-8-6-14-4-2-12/h11-12H,1-10H2 | JLFNLZLINWHATN-UHFFFAOYSA-N | CON00072 | Polyether | 62551 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 239.1489 | 239.1489 | [M+H]+
PEG.5 | C10H22O6 | 238.1416 | InChI=1S/C10H22O6/c11-1-3-13-5-7-15-9-10-16-8-6-14-4-2-12/h11-12H,1-10H2 | JLFNLZLINWHATN-UHFFFAOYSA-N | CON00072 | Polyether | 62551 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 261.1309 | 261.1309 | [M+Na]+
PEG.5 | C10H22O6 | 238.1416 | InChI=1S/C10H22O6/c11-1-3-13-5-7-15-9-10-16-8-6-14-4-2-12/h11-12H,1-10H2 | JLFNLZLINWHATN-UHFFFAOYSA-N | CON00072 | Polyether | 62551 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 277.1048 | 277.1048 | [M+K]+
PEG.6 | C12H26O7 | 282.1679 | InChI=1S/C12H26O7/c13-1-3-15-5-7-17-9-11-19-12-10-18-8-6-16-4-2-14/h13-14H,1-12H2 | IIRDTKBZINWQAW-UHFFFAOYSA-N | CON00073 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 283.1751 | 283.1751 | [M+H]+
PEG.6 | C12H26O7 | 282.1679 | InChI=1S/C12H26O7/c13-1-3-15-5-7-17-9-11-19-12-10-18-8-6-16-4-2-14/h13-14H,1-12H2 | IIRDTKBZINWQAW-UHFFFAOYSA-N | CON00073 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 305.1571 | 305.1571 | [M+Na]+
PEG.6 | C12H26O7 | 282.1679 | InChI=1S/C12H26O7/c13-1-3-15-5-7-17-9-11-19-12-10-18-8-6-16-4-2-14/h13-14H,1-12H2 | IIRDTKBZINWQAW-UHFFFAOYSA-N | CON00073 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 321.131 | 321.131 | [M+K]+
PEG.7 | C14H30O8 | 326.1941 | InChI=1S/C14H30O8/c15-1-3-17-5-7-19-9-11-21-13-14-22-12-10-20-8-6-18-4-2-16/h15-16H,1-14H2 | XPJRQAIZZQMSCM-UHFFFAOYSA-N | CON00074 | Polyether | 79718 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 365.1572 | 365.1572 | [M+K]+
PEG.7 | C14H30O8 | 326.1941 | InChI=1S/C14H30O8/c15-1-3-17-5-7-19-9-11-21-13-14-22-12-10-20-8-6-18-4-2-16/h15-16H,1-14H2 | XPJRQAIZZQMSCM-UHFFFAOYSA-N | CON00074 | Polyether | 79718 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 327.2013 | 327.2013 | [M+H]+
PEG.7 | C14H30O8 | 326.1941 | InChI=1S/C14H30O8/c15-1-3-17-5-7-19-9-11-21-13-14-22-12-10-20-8-6-18-4-2-16/h15-16H,1-14H2 | XPJRQAIZZQMSCM-UHFFFAOYSA-N | CON00074 | Polyether | 79718 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 349.1833 | 349.1833 | [M+Na]+
PEG.8 | C16H34O9 | 370.2203 | InChI=1S/C16H34O9/c17-1-3-19-5-7-21-9-11-23-13-15-25-16-14-24-12-10-22-8-6-20-4-2-18/h17-18H,1-16H2 | GLZWNFNQMJAZGY-UHFFFAOYSA-N | CON00075 | Polyether | 78798 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 371.2276 | 371.2276 | [M+H]+
PEG.8 | C16H34O9 | 370.2203 | InChI=1S/C16H34O9/c17-1-3-19-5-7-21-9-11-23-13-15-25-16-14-24-12-10-22-8-6-20-4-2-18/h17-18H,1-16H2 | GLZWNFNQMJAZGY-UHFFFAOYSA-N | CON00075 | Polyether | 78798 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 393.2095 | 393.2095 | [M+Na]+
PEG.8 | C16H34O9 | 370.2203 | InChI=1S/C16H34O9/c17-1-3-19-5-7-21-9-11-23-13-15-25-16-14-24-12-10-22-8-6-20-4-2-18/h17-18H,1-16H2 | GLZWNFNQMJAZGY-UHFFFAOYSA-N | CON00075 | Polyether | 78798 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 409.1834 | 409.1834 | [M+K]+
PEG.9 | C18H38O10 | 414.2465 | InChI=1S/C18H38O10/c19-1-3-21-5-7-23-9-11-25-13-15-27-17-18-28-16-14-26-12-10-24-8-6-22-4-2-20/h19-20H,1-18H2 | YZUUTMGDONTGTN-UHFFFAOYSA-N | CON00076 | Polyether | 4867 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 415.2538 | 415.2538 | [M+H]+
PEG.9 | C18H38O10 | 414.2465 | InChI=1S/C18H38O10/c19-1-3-21-5-7-23-9-11-25-13-15-27-17-18-28-16-14-26-12-10-24-8-6-22-4-2-20/h19-20H,1-18H2 | YZUUTMGDONTGTN-UHFFFAOYSA-N | CON00076 | Polyether | 4867 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 437.2357 | 437.2357 | [M+Na]+
PEG.9 | C18H38O10 | 414.2465 | InChI=1S/C18H38O10/c19-1-3-21-5-7-23-9-11-25-13-15-27-17-18-28-16-14-26-12-10-24-8-6-22-4-2-20/h19-20H,1-18H2 | YZUUTMGDONTGTN-UHFFFAOYSA-N | CON00076 | Polyether | 4867 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 453.2097 | 453.2097 | [M+K]+
Pentafluoropropionic acid | C3HF5O2 | 163.9897 | InChI=1S/C3HF5O2/c4-2(5,1(9)10)3(6,7)8/h(H,9,10) | LRMSQVBRUNSOJL-UHFFFAOYSA-N | CON00077 |  | 62356 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | NEG | 162.9824 | 162.9824 | [M-H]-
Phenyldiethylamine | C10H15N | 149.1204 | InChI=1S/C10H15N/c1-3-11(4-2)10-8-6-5-7-9-10/h5-9H,3-4H2,1-2H3 | GGSUCNLOZRCGPQ-UHFFFAOYSA-N | CON00078 | Solvent | 7061 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 150 | 150.1277 | [M+H]+
Phosphoric acid | H3PO4 | 97.9769 | InChI=1S/H3O4P/c1-5(2,3)4/h(H3,1,2,3,4) | NBIIXXVUZAFLBC-UHFFFAOYSA-N | CON00079 |  | 1004 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | NEG | 79 | 78.9591 | [M-H3O]-
Phosphoric acid | H3PO4 | 97.9769 | InChI=1S/H3O4P/c1-5(2,3)4/h(H3,1,2,3,4) | NBIIXXVUZAFLBC-UHFFFAOYSA-N | CON00079 |  |  | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | NEG | 97 | 96.9696 | [M-H]-
Phthalic acid | C8H6O4 | 166.0266 | InChI=1S/C8H6O4/c9-7(10)5-3-1-2-4-6(5)8(11)12/h1-4H,(H,9,10)(H,11,12) | XNGIFLGASWRNHJ-UHFFFAOYSA-N | CON00080 |  | 1017 | Guo, X. H., Bruins, A. P., & Covey, T. R. (2006). Characterization of typical chemical background interferences in atmospheric pressure ionization liquid chromatography-mass spectrometry. Rapid Communications in Mass Spectrometry, 20(20), 3145-3150, doi:10.1002/rcm.2715. | POS | 167 | 167.0339 | [M+H]+ | Triple quadrupole | API366 | LC
Phthalic anhydride | C8H4O3 | 148.016 | InChI=1S/C8H4O3/c9-7-5-3-1-2-4-6(5)8(10)11-7/h1-4H | LGRFSURHDFAFJT-UHFFFAOYSA-N | CON00081 | Plasticiser | 6811 | Guo, X. H., Bruins, A. P., & Covey, T. R. (2006). Characterization of typical chemical background interferences in atmospheric pressure ionization liquid chromatography-mass spectrometry. Rapid Communications in Mass Spectrometry, 20(20), 3145-3150, doi:10.1002/rcm.2715. | POS | 149 | 149.0233 | [M+H]+ | Triple quadrupole | API365 | LC
Phthalic anhydride | C8H4O3 | 148.016 | InChI=1S/C8H4O3/c9-7-5-3-1-2-4-6(5)8(10)11-7/h1-4H | LGRFSURHDFAFJT-UHFFFAOYSA-N | CON00081 | Plasticiser | 6811 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 149.0233 | 149.0233 | [M+H]+
Phthalic anhydride | C8H4O3 | 148.016 | InChI=1S/C8H4O3/c9-7-5-3-1-2-4-6(5)8(10)11-7/h1-4H | LGRFSURHDFAFJT-UHFFFAOYSA-N | CON00081 | Plasticiser | 6811 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 171 | 171.0053 | [M+Na]+
Phthalic anhydride | C8H4O3 | 148.016 | InChI=1S/C8H4O3/c9-7-5-3-1-2-4-6(5)8(10)11-7/h1-4H | LGRFSURHDFAFJT-UHFFFAOYSA-N | CON00081 | Plasticiser | 6811 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 171.0053 | 171.0053 | [M+Na]+
Phthalic anhydride | C8H4O3 | 148.016 | InChI=1S/C8H4O3/c9-7-5-3-1-2-4-6(5)8(10)11-7/h1-4H | LGRFSURHDFAFJT-UHFFFAOYSA-N | CON00081 | Plasticiser | 6811 | Guo, X. H., Bruins, A. P., & Covey, T. R. (2006). Characterization of typical chemical background interferences in atmospheric pressure ionization liquid chromatography-mass spectrometry. Rapid Communications in Mass Spectrometry, 20(20), 3145-3150, doi:10.1002/rcm.2715. | POS | 205 | 204.0781 | [M+C4H8]+ | Triple quadrupole | API365 | LC
PPG.1 | C3H8O2 | 76.0524 | InChI=1S/C3H8O2/c1-3(5)2-4/h3-5H,2H2,1H3 | DNIAPMSPPWPWGF-UHFFFAOYSA-N | CON00082 | Polyether | 439846 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 77.0597 | 77.0597 | [M+H]+
PPG.1 | C3H8O2 | 76.0524 | InChI=1S/C3H8O2/c1-3(5)2-4/h3-5H,2H2,1H3 | DNIAPMSPPWPWGF-UHFFFAOYSA-N | CON00082 | Polyether | 439846 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 99.0416 | 99.0417 | [M+Na]+
PPG.1 | C3H8O2 | 76.0524 | InChI=1S/C3H8O2/c1-3(5)2-4/h3-5H,2H2,1H3 | DNIAPMSPPWPWGF-UHFFFAOYSA-N | CON00082 | Polyether | 439846 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 115.0156 | 115.0156 | [M+K]+
PPG.10 | C30H62O11 | 598.4292 |  |  | CON00083 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 599.4365 | 599.4365 | [M+H]+
PPG.10 | C30H62O11 | 598.4292 |  |  | CON00083 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 621.4184 | 621.4184 | [M+Na]+
PPG.10 | C30H62O11 | 598.4292 |  |  | CON00083 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 637.3924 | 637.3924 | [M+K]+
PPG.11 | C33H68O12 | 656.4711 |  |  | CON00084 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 657.4784 | 657.4784 | [M+H]+
PPG.11 | C33H68O12 | 656.4711 |  |  | CON00084 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 679.4603 | 679.4603 | [M+Na]+
PPG.11 | C33H68O12 | 656.4711 |  |  | CON00084 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 695.4342 | 695.4342 | [M+K]+
PPG.12 | C36H74O13 | 714.5129 |  |  | CON00085 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 715.5202 | 715.5202 | [M+H]+
PPG.12 | C36H74O13 | 714.5129 |  |  | CON00085 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 737.5022 | 737.5022 | [M+Na]+
PPG.12 | C36H74O13 | 714.5129 |  |  | CON00085 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 753.4761 | 753.4761 | [M+K]+
PPG.13 | C39H80O14 | 772.5548 |  |  | CON00086 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 773.5621 | 773.5621 | [M+H]+
PPG.13 | C39H80O14 | 772.5548 |  |  | CON00086 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 795.544 | 795.544 | [M+Na]+
PPG.13 | C39H80O14 | 772.5548 |  |  | CON00086 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 811.518 | 811.518 | [M+K]+
PPG.14 | C42H86O15 | 830.5967 |  |  | CON00087 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 831.604 | 831.604 | [M+H]+
PPG.14 | C42H86O15 | 830.5967 |  |  | CON00087 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 853.5859 | 853.5859 | [M+Na]+
PPG.14 | C42H86O15 | 830.5967 |  |  | CON00087 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 869.5598 | 869.5598 | [M+K]+
PPG.15 | C45H92O16 | 888.6385 |  |  | CON00088 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 889.6458 | 889.6458 | [M+H]+
PPG.15 | C45H92O16 | 888.6385 |  |  | CON00088 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 911.6278 | 911.6278 | [M+Na]+
PPG.15 | C45H92O16 | 888.6385 |  |  | CON00088 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 927.6017 | 927.6017 | [M+K]+
PPG.16 | C48H98O17 | 946.6804 |  |  | CON00089 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 969.6696 | 969.6696 | [M+Na]+
PPG.16 | C48H98O17 | 946.6804 |  |  | CON00089 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 947.6877 | 947.6877 | [M+H]+
PPG.16 | C48H98O17 | 946.6804 |  |  | CON00089 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 985.6236 | 985.6436 | [M+K]+
PPG.17 | C51H104O18 | 1004.7223 |  |  | CON00090 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1005.7295 | 1005.7295 | [M+H]+
PPG.17 | C51H104O18 | 1004.7223 |  |  | CON00090 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1027.7115 | 1027.7115 | [M+Na]+
PPG.17 | C51H104O18 | 1004.7223 |  |  | CON00090 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1043.6854 | 1043.6854 | [M+K]+
PPG.18 | C54H110O19 | 1062.7641 |  |  | CON00091 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1063.7714 | 1063.7714 | [M+H]+
PPG.18 | C54H110O19 | 1062.7641 |  |  | CON00091 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1085.7534 | 1085.7534 | [M+Na]+
PPG.18 | C54H110O19 | 1062.7641 |  |  | CON00091 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1101.7273 | 1101.7273 | [M+K]+
PPG.19 | C57H116O20 | 1120.806 |  |  | CON00092 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1121.8133 | 1121.8133 | [M+H]+
PPG.19 | C57H116O20 | 1120.806 |  |  | CON00092 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1143.7952 | 1143.7952 | [M+Na]+
PPG.19 | C57H116O20 | 1120.806 |  |  | CON00092 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1159.7692 | 1159.7692 | [M+K]+
PPG.20 | C60H122O21 | 1178.8479 |  |  | CON00093 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1179.8551 | 1179.8551 | [M+H]+
PPG.20 | C60H122O21 | 1178.8479 |  |  | CON00093 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1201.8371 | 1201.8371 | [M+Na]+
PPG.20 | C60H122O21 | 1178.8479 |  |  | CON00093 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1217.811 | 1217.811 | [M+K]+
PPG.2 | C6H14O3 | 134.0943 | InChI=1S/C6H14O3/c1-5(8)4-9-6(2)3-7/h5-8H,3-4H2,1-2H3 | DUFKCOQISQKSAV-UHFFFAOYSA-N | CON00094 | Polyether | 32881 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 135.1016 | 135.1016 | [M+H]+
PPG.2 | C6H14O3 | 134.0943 | InChI=1S/C6H14O3/c1-5(8)4-9-6(2)3-7/h5-8H,3-4H2,1-2H3 | DUFKCOQISQKSAV-UHFFFAOYSA-N | CON00094 | Polyether | 32881 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 157.0835 | 157.0835 | [M+Na]+
PPG.2 | C6H14O3 | 134.0943 | InChI=1S/C6H14O3/c1-5(8)4-9-6(2)3-7/h5-8H,3-4H2,1-2H3 | DUFKCOQISQKSAV-UHFFFAOYSA-N | CON00094 | Polyether | 32881 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 173.0575 | 173.0575 | [M+K]+
PPG.3 | C9H20O4 | 192.1362 | InChI=1S/C9H20O4/c1-7(11)5-12-9(3)6-13-8(2)4-10/h7-11H,4-6H2,1-3H3 | LCZVSXRMYJUNFX-UHFFFAOYSA-N | CON00095 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 193.1434 | 193.1434 | [M+H]+
PPG.3 | C9H20O4 | 192.1362 | InChI=1S/C9H20O4/c1-7(11)5-12-9(3)6-13-8(2)4-10/h7-11H,4-6H2,1-3H3 | LCZVSXRMYJUNFX-UHFFFAOYSA-N | CON00095 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 215.1254 | 215.1254 | [M+Na]+
PPG.3 | C9H20O4 | 192.1362 | InChI=1S/C9H20O4/c1-7(11)5-12-9(3)6-13-8(2)4-10/h7-11H,4-6H2,1-3H3 | LCZVSXRMYJUNFX-UHFFFAOYSA-N | CON00095 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 231.0993 | 231.0993 | [M+K]+
PPG.4 | C12H26O5 | 250.178 | InChI=1S/C12H26O5/c1-9(14)6-15-11(3)8-17-12(4)7-16-10(2)5-13/h9-14H,5-8H2,1-4H3 | QVHMSMOUDQXMRS-UHFFFAOYSA-N | CON00096 | Polyether | 91234 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 251.1853 | 251.1853 | [M+H]+
PPG.4 | C12H26O5 | 250.178 | InChI=1S/C12H26O5/c1-9(14)6-15-11(3)8-17-12(4)7-16-10(2)5-13/h9-14H,5-8H2,1-4H3 | QVHMSMOUDQXMRS-UHFFFAOYSA-N | CON00096 | Polyether | 91234 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 273.1672 | 273.1672 | [M+Na]+
PPG.4 | C12H26O5 | 250.178 | InChI=1S/C12H26O5/c1-9(14)6-15-11(3)8-17-12(4)7-16-10(2)5-13/h9-14H,5-8H2,1-4H3 | QVHMSMOUDQXMRS-UHFFFAOYSA-N | CON00096 | Polyether | 91234 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 289.1412 | 289.1412 | [M+K]+
PPG.5 | C15H32O6 | 308.2199 | InChI=1S/C15H32O6/c1-11(17)7-18-13(3)9-20-15(5)10-21-14(4)8-19-12(2)6-16/h11-17H,6-10H2,1-5H3 | AQRQHYITOOVBTO-UHFFFAOYSA-N | CON00097 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 347.183 | 347.183 | [M+K]+
PPG.5 | C15H32O6 | 308.2199 | InChI=1S/C15H32O6/c1-11(17)7-18-13(3)9-20-15(5)10-21-14(4)8-19-12(2)6-16/h11-17H,6-10H2,1-5H3 | AQRQHYITOOVBTO-UHFFFAOYSA-N | CON00097 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 309.2272 | 309.2272 | [M+H]+
PPG.5 | C15H32O6 | 308.2199 | InChI=1S/C15H32O6/c1-11(17)7-18-13(3)9-20-15(5)10-21-14(4)8-19-12(2)6-16/h11-17H,6-10H2,1-5H3 | AQRQHYITOOVBTO-UHFFFAOYSA-N | CON00097 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 331.2091 | 331.2091 | [M+Na]+
PPG.6 | C18H38O7 | 366.2618 |  |  | CON00098 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 367.269 | 367.269 | [M+H]+
PPG.6 | C18H38O7 | 366.2618 |  |  | CON00098 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 389.251 | 389.251 | [M+Na]+
PPG.6 | C18H38O7 | 366.2618 |  |  | CON00098 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 405.2249 | 405.2249 | [M+K]+
PPG.7 | C21H44O8 | 424.3036 | InChI=1S/C21H44O8/c1-15(22)8-24-10-17(3)26-12-19(5)28-14-21(7)29-13-20(6)27-11-18(4)25-9-16(2)23/h15-23H,8-14H2,1-7H3/t15-,16+,17+,18-,19+,20-,21+/m1/s1 |  | CON00099 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 425.3109 | 425.3109 | [M+H]+
PPG.7 | C21H44O8 | 424.3036 | InChI=1S/C21H44O8/c1-15(22)8-24-10-17(3)26-12-19(5)28-14-21(7)29-13-20(6)27-11-18(4)25-9-16(2)23/h15-23H,8-14H2,1-7H3/t15-,16+,17+,18-,19+,20-,21+/m1/s1 |  | CON00099 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 447.2928 | 447.2928 | [M+Na]+
PPG.7 | C21H44O8 | 424.3036 | InChI=1S/C21H44O8/c1-15(22)8-24-10-17(3)26-12-19(5)28-14-21(7)29-13-20(6)27-11-18(4)25-9-16(2)23/h15-23H,8-14H2,1-7H3/t15-,16+,17+,18-,19+,20-,21+/m1/s1 |  | CON00099 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 463.2668 | 463.2668 | [M+K]+
PPG.8 | C24H50O9 | 482.3455 |  |  | CON00100 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 483.3528 | 483.3528 | [M+H]+
PPG.8 | C24H50O9 | 482.3455 |  |  | CON00100 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 505.3347 | 505.3347 | [M+Na]+
PPG.8 | C24H50O9 | 482.3455 |  |  | CON00100 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 521.3086 | 521.3086 | [M+K]+
PPG.9 | C27H56O10 | 540.3873 |  |  | CON00101 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 541.3946 | 541.3946 | [M+H]+
PPG.9 | C27H56O10 | 540.3873 |  |  | CON00101 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 563.3766 | 563.3766 | [M+Na]+
PPG.9 | C27H56O10 | 540.3873 |  |  | CON00101 | Polyether |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 579.3505 | 579.3505 | [M+K]+
Propionic acid | C3H6O2 | 74.0368 | InChI=1S/C3H6O2/c1-2-3(4)5/h2H2,1H3,(H,4,5) | XBDQKXXYIPTUBI-UHFFFAOYSA-N | CON00102 |  | 1032 | Ijames, C. F., Dutky, R. C., & Fales, H. M. (1995). Iron carboxylate oxygen-centered-triangle complexes detected during electrospray use of organic acid modifiers with a comment on the Finnigan TSQ-700 electrospray inlet system. Journal of the American Society for Mass Spectrometry, 6(12), 1226-1231, doi:10.1016/1044-0305(95)00579-x. | POS | 621.9722 | 621.9729 | [M6-H6+Fe3+O]+ | Triple quadrupole | Finnigan TSQ-700 |  | ESI
Propionic acid | C3H6O2 | 74.0368 | InChI=1S/C3H6O2/c1-2-3(4)5/h2H2,1H3,(H,4,5) | XBDQKXXYIPTUBI-UHFFFAOYSA-N | CON00102 |  | 1032 | Ijames, C. F., Dutky, R. C., & Fales, H. M. (1995). Iron carboxylate oxygen-centered-triangle complexes detected during electrospray use of organic acid modifiers with a comment on the Finnigan TSQ-700 electrospray inlet system. Journal of the American Society for Mass Spectrometry, 6(12), 1226-1231, doi:10.1016/1044-0305(95)00579-x. | POS | 696 | 696.0097 | [M7-H6+Fe3+O]+ | Triple quadrupole | Finnigan TSQ-700 |  | ESI
Propionic acid | C3H6O2 | 74.0368 | InChI=1S/C3H6O2/c1-2-3(4)5/h2H2,1H3,(H,4,5) | XBDQKXXYIPTUBI-UHFFFAOYSA-N | CON00102 |  | 1032 | Ijames, C. F., Dutky, R. C., & Fales, H. M. (1995). Iron carboxylate oxygen-centered-triangle complexes detected during electrospray use of organic acid modifiers with a comment on the Finnigan TSQ-700 electrospray inlet system. Journal of the American Society for Mass Spectrometry, 6(12), 1226-1231, doi:10.1016/1044-0305(95)00579-x. | POS | 770 | 770.0465 | [M8-H6+Fe3+O]+ | Triple quadrupole | Finnigan TSQ-700 |  | ESI
Sodium acetate | C2H3O2Na | 82.0031 | InChI=1S/C2H4O2.Na/c1-2(3)4;/h1H3,(H,3,4);/q;+1/p-1 | VMHLLURERBWHNL-UHFFFAOYSA-M | CON00103 | Solvent | 517045 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 104.9923 | 104.9923 | [M+Na]+
Sodium trifluoroacetate | C2F3O2Na | 135.9748 | InChI=1S/C2HF3O2.Na/c3-2(4,5)1(6)7;/h(H,6,7);/q;+1/p-1 | UYCAUPASBSROMS-UHFFFAOYSA-M | CON00104 | Salt | 51709 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 159 | 158.964 | [M+Na]+
Stearamide | C18H37NO | 283.2875 | InChI=1S/C18H37NO/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18(19)20/h2-17H2,1H3,(H2,19,20) | LYRFLYHAGKPMFH-UHFFFAOYSA-N | CON00105 | Slip agent | 31292 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 284 | 284.2948 | [M+H]+
Stearamide | C18H37NO | 283.2875 | InChI=1S/C18H37NO/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18(19)20/h2-17H2,1H3,(H2,19,20) | LYRFLYHAGKPMFH-UHFFFAOYSA-N | CON00105 | Slip agent | 31292 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 306.2767 | 306.2767 | [M+Na]+
Stearic acid | C18H36O2 | 284.2715 | InChI=1S/C18H36O2/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18(19)20/ | QIQXTHQIDYTFRH-UHFFFAOYSA-N | CON00106 |  | 5281 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | NEG | 283.2643 | 283.2643 | [M-H]-
Stearyl-palmityldimethylammonium chloride | C36H76NCl | 557.5666 | h2-17H2,1H3,(H,19,20) | UKAFOQJMXIZSMD-UHFFFAOYSA-M | CON00107 | Personal care products | 161744 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 522.5972 | 522.5972 | [M-Cl]+
Sulphuric acid | H2SO4 | 97.9674 | InChI=1S/H2O4S/c1-5(2,3)4/h(H2,1,2,3,4) | QAOWNCQODCNURD-UHFFFAOYSA-N | CON00108 |  | 1118 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | NEG | 80 | 79.9574 | [M-H2O]-
Tetrabutylammonium (TBA) | C16H36N | 242.2848 | InChI=1S/C16H36N/c1-5-9-13-17(14-10-6-2,15-11-7-3)16-12-8-4/h5-16H2,1-4H3/q+1 | DZLFLBLQUQXARW-UHFFFAOYSA-N | CON00109 | Buffer | 70681 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 242.2842 | 242.2842 | [M]+
Tetradecamethylcycloheptasiloxane | C14H42Si7O7 | 518.1315 | InChI=1S/C14H42O7Si7/c1-22(2)15-23(3,4)17-25(7,8)19-27(11,12)21-28(13,14)20-26(9,10)18-24(5,6)16-22/h1-14H3 | GSANOGQCVHBHIF-UHFFFAOYSA-N | CON00110 | Airborne Contaminant | 7874 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 519 | 519.1388 | [M+H]+
Tetradecamethylcycloheptasiloxane | C14H42Si7O7 | 518.1315 | InChI=1S/C14H42O7Si7/c1-22(2)15-23(3,4)17-25(7,8)19-27(11,12)21-28(13,14)20-26(9,10)18-24(5,6)16-22/h1-14H3 | GSANOGQCVHBHIF-UHFFFAOYSA-N | CON00110 | Airborne Contaminant | 7874 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 519.127 | 519.1388 | [M+H]+
Tetradecamethylcycloheptasiloxane | C14H42Si7O7 | 518.1315 | InChI=1S/C14H42O7Si7/c1-22(2)15-23(3,4)17-25(7,8)19-27(11,12)21-28(13,14)20-26(9,10)18-24(5,6)16-22/h1-14H3 | GSANOGQCVHBHIF-UHFFFAOYSA-N | CON00110 | Airborne Contaminant | 7874 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 503 | 503.1075 | [M+H-CH4]+
Tetradecamethylcycloheptasiloxane | C14H42Si7O7 | 518.1315 | InChI=1S/C14H42O7Si7/c1-22(2)15-23(3,4)17-25(7,8)19-27(11,12)21-28(13,14)20-26(9,10)18-24(5,6)16-22/h1-14H3 | GSANOGQCVHBHIF-UHFFFAOYSA-N | CON00110 | Airborne Contaminant | 7874 | Schlosser, A., & Volkmer-Engert, R. (2003). Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in nanoelectrospray mass spectrometry. Journal of Mass Spectrometry, 38(5), 523-525, doi:10.1002/jms.465. | POS | 536.127 | 536.1654 | [M+NH4]+
Tributyl Tin Formate | C13H28O2Sn | 336.1111 | InChI=1S/3C4H9.CH2O2.Sn/c3*1-3-4-2;2-1-3;/h3*1,3-4H2,2H3;1H,(H,2,3);/q;;;;+1/p-1 | POUGAXCXKDLEMK-UHFFFAOYSA-M | CON00111 | Catalyst | 22167734 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 336 | 337.1184 | [M+H]+
Tributylamine | C12H27N | 185.2143 | InChI=1S/C12H27N/c1-4-7-10-13(11-8-5-2)12-9-6-3/h4-12H2,1-3H3 | IMFACGCPASFAPR-UHFFFAOYSA-N | CON00112 | Solvent | 7622 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 186 | 186.2216 | [M+H]+
Tributylphosphate | C12H27O4P | 266.1647 | InChI=1S/C12H27O4P/c1-4-7-10-14-17(13,15-11-8-5-2)16-12-9-6-3/h4-12H2,1-3H3 | STCOOQWBFONSKY-UHFFFAOYSA-N | CON00113 | Plasticiser | 31357 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 267 | 267.172 | [M+H]+
Triethylamine | C6H15N | 101.1204 | InChI=1S/C6H15N/c1-4-7(5-2)6-3/h4-6H2,1-3H3 | ZMANZCXQSJIPKH-UHFFFAOYSA-N | CON00114 | Buffer | 8471 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 102 | 102.1277 | [M+H]+
Triethylamine Hydrochloride (35) | C6H15N | 101.1204 | InChI=1S/C6H15N/c1-4-7(5-2)6-3/h4-6H2,1-3H3 | ZMANZCXQSJIPKH-UHFFFAOYSA-N | CON00115 | Buffer | 8471 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 239 | 239.2249 | [M2+H2+35Cl2-35Cl]+
Triethylamine Hydrochloride (37) | C6H15N | 101.1204 | InChI=1S/C6H15N/c1-4-7(5-2)6-3/h4-6H2,1-3H3 | ZMANZCXQSJIPKH-UHFFFAOYSA-N | CON00116 | Buffer | 8471 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 241 | 241.2219 | [M2+H2+37Cl2-37Cl]+
Trifluoroacetic acid | C2HO2F3 | 113.9929 | InChI=1S/C2HF3O2/c3-2(4,5)1(6)7/h(H,6,7) | DTQVDTLACAAQTR-UHFFFAOYSA-N | CON00117 |  | 6422 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | NEG | 227 | 226.9785 | [M2-H]-
Trifluoroacetic acid | C2HO2F3 | 113.9929 | InChI=1S/C2HF3O2/c3-2(4,5)1(6)7/h(H,6,7) | DTQVDTLACAAQTR-UHFFFAOYSA-N | CON00117 |  | 6422 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | NEG | 249 | 248.9604 | [M2+Na-H2]-
Trifluoroacetic acid | C2HO2F3 | 113.9929 | InChI=1S/C2HF3O2/c3-2(4,5)1(6)7/h(H,6,7) | DTQVDTLACAAQTR-UHFFFAOYSA-N | CON00117 |  | 6422 | Tong, H., Bell, D., Tabei, K., & Siegel, M. M. (1999). Automated data massaging, interpretation, and E-mailing modules for high throughput open access mass spectrometry. Journal of the American Society for Mass Spectrometry, 10(11), 1174-1187, doi:10.1016/s1044-0305(99)00090-2. | NEG | 113 | 112.9856 | [M-H]- | Ion trap | Micromass Platform II | LC | ESI
Triphenyl phosphate | C18H15O4P | 326.0708 | InChI=1S/C18H15O4P/c19-23(20-16-10-4-1-5-11-16,21-17-12-6-2-7-13-17)22-18-14-8-3-9-15-18/h1-15H | XZZNDPSIHUTMOC-UHFFFAOYSA-N | CON00118 | Flame retardant in plastics | 8289 | Guo, X. H., Bruins, A. P., & Covey, T. R. (2006). Characterization of typical chemical background interferences in atmospheric pressure ionization liquid chromatography-mass spectrometry. Rapid Communications in Mass Spectrometry, 20(20), 3145-3150, doi:10.1002/rcm.2715. | POS | 327 | 327.0781 | [M+H]+ | Triple quadrupole | API365 | LC
Triphenylphosphine oxide | C18H15OP | 278.0861 | InChI=1S/C18H15OP/c19-20(16-10-4-1-5-11-16,17-12-6-2-7-13-17)18-14-8-3-9-15-18/h1-15H | FIQMHBFVRAXMOP-UHFFFAOYSA-N | CON00119 | Crystallizing agent | 13097 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 279 | 279.0933 | [M+H]+
Tripropylamine | C9H21N | 143.1674 | InChI=1S/C9H21N/c1-4-7-10(8-5-2)9-6-3/h4-9H2,1-3H3 | YFTHZRPMJXBUME-UHFFFAOYSA-N | CON00120 | Solvent | 7616 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 144 | 144.1747 | [M+H]+
TRIS | C4H11NO3 | 121.0739 | InChI=1S/C4H11NO3/c5-4(1-6,2-7)3-8/h6-8H,1-3,5H2 | LENZDBCJOHFCAS-UHFFFAOYSA-N | CON00121 | Buffer | 6503 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 122 | 122.0812 | [M+H]+
Triton X-100.1 | C16H26O2 | 250.1933 | InChI=1S/C16H26O2/c1-15(2,3)12-16(4,5)13-6-8-14(9-7-13)18-11-10-17/h6-9,17H,10-12H2,1-5H3 | JYCQQPHGFMYQCF-UHFFFAOYSA-N | CON00122 | Buffer | 5590 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 251.2006 | 251.2006 | [M+H]+
Triton X-100.1 | C16H26O2 | 250.1933 | InChI=1S/C16H26O2/c1-15(2,3)12-16(4,5)13-6-8-14(9-7-13)18-11-10-17/h6-9,17H,10-12H2,1-5H3 | JYCQQPHGFMYQCF-UHFFFAOYSA-N | CON00122 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 273.1825 | 273.1825 | [M+Na]+
Triton X-100.10 | C34H62O11 | 646.4292 | InChI=1S/C34H62O11/c1-33(2,3)30-34(4,5)31-6-8-32(9-7-31)45-29-28-44-27-26-43-25-24-42-23-22-41-21-20-40-19-18-39-17-16-38-15-14-37-13-12-36-11-10-35/h6-9,35H,10-30H2,1-5H3 | IVKNZCBNXPYYKL-UHFFFAOYSA-N | CON00123 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 647.4365 | 647.4365 | [M+H]+
Triton X-100.10 | C34H62O11 | 646.4292 | InChI=1S/C34H62O11/c1-33(2,3)30-34(4,5)31-6-8-32(9-7-31)45-29-28-44-27-26-43-25-24-42-23-22-41-21-20-40-19-18-39-17-16-38-15-14-37-13-12-36-11-10-35/h6-9,35H,10-30H2,1-5H3 | IVKNZCBNXPYYKL-UHFFFAOYSA-N | CON00123 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 669.4184 | 669.4184 | [M+Na]+
Triton X-100.11 | C36H66O12 | 690.4554 |  |  | CON00124 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 691.4627 | 691.4627 | [M+H]+
Triton X-100.11 | C36H66O12 | 690.4554 |  |  | CON00124 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 713.4446 | 713.4447 | [M+Na]+
Triton X-100.12 | C38H70O13 | 734.4816 |  |  | CON00125 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 735.4889 | 735.4889 | [M+H]+
Triton X-100.12 | C38H70O13 | 734.4816 |  |  | CON00125 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 757.4709 | 757.4709 | [M+Na]+
Triton X-100.13 | C40H74O14 | 778.5079 |  |  | CON00126 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 779.5151 | 779.5151 | [M+H]+
Triton X-100.13 | C40H74O14 | 778.5079 |  |  | CON00126 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 801.4971 | 801.4971 | [M+Na]+
Triton X-100.14 | C42H78O15 | 822.5341 |  |  | CON00127 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 823.5414 | 823.5414 | [M+H]+
Triton X-100.14 | C42H78O15 | 822.5341 |  |  | CON00127 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 845.5233 | 845.5233 | [M+Na]+
Triton X-100.15 | C44H82O16 | 866.5603 |  |  | CON00128 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 867.5676 | 867.5676 | [M+H]+
Triton X-100.15 | C44H82O16 | 866.5603 |  |  | CON00128 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 889.5495 | 889.5495 | [M+Na]+
Triton X-100.16 | C46H86O17 | 910.5865 |  |  | CON00129 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 911.5938 | 911.5938 | [M+H]+
Triton X-100.16 | C46H86O17 | 910.5865 |  |  | CON00129 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 933.5757 | 933.5757 | [M+Na]+
Triton X-100.17 | C48H90O18 | 954.6127 |  |  | CON00130 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 955.62 | 955.62 | [M+H]+
Triton X-100.17 | C48H90O18 | 954.6127 |  |  | CON00130 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 977.6019 | 977.6019 | [M+Na]+
Triton X-100.18 | C50H94O19 | 998.6389 |  |  | CON00131 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 999.6462 | 999.6462 | [M+H]+
Triton X-100.18 | C50H94O19 | 998.6389 |  |  | CON00131 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1021.6281 | 1021.6282 | [M+Na]+
Triton X-100.19 | C52H98O20 | 1042.6652 |  |  | CON00132 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1043.6724 | 1043.6724 | [M+H]+
Triton X-100.19 | C52H98O20 | 1042.6652 |  |  | CON00132 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1065.6543 | 1065.6544 | [M+Na]+
Triton X-100.20 | C54H102O21 | 1086.6914 |  |  | CON00133 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1087.6986 | 1087.6986 | [M+H]+
Triton X-100.20 | C54H102O21 | 1086.6914 |  |  | CON00133 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1109.6805 | 1109.6806 | [M+Na]+
Triton X-100.2 | C18H30O3 | 294.2195 | InChI=1S/C18H30O3/c1-17(2,3)14-18(4,5)15-6-8-16(9-7-15)21-13-12-20-11-10-19/h6-9,19H,10-14H2,1-5H3 | LBCZOTMMGHGTPH-UHFFFAOYSA-N | CON00134 | Buffer | 24775 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 295.2268 | 295.2268 | [M+H]+
Triton X-100.2 | C18H30O3 | 294.2195 |  |  | CON00134 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 317.2087 | 317.2087 | [M+Na]+
Triton X-100.3 | C20H34O4 | 338.2457 | InChI=1S/C20H34O4/c1-19(2,3)16-20(4,5)17-6-8-18(9-7-17)24-15-14-23-13-12-22-11-10-21/h6-9,21H,10-16H2,1-5H3 | OLOHNJXENGJYTB-UHFFFAOYSA-N | CON00135 | Buffer | 448876 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 339.253 | 339.253 | [M+H]+
Triton X-100.3 | C20H34O4 | 338.2457 |  |  | CON00135 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 361.2349 | 361.2349 | [M+Na]+
Triton X-100.4 | C22H38O5 | 382.2719 | InChI=1S/C22H38O5/c1-21(2,3)18-22(4,5)19-6-8-20(9-7-19)27-17-16-26-15-14-25-13-12-24-11-10-23/h6-9,23H,10-18H2,1-5H3 | UYDLBVPAAFVANX-UHFFFAOYSA-N | CON00136 | Buffer | 628327 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 383.2792 | 383.2792 | [M+H]+
Triton X-100.4 | C22H38O5 | 382.2719 |  |  | CON00136 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 405.2611 | 405.2611 | [M+Na]+
Triton X-100.5 | C24H42O6 | 426.2981 | InChI=1S/C24H42O6/c1-23(2,3)20-24(4,5)21-6-8-22(9-7-21)30-19-18-29-17-16-28-15-14-27-13-12-26-11-10-25/h6-9,25H,10-20H2,1-5H3 | WQBPSJIAFZCNBR-UHFFFAOYSA-N | CON00137 | Buffer | 560774 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 427.3054 | 427.3054 | [M+H]+
Triton X-100.5 | C24H42O6 | 426.2981 |  |  | CON00137 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 449.2874 | 449.2874 | [M+Na]+
Triton X-100.6 | C26H46O7 | 470.3244 | InChI=1S/C26H46O7/c1-25(2,3)22-26(4,5)23-6-8-24(9-7-23)33-21-20-32-19-18-31-17-16-30-15-14-29-13-12-28-11-10-27/h6-9,27H,10-22H2,1-5H3 | SQAUJCLQNBQZMT-UHFFFAOYSA-N | CON00138 | Buffer | 17245 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 471.3316 | 471.3316 | [M+H]+
Triton X-100.6 | C26H46O7 | 470.3244 |  |  | CON00138 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 493.3136 | 493.3136 | [M+Na]+
Triton X-100.7 | C28H50O8 | 514.3506 | InChI=1S/C28H50O8/c1-27(2,3)24-28(4,5)25-6-8-26(9-7-25)36-23-22-35-21-20-34-19-18-33-17-16-32-15-14-31-13-12-30-11-10-29/h6-9,29H,10-24H2,1-5H3 | HNLXNOZHXNSSPN-UHFFFAOYSA-N | CON00139 | Buffer | 75622 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 515.3578 | 515.3578 | [M+H]+
Triton X-100.7 | C28H50O8 | 514.3506 |  |  | CON00139 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 537.3398 | 537.3398 | [M+Na]+
Triton X-100.8 | C30H54O9 | 558.3768 |  |  | CON00140 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 559.3841 | 559.3841 | [M+H]+
Triton X-100.8 | C30H54O9 | 558.3768 |  |  | CON00140 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 581.366 | 581.366 | [M+Na]+
Triton X-100.9 | C32H58O10 | 602.403 | InChI=1S/C32H58O10/c1-31(2,3)28-32(4,5)29-6-8-30(9-7-29)42-27-26-41-25-24-40-23-22-39-21-20-38-19-18-37-17-16-36-15-14-35-13-12-34-11-10-33/h6-9,33H,10-28H2,1-5H3 | VJYAJQFKKLYARJ-UHFFFAOYSA-N | CON00141 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 603.4103 | 603.4103 | [M+H]+
Triton X-100.9 | C32H58O10 | 602.403 |  |  | CON00141 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 625.3922 | 625.3922 | [M+Na]+
Triton X-100R.1 | C16H32O2 | 256.2402 |  |  | CON00142 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 279.2294 | 279.2295 | [M+Na]+
Triton X-100R.10 | C34H68O11 | 652.4762 | InChI=1S/C34H68O11/c1-33(2,3)30-34(4,5)31-6-8-32(9-7-31)45-29-28-44-27-26-43-25-24-42-23-22-41-21-20-40-19-18-39-17-16-38-15-14-37-13-12-36-11-10-35/h31-32,35H,6-30H2,1-5H3 | AXMJGXMRGDCRND-UHFFFAOYSA-N | CON00143 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 675.4654 | 675.4654 | [M+Na]+
Triton X-100R.11 | C36H72O12 | 696.5024 |  |  | CON00144 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 719.4916 | 719.4916 | [M+Na]+
Triton X-100R.12 | C38H76O13 | 740.5286 |  |  | CON00145 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 763.5178 | 763.5178 | [M+Na]+
Triton X-100R.13 | C40H80O14 | 784.5548 |  |  | CON00146 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 807.544 | 807.544 | [M+Na]+
Triton X-100R.14 | C42H84O15 | 828.581 |  |  | CON00147 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 851.5702 | 851.5702 | [M+Na]+
Triton X-100R.15 | C44H88O16 | 872.6072 |  |  | CON00148 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 895.5965 | 895.5965 | [M+Na]+
Triton X-100R.16 | C46H92O17 | 916.6335 |  |  | CON00149 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 939.6227 | 939.6227 | [M+Na]+
Triton X-100R.17 | C48H96O18 | 960.6597 |  |  | CON00150 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 983.6489 | 983.6489 | [M+Na]+
Triton X-100R.18 | C50H100O19 | 1004.6859 |  |  | CON00151 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1027.6751 | 1027.6751 | [M+Na]+
Triton X-100R.19 | C52H104O20 | 1048.7121 |  |  | CON00152 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1071.7013 | 1071.7013 | [M+Na]+
Triton X-100R.20 | C54H108O21 | 1092.7383 |  |  | CON00153 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1115.7275 | 1115.7275 | [M+Na]+
Triton X-100R.2 | C18H36O3 | 300.2664 | InChI=1S/C18H36O3/c1-17(2,3)14-18(4,5)15-6-8-16(9-7-15)21-13-12-20-11-10-19/h15-16,19H,6-14H2,1-5H3 | PNAKTNFERWFEIX-UHFFFAOYSA-N | CON00154 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 323.2557 | 323.2557 | [M+Na]+
Triton X-100R.3 | C20H40O4 | 344.2927 |  |  | CON00155 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 367.2819 | 367.2819 | [M+Na]+
Triton X-100R.4 | C22H44O5 | 388.3189 |  |  | CON00156 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 411.3081 | 411.3081 | [M+Na]+
Triton X-100R.5 | C24H48O6 | 432.3451 |  |  | CON00157 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 455.3343 | 455.3343 | [M+Na]+
Triton X-100R.6 | C26H52O7 | 476.3713 |  |  | CON00158 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 499.3605 | 499.3605 | [M+Na]+
Triton X-100R.7 | C28H56O8 | 520.3975 |  |  | CON00159 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 543.3867 | 543.3867 | [M+Na]+
Triton X-100R.8 | C30H60O9 | 564.4237 | InChI=1S/C30H60O9/c1-29(2,3)26-30(4,5)27-6-8-28(9-7-27)39-25-24-38-23-22-37-21-20-36-19-18-35-17-16-34-15-14-33-13-12-32-11-10-31/h27-28,31H,6-26H2,1-5H3 | LVHZSSVOODJLAY-UHFFFAOYSA-N | CON00160 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 587.413 | 587.413 | [M+Na]+
Triton X-100R.9 | C32H64O10 | 608.4499 |  |  | CON00161 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 631.4392 | 631.4392 | [M+Na]+
Triton X-101.1 | C17H28O2 | 264.2089 |  |  | CON00162 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 265.2162 | 265.2162 | [M+H]+
Triton X-101.1 | C17H28O2 | 264.2089 |  |  | CON00162 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 287.1981 | 287.1982 | [M+Na]+
Triton X-101.10 | C35H64O11 | 660.4449 |  |  | CON00163 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 661.4521 | 661.4521 | [M+H]+
Triton X-101.10 | C35H64O11 | 660.4449 |  |  | CON00163 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 683.4341 | 683.4341 | [M+Na]+
Triton X-101.11 | C37H68O12 | 704.4711 |  |  | CON00164 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 705.4784 | 705.4784 | [M+H]+
Triton X-101.11 | C37H68O12 | 704.4711 |  |  | CON00164 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 727.4603 | 727.4603 | [M+Na]+
Triton X-101.12 | C39H72O13 | 748.4973 |  |  | CON00165 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 749.5046 | 749.5046 | [M+H]+
Triton X-101.12 | C39H72O13 | 748.4973 |  |  | CON00165 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 771.4865 | 771.4865 | [M+Na]+
Triton X-101.13 | C41H76O14 | 792.5235 |  |  | CON00166 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 793.5308 | 793.5308 | [M+H]+
Triton X-101.13 | C41H76O14 | 792.5235 |  |  | CON00166 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 815.5127 | 815.5127 | [M+Na]+
Triton X-101.14 | C43H80O15 | 836.5497 |  |  | CON00167 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 837.557 | 837.557 | [M+H]+
Triton X-101.14 | C43H80O15 | 836.5497 |  |  | CON00167 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 859.5389 | 859.5389 | [M+Na]+
Triton X-101.15 | C45H84O16 | 880.5759 |  |  | CON00168 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 881.5832 | 881.5832 | [M+H]+
Triton X-101.15 | C45H84O16 | 880.5759 |  |  | CON00168 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 903.5652 | 903.5652 | [M+Na]+
Triton X-101.16 | C47H88O17 | 924.6022 |  |  | CON00169 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 925.6094 | 925.6094 | [M+H]+
Triton X-101.16 | C47H88O17 | 924.6022 |  |  | CON00169 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 947.5914 | 947.5914 | [M+Na]+
Triton X-101.17 | C49H92O18 | 968.6284 |  |  | CON00170 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 969.6356 | 969.6356 | [M+H]+
Triton X-101.17 | C49H92O18 | 968.6284 |  |  | CON00170 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 991.6176 | 991.6176 | [M+Na]+
Triton X-101.18 | C51H96O19 | 1012.6546 |  |  | CON00171 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1013.6619 | 1013.6619 | [M+H]+
Triton X-101.18 | C51H96O19 | 1012.6546 |  |  | CON00171 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1035.6438 | 1035.6438 | [M+Na]+
Triton X-101.19 | C53H100O20 | 1056.6808 |  |  | CON00172 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1057.6881 | 1057.6881 | [M+H]+
Triton X-101.19 | C53H100O20 | 1056.6808 |  |  | CON00172 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1079.67 | 1079.67 | [M+Na]+
Triton X-101.20 | C55H104O21 | 1100.707 |  |  | CON00173 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1101.7143 | 1101.7143 | [M+H]+
Triton X-101.20 | C55H104O21 | 1100.707 |  |  | CON00173 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1123.6962 | 1123.6962 | [M+Na]+
Triton X-101.2 | C19H32O3 | 308.2351 |  |  | CON00174 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 309.2424 | 309.2424 | [M+H]+
Triton X-101.2 | C19H32O3 | 308.2351 |  |  | CON00174 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 331.2244 | 331.2244 | [M+Na]+
Triton X-101.3 | C21H36O4 | 352.2614 |  |  | CON00175 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 353.2686 | 353.2686 | [M+H]+
Triton X-101.3 | C21H36O4 | 352.2614 |  |  | CON00175 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 375.2506 | 375.2506 | [M+Na]+
Triton X-101.4 | C23H40O5 | 396.2876 |  |  | CON00176 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 397.2949 | 397.2949 | [M+H]+
Triton X-101.4 | C23H40O5 | 396.2876 |  |  | CON00176 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 419.2768 | 419.2768 | [M+Na]+
Triton X-101.5 | C25H44O6 | 440.3138 |  |  | CON00177 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 441.3211 | 441.3211 | [M+H]+
Triton X-101.5 | C25H44O6 | 440.3138 |  |  | CON00177 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 463.303 | 463.303 | [M+Na]+
Triton X-101.6 | C27H48O7 | 484.34 |  |  | CON00178 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 485.3473 | 485.3473 | [M+H]+
Triton X-101.6 | C27H48O7 | 484.34 |  |  | CON00178 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 507.3292 | 507.3292 | [M+Na]+
Triton X-101.7 | C29H52O8 | 528.3662 |  |  | CON00179 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 529.3735 | 529.3735 | [M+H]+
Triton X-101.7 | C29H52O8 | 528.3662 |  |  | CON00179 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 551.3554 | 551.3554 | [M+Na]+
Triton X-101.8 | C31H56O9 | 572.3924 |  |  | CON00180 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 573.3997 | 573.3997 | [M+H]+
Triton X-101.8 | C31H56O9 | 572.3924 |  |  | CON00180 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 595.3817 | 595.3817 | [M+Na]+
Triton X-101.9 | C33H60O10 | 616.4186 |  |  | CON00181 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 617.4259 | 617.4259 | [M+H]+
Triton X-101.9 | C33H60O10 | 616.4186 |  |  | CON00181 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 639.4079 | 639.4079 | [M+Na]+
Triton X-101R.1 | C17H34O2 | 270.2559 |  |  | CON00182 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 293.2451 | 293.2451 | [M+Na]+
Triton X-101R.10 | C35H70O11 | 666.4918 |  |  | CON00183 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 689.481 | 689.481 | [M+Na]+
Triton X-101R.11 | C37H74O12 | 710.518 |  |  | CON00184 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 733.5072 | 733.5073 | [M+Na]+
Triton X-101R.12 | C39H78O13 | 754.5442 |  |  | CON00185 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 777.5335 | 777.5335 | [M+Na]+
Triton X-101R.13 | C41H82O14 | 798.5705 |  |  | CON00186 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 821.5597 | 821.5597 | [M+Na]+
Triton X-101R.14 | C43H86O15 | 842.5967 |  |  | CON00187 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 865.5859 | 865.5859 | [M+Na]+
Triton X-101R.15 | C45H90O16 | 886.6229 |  |  | CON00188 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 909.6121 | 909.6121 | [M+Na]+
Triton X-101R.16 | C47H94O17 | 930.6491 |  |  | CON00189 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 953.6386 | 953.6383 | [M+Na]+
Triton X-101R.17 | C49H98O18 | 974.6753 |  |  | CON00190 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 997.6645 | 997.6645 | [M+Na]+
Triton X-101R.18 | C51H102O19 | 1018.7015 |  |  | CON00191 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1041.6904 | 1041.6908 | [M+Na]+
Triton X-101R.19 | C53H106O20 | 1062.7278 |  |  | CON00192 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1085.7163 | 1085.717 | [M+Na]+
Triton X-101R.20 | C55H110O21 | 1106.754 |  |  | CON00193 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1129.7422 | 1129.7432 | [M+Na]+
Triton X-101R.2 | C19H38O3 | 314.2821 |  |  | CON00194 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 337.2713 | 337.2713 | [M+Na]+
Triton X-101R.3 | C21H42O4 | 358.3083 |  |  | CON00195 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 381.2975 | 381.2975 | [M+Na]+
Triton X-101R.4 | C23H46O5 | 402.3345 |  |  | CON00196 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 425.3237 | 425.3237 | [M+Na]+
Triton X-101R.5 | C25H50O6 | 446.3607 |  |  | CON00197 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 469.35 | 469.35 | [M+Na]+
Triton X-101R.6 | C27H54O7 | 490.387 |  |  | CON00198 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 513.3762 | 513.3762 | [M+Na]+
Triton X-101R.7 | C29H58O8 | 534.4132 |  |  | CON00199 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 557.4024 | 557.4024 | [M+Na]+
Triton X-101R.8 | C31H62O9 | 578.4394 |  |  | CON00200 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 601.4286 | 601.4286 | [M+Na]+
Triton X-101R.9 | C33H66O10 | 622.4656 |  |  | CON00201 | Detergent |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 645.4548 | 645.4548 | [M+Na]+
Trityl cation | C19H15 | 243.1174 | InChI=1S/C19H15/c1-4-10-16(11-5-1)19(17-12-6-2-7-13-17)18-14-8-3-9-15-18/h1-15H/q+1 | OLFPYUPGPBITMH-UHFFFAOYSA-N | CON00202 | Catalyst | 84899 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 243 | 243.1168 | [M]+
Tween 20.10 | C38H74O16 | 786.4977 |  |  | CON00203 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 809.4869 | 809.4869 | [M+Na]+
Tween 20.11 | C40H78O17 | 830.5239 |  |  | CON00204 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 853.5131 | 853.5131 | [M+Na]+
Tween 20.12 | C42H82O18 | 874.5501 |  |  | CON00205 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 897.5393 | 897.5393 | [M+Na]+
Tween 20.13 | C44H86O19 | 918.5763 |  |  | CON00206 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 941.5656 | 941.5656 | [M+Na]+
Tween 20.14 | C46H90O20 | 962.6025 |  |  | CON00207 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 985.5918 | 985.5918 | [M+Na]+
Tween 20.15 | C48H94O21 | 1006.6288 |  |  | CON00208 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1029.618 | 1029.618 | [M+Na]+
Tween 20.16 | C50H98O22 | 1050.655 |  |  | CON00209 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1073.6442 | 1073.6442 | [M+Na]+
Tween 20.17 | C52H102O23 | 1094.6812 |  |  | CON00210 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1117.6704 | 1117.6704 | [M+Na]+
Tween 20.18 | C54H106O24 | 1138.7074 |  |  | CON00211 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1161.6966 | 1161.6966 | [M+Na]+
Tween 20.19 | C56H110O25 | 1182.7336 |  |  | CON00212 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1205.7228 | 1205.7228 | [M+Na]+
Tween 20.20 | C58H114O26 | 1226.7598 |  |  | CON00213 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1249.749 | 1249.7491 | [M+Na]+
Tween 20.21 | C60H118O27 | 1270.7861 |  |  | CON00214 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1293.7752 | 1293.7753 | [M+Na]+
Tween 20.22 | C62H122O28 | 1314.8123 |  |  | CON00215 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1337.8014 | 1337.8015 | [M+Na]+
Tween 20.23 | C64H126O29 | 1358.8385 |  |  | CON00216 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1381.8276 | 1381.8277 | [M+Na]+
Tween 20.24 | C66H130O30 | 1402.8647 |  |  | CON00217 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1425.8538 | 1425.8539 | [M+Na]+
Tween 20.25 | C68H134O31 | 1446.8909 |  |  | CON00218 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1469.88 | 1469.8801 | [M+Na]+
Tween 20.26 | C70H138O32 | 1490.9171 |  |  | CON00219 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1513.9062 | 1513.9064 | [M+Na]+
Tween 20.27 | C72H142O33 | 1534.9433 |  |  | CON00220 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1557.9324 | 1557.9326 | [M+Na]+
Tween 20.28 | C74H146O34 | 1578.9696 |  |  | CON00221 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1601.9586 | 1601.9588 | [M+Na]+
Tween 20.29 | C76H150O35 | 1622.9958 |  |  | CON00222 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1645.9848 | 1645.985 | [M+Na]+
Tween 20.30 | C78H154O36 | 1667.022 |  |  | CON00223 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1690.011 | 1690.0112 | [M+Na]+
Tween 20.31 | C80H158O37 | 1711.0482 |  |  | CON00224 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1734.0372 | 1734.0374 | [M+Na]+
Tween 20.32 | C82H162O38 | 1755.0744 |  |  | CON00225 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1778.0634 | 1778.0636 | [M+Na]+
Tween 40.10 | C42H82O16 | 842.5603 |  |  | CON00226 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 865.5495 | 865.5495 | [M+Na]+
Tween 40.11 | C44H86O17 | 886.5865 |  |  | CON00227 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 909.5757 | 909.5757 | [M+Na]+
Tween 40.12 | C46H90O18 | 930.6127 |  |  | CON00228 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 953.6019 | 953.6019 | [M+Na]+
Tween 40.13 | C48H94O19 | 974.6389 |  |  | CON00229 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 997.6282 | 997.6282 | [M+Na]+
Tween 40.14 | C50H98O20 | 1018.6652 |  |  | CON00230 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1041.6545 | 1041.6544 | [M+Na]+
Tween 40.15 | C52H102O21 | 1062.6914 |  |  | CON00231 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1085.6808 | 1085.6806 | [M+Na]+
Tween 40.16 | C54H106O22 | 1106.7176 |  |  | CON00232 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1129.7071 | 1129.7068 | [M+Na]+
Tween 40.17 | C56H110O23 | 1150.7438 |  |  | CON00233 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1173.7334 | 1173.733 | [M+Na]+
Tween 40.18 | C58H114O24 | 1194.77 |  |  | CON00234 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1217.7597 | 1217.7592 | [M+Na]+
Tween 40.19 | C60H118O25 | 1238.7962 |  |  | CON00235 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1261.786 | 1261.7854 | [M+Na]+
Tween 40.20 | C62H122O26 | 1282.8224 |  |  | CON00236 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1305.8123 | 1305.8117 | [M+Na]+
Tween 40.21 | C64H126O27 | 1326.8487 |  |  | CON00237 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1349.8386 | 1349.8379 | [M+Na]+
Tween 40.22 | C66H130O28 | 1370.8749 |  |  | CON00238 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1393.8649 | 1393.8641 | [M+Na]+
Tween 40.23 | C68H134O29 | 1414.9011 |  |  | CON00239 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1437.8912 | 1437.8903 | [M+Na]+
Tween 40.24 | C70H138O30 | 1458.9273 |  |  | CON00240 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1481.9175 | 1481.9165 | [M+Na]+
Tween 40.25 | C72H142O31 | 1502.9535 |  |  | CON00241 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1525.9438 | 1525.9427 | [M+Na]+
Tween 40.26 | C74H146O32 | 1546.9797 |  |  | CON00242 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1569.9701 | 1569.969 | [M+Na]+
Tween 40.27 | C76H150O33 | 1591.0059 |  |  | CON00243 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1613.9964 | 1613.9952 | [M+Na]+
Tween 40.28 | C78H154O34 | 1635.0322 |  |  | CON00244 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1658.0227 | 1658.0214 | [M+Na]+
Tween 40.29 | C80H158O35 | 1679.0584 |  |  | CON00245 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1702.049 | 1702.0476 | [M+Na]+
Tween 40.30 | C82H162O36 | 1723.0846 |  |  | CON00246 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1746.0753 | 1746.0738 | [M+Na]+
Tween 40.31 | C84H166O37 | 1767.1108 |  |  | CON00247 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1790.1016 | 1790.1 | [M+Na]+
Tween 40.32 | C86H170O38 | 1811.137 |  |  | CON00248 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1834.1279 | 1834.1262 | [M+Na]+
Tween 60.10 | C44H86O16 | 870.5916 |  |  | CON00249 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 893.5808 | 893.5808 | [M+Na]+
Tween 60.11 | C46H90O17 | 914.6178 |  |  | CON00250 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 937.607 | 937.607 | [M+Na]+
Tween 60.12 | C48H94O18 | 958.644 |  |  | CON00251 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 981.6332 | 981.6332 | [M+Na]+
Tween 60.13 | C50H98O19 | 1002.6702 |  |  | CON00252 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1025.6594 | 1025.6595 | [M+Na]+
Tween 60.14 | C52H102O20 | 1046.6964 |  |  | CON00253 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1069.6856 | 1069.6857 | [M+Na]+
Tween 60.15 | C54H106O21 | 1090.7227 |  |  | CON00254 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1113.7118 | 1113.7119 | [M+Na]+
Tween 60.16 | C56H110O22 | 1134.7489 |  |  | CON00255 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1157.738 | 1157.7381 | [M+Na]+
Tween 60.17 | C58H114O23 | 1178.7751 |  |  | CON00256 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1201.7642 | 1201.7643 | [M+Na]+
Tween 60.18 | C60H118O24 | 1222.8013 |  |  | CON00257 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1245.7904 | 1245.7905 | [M+Na]+
Tween 60.19 | C62H122O25 | 1266.8275 |  |  | CON00258 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1289.8166 | 1289.8167 | [M+Na]+
Tween 60.20 | C64H126O26 | 1310.8537 |  |  | CON00259 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1333.8428 | 1333.843 | [M+Na]+
Tween 60.21 | C66H130O27 | 1354.88 |  |  | CON00260 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1377.869 | 1377.8692 | [M+Na]+
Tween 60.22 | C68H134O28 | 1398.9062 |  |  | CON00261 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1421.8952 | 1421.8954 | [M+Na]+
Tween 60.23 | C70H138O29 | 1442.9324 |  |  | CON00262 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1465.9214 | 1465.9216 | [M+Na]+
Tween 60.24 | C72H142O30 | 1486.9586 |  |  | CON00263 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1509.9476 | 1509.9478 | [M+Na]+
Tween 60.25 | C74H146O31 | 1530.9848 |  |  | CON00264 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1553.9738 | 1553.974 | [M+Na]+
Tween 60.26 | C76H150O32 | 1575.011 |  |  | CON00265 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1598 | 1598.0003 | [M+Na]+
Tween 60.27 | C78H154O33 | 1619.0372 |  |  | CON00266 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1642.0262 | 1642.0265 | [M+Na]+
Tween 60.28 | C80H158O34 | 1663.0635 |  |  | CON00267 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1686.0524 | 1686.0527 | [M+Na]+
Tween 60.29 | C82H162O35 | 1707.0897 |  |  | CON00268 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1730.0786 | 1730.0789 | [M+Na]+
Tween 60.30 | C84H166O36 | 1751.1159 |  |  | CON00269 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1774.1048 | 1774.1051 | [M+Na]+
Tween 60.31 | C86H170O37 | 1795.1421 |  |  | CON00270 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1818.131 | 1818.1313 | [M+Na]+
Tween 60.32 | C88H174O38 | 1839.1683 |  |  | CON00271 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1862.1572 | 1862.1575 | [M+Na]+
Tween 80.10 | C44H84O16 | 868.5759 |  |  | CON00272 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 891.5652 | 891.5652 | [M+Na]+
Tween 80.11 | C46H88O17 | 912.6022 |  |  | CON00273 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 935.5914 | 935.5914 | [M+Na]+
Tween 80.12 | C48H92O18 | 956.6284 |  |  | CON00274 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 979.6176 | 979.6176 | [M+Na]+
Tween 80.13 | C50H96O19 | 1000.6546 |  |  | CON00275 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1023.6438 | 1023.6438 | [M+Na]+
Tween 80.14 | C52H100O20 | 1044.6808 |  |  | CON00276 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1067.67 | 1067.67 | [M+Na]+
Tween 80.15 | C54H104O21 | 1088.707 |  |  | CON00277 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1111.6962 | 1111.6962 | [M+Na]+
Tween 80.16 | C56H108O22 | 1132.7332 |  |  | CON00278 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1155.7224 | 1155.7225 | [M+Na]+
Tween 80.17 | C58H112O23 | 1176.7594 |  |  | CON00279 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1199.7486 | 1199.7487 | [M+Na]+
Tween 80.18 | C60H116O24 | 1220.7857 |  |  | CON00280 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1243.7748 | 1243.7749 | [M+Na]+
Tween 80.19 | C62H120O25 | 1264.8119 |  |  | CON00281 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1287.801 | 1287.8011 | [M+Na]+
Tween 80.20 | C64H124O26 | 1308.8381 |  |  | CON00282 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1331.8272 | 1331.8273 | [M+Na]+
Tween 80.21 | C66H128O27 | 1352.8643 |  |  | CON00283 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1375.8534 | 1375.8535 | [M+Na]+
Tween 80.22 | C68H132O28 | 1396.8905 |  |  | CON00284 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1419.8796 | 1419.8797 | [M+Na]+
Tween 80.23 | C70H136O29 | 1440.9167 |  |  | CON00285 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1463.9058 | 1463.906 | [M+Na]+
Tween 80.24 | C72H140O30 | 1484.943 |  |  | CON00286 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1507.932 | 1507.9322 | [M+Na]+
Tween 80.25 | C74H144O31 | 1528.9692 |  |  | CON00287 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1551.9582 | 1551.9584 | [M+Na]+
Tween 80.26 | C76H148O32 | 1572.9954 |  |  | CON00288 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1595.9844 | 1595.9846 | [M+Na]+
Tween 80.27 | C78H152O33 | 1617.0216 |  |  | CON00289 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1640.0106 | 1640.0108 | [M+Na]+
Tween 80.28 | C80H156O34 | 1661.0478 |  |  | CON00290 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1684.0368 | 1684.037 | [M+Na]+
Tween 80.29 | C82H160O35 | 1705.074 |  |  | CON00291 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1728.063 | 1728.0632 | [M+Na]+
Tween 80.30 | C84H164O36 | 1749.1002 |  |  | CON00292 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1772.0892 | 1772.0895 | [M+Na]+
Tween 80.31 | C86H168O37 | 1793.1265 |  |  | CON00293 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1816.1154 | 1816.1157 | [M+Na]+
Tween 80.32 | C88H172O38 | 1837.1527 |  |  | CON00294 | Buffer |  | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1860.1416 | 1860.1419 | [M+Na]+
1,3-Dicyclohexylurea | C13H24N2O | 224.1889 | InChI=1S/C13H24N2O/c16-13(14-11-7-3-1-4-8-11)15-12-9-5-2-6-10-12/h11-12H,1-10H2,(H2,14,15,16) | ADFXKUOMJKEIND-UHFFFAOYSA-N | CON00295 |  | 4277 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 225 | 225.1961 | [M+H]+
1,3-Dicyclohexylurea | C13H24N2O | 224.1889 | InChI=1S/C13H24N2O/c16-13(14-11-7-3-1-4-8-11)15-12-9-5-2-6-10-12/h11-12H,1-10H2,(H2,14,15,16) | ADFXKUOMJKEIND-UHFFFAOYSA-N | CON00295 |  | 4277 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 449 | 449.385 | [M2+H]+
1,8-Diazabicycloundec-7-ene | C9H16N2 | 152.1313 | InChI=1S/C9H16N2/c1-2-5-9-10-6-4-8-11(9)7-3-1/h1-8H2 | GQHTUMJGOHRCHB-UHFFFAOYSA-N | CON00296 | Catalyst | 81184 | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 153 | 153.1386 | [M+H]+
3-((3-cholamidopropyl)dimethylammonium)-1-propanesulfonate (CHAPS) | C32H58N2O7S | 614.3965 | InChI=1S/C32H58N2O7S/c1-21(8-11-29(38)33-14-6-15-34(4,5)16-7-17-42(39,40)41)24-9-10-25-30-26(20-28(37)32(24,25)3)31(2)13-12-23(35)18-22(31)19-27(30)36/h21-28,30,35-37H,6-20H2,1-5H3,(H-,33,38,39,40,41)/t21-,22+,23-,24-,25+,26+,27-,28+,30+,31+,32-/m1/s1 | UMCMPZBLKLEWAF-BCTGSCMUSA-N | CON00297 | Solvent | 107670 | Thermo Corporation, Notes on Troubleshooting LC/MS Contamination | POS | 615.7 | 615.4038 | [M+H]+
3-((3-cholamidopropyl)dimethylammonium)-1-propanesulfonate (CHAPS) | C32H58N2O7S | 614.3965 | InChI=1S/C32H58N2O7S/c1-21(8-11-29(38)33-14-6-15-34(4,5)16-7-17-42(39,40)41)24-9-10-25-30-26(20-28(37)32(24,25)3)31(2)13-12-23(35)18-22(31)19-27(30)36/h21-28,30,35-37H,6-20H2,1-5H3,(H-,33,38,39,40,41)/t21-,22+,23-,24-,25+,26+,27-,28+,30+,31+,32-/m1/s1 | UMCMPZBLKLEWAF-BCTGSCMUSA-N | CON00297 | Solvent | 107670 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 1229.8002 | 1229.8002 | [M2+H]+
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 190.0499 | 190.0499 | [M+H]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 212.0318 | 212.0318 | [M+Na]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 228.0058 | 228.0058 | [M+K]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 379.0925 | 379.0925 | [M2+H]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 441.0148 | 441.0142 | [M2+63Cu(I)]+
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., Suj, J., Young, A. B., & Whittal, R. M. (2008). Interferences and contaminants encountered in modern mass spectrometry. Analytica Chimica Acta, 627(1), 71-81, doi:10.1016/j.aca.2008.04.043. | POS | 443.013 | 443.0124 | [M2+65Cu(I)]+
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 568.1351 | 568.1351 | [M3+H]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. | POS | 644.0473 | 644.0468 | [M3-H+K2]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. | POS | 666.0293 | 666.0288 | [M3-H2+K2+Na]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. | POS | 680.9954 | 680.9949 | [M3-H3+K3]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. | POS | 833.0899 | 833.0894 | [M4-H+K2]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 839.0974 | 839.0974 | [M4-H2+K+Na2]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 845.1054 | 845.1054 | [M4-H3+Na4]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. | POS | 855.0719 | 855.0714 | [M4-H2+K2+Na]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 867.0874 | 867.0874 | [M4-H4+Na5]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 871.0453 | 871.0453 | [M4-H2+K3]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. | POS | 871.0458 | 871.0453 | [M4-H2+K3]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Harris, W. A., Janecki, D. J., & Reilly, J. P. (2002). Use of matrix clusters and trypsin autolysis fragments as mass calibrants in matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid Commun Mass Spectrom, 16(18), 1714-1722, doi:10.1002/rcm.775. | POS | 907.9938 | 907.9934 | [M4-H4+K4]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 994.1555 | 994.1555 | [M5-H2O-H2+Na3]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 172.0393 | 172.0393 | [M-H2O+H]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 606.0915 | 606.0909 | [M3+K]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 861.0794 | 861.0794 | [M4-H3+Na3+K]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 1036.0828 | 1036.0828 | [M5-(H2O)2-H4+Na4+K]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 1044.114 | 1044.114 | [M5-H2+Na+K2]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 1054.0933 | 1054.0933 | [M5-H2O-H4+Na4+K]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 1060.0879 | 1060.0879 | [M5-H2+K3]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 1066.0959 | 1066.0965 | [M5-H3+K2+Na2]+ | ToF |  |  | MALDI
4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | CON00298 | Matrix compound | 5328791 | Keller, B. O., & Li, L. (2000). Discerning matrix-cluster peaks in matrix-assisted laser desorption/ionization time-of-flight mass spectra of dilute peptide mixtures. Journal of the American Society for Mass Spectrometry, 11(1), 88-93, doi:10.1016/s1044-0305(99)00126-9. | POS | 1225.1254 | 1225.1254 | [M6-(H2O)2-H4+Na4+K]+ | ToF |  |  | MALDI
unknown | unknown | 0 |  |  | CON00299 | Plasticiser |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 425 | 0 | [M+Na]+
unknown | unknown | 0 |  |  | CON00300 |  |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 522 | 0
unknown | unknown | 0 |  |  | CON00301 |  |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 587 | 0 | [M+H]+
unknown | unknown | 0 |  |  | CON00302 |  |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 609 | 0 | [M+Na]+
unknown | unknown | 0 |  |  | CON00303 |  |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 625 | 0 | [M+K]+
unknown | unknown | 0 |  |  | CON00304 |  |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 610 | 0
unknown | unknown | 0 |  |  | CON00305 |  |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 638 | 0
unknown | unknown | 0 |  |  | CON00306 |  |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 675 | 0
unknown | unknown | 0 |  |  | CON00307 |  |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 684 | 0
unknown | unknown | 0 |  |  | CON00308 | Filter |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 696 | 0 | [M+NH4]+
unknown | unknown | 0 |  |  | CON00309 | Filter |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 701 | 0 | [M+Na]+
unknown | unknown | 0 |  |  | CON00310 |  |  | Waters - ESI+ Common Background Ions. http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf. | POS | 786 | 0 | [M+Ti]+
unknown | unknown | 0 |  |  | CON00311 |  |  | NewObjective Common Background Ions for Electrospray (Positive Ion). http://www.newobjective.com/downloads/technotes/PV-3.pdf. | POS | 217 | 0
unknown | unknown | 0 |  |  | CON00312 | Plasticiser |  | NewObjective Common Background Ions for Electrospray (Positive Ion). http://www.newobjective.com/downloads/technotes/PV-3.pdf. | POS | 282 | 0
unknown | unknown | 0 |  |  | CON00313 |  |  | NewObjective Common Background Ions for Electrospray (Positive Ion). http://www.newobjective.com/downloads/technotes/PV-3.pdf. | POS | 599 | 0
unknown | unknown | 0 |  |  | CON00314 |  |  | NewObjective Common Background Ions for Electrospray (Positive Ion). http://www.newobjective.com/downloads/technotes/PV-3.pdf. | POS | 879 | 0
unknown | unknown | 0 |  |  | CON00315 |  |  | NewObjective Common Background Ions for Electrospray (Positive Ion). http://www.newobjective.com/downloads/technotes/PV-3.pdf. | POS | 933 | 0
id | name | formula | exact_mass | std_inchi | std_inchi_key | pubchem_cid | type_of_contaminant
CON00001 | Acetic Acid | C2H4O2 | 60.0211 | InChI=1S/C2H4O2/c1-2(3)4/h1H3,(H,3,4) | QTBSBXVTEAMEQO-UHFFFAOYSA-N | 176 | Solvent
CON00002 | Acetonitrile (fragment) | CN | 26.0031
CON00003 | Acetonitrile | CH3CN | 41.0265 | InChI=1S/C2H3N/c1-2-3/H1h3 | WEVYAHXRMPXWCK-UHFFFAOYSA-N | 6342 | Solvent
CON00004 | Acetonitrile.1.Acetic acid.1 | C4H7NO2 | 101.0477 |  |  |  | Solvent
CON00005 | Acetonitrile.1.Methanol.1 | C3H7NO | 73.0528 |  |  |  | Solvent
CON00006 | Acetonitrile.2.Methanol.2 | C6H14N2O2 | 146.1055 |  |  |  | Solvent
CON00007 | Behentrimonium chloride | C25H54NCl | 403.3945 | InChI=1S/C25H54N.ClH/c1-5-6-7-8-9-10-11-12-13-14-15-16-17-18-19-20-21-22-23-24-25-26(2,3)4;/h5-25H2,1-4H3;1H/q+1;/p-1 | YSJGOMATDFSEED-UHFFFAOYSA-M |  | Personal care products
CON00008 | Bis(2-ethylhexyl) Adipate | C22H42O4 | 370.3083 | InChI=1S/C22H42O4/c1-5-9-13-19(7-3)17-25-21(23)15-11-12-16-22(24)26-18-20(8-4)14-10-6-2/h19-20H,5-18H2,1-4H3 | SAOKZLXYCUGLFA-UHFFFAOYSA-N | 7641 | Plasticiser
CON00009 | Bis(2-ethylhexyl) Phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-5-9-13-19(7-3)17-27-23(25)21-15-11-12-16-22(21)24(26)28-18-20(8-4)14-10-6-2/h11-12,15-16,19-20H,5-10,13-14,17-18H2,1-4H3 | BJQHLKABXJIVAM-UHFFFAOYSA-N | 8343 | Plasticiser
CON00010 | Butyl Carbitol | C8H18O3 | 162.1256 | InChI=1S/C8H18O3/c1-2-3-5-10-7-8-11-6-4-9/h9H,2-8H2,1H3 | OAYXUHPQHDHDDZ-UHFFFAOYSA-N | 8177 | Scintillation cocktail
CON00011 | Butylated Hydroxyanisole | C11H16O2 | 180.115 | InChI=1S/C11H16O2/c1-11(2,3)9-7-8(13-4)5-6-10(9)12/h5-7,12H,1-4H3 | MRBKEAMVRSLQPH-UHFFFAOYSA-N | 11954184 | Antioxidant
CON00012 | Butylated Hydroxytoluene | C15H24O | 220.1827 | InChI=1S/C15H24O/c1-10-8-11(14(2,3)4)13(16)12(9-10)15(5,6)7/h8-9,16H,1-7H3 | NLZUEZXRPGMBCV-UHFFFAOYSA-N | 31404 | Antioxidant
CON00013 | Copper catalyst iodide adduct (63Cu(I)) | 63Cu(I) | 189.8341
CON00014 | Copper catalyst iodide adduct (65Cu(I)) | 65Cu(I) | 191.8323
CON00015 | Cs-133 | Cs | 132.9055
CON00016 | d6-Dimethylsulfoxide | C2D6OS | 84.0516 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3/i1D3,2D3 | IAZDPXIOMUYVGZ-WFGJKAKNSA-N | 75151 | Solvent
CON00017 | DDTDP | C30H58O4S | 514.4056
CON00018 | Decamethylcyclopentasiloxane | C10H30Si5O5 | 370.094 | InChI=1S/C10H30O5Si5/c1-16(2)11-17(3,4)13-19(7,8)15-20(9,10)14-18(5,6)12-16/h1-10H3 | XMSXQFUHVRWGNA-UHFFFAOYSA-N | 10193 | Airborne Contaminant
CON00019 | Dibutyl phthalate | C16H22O4 | 278.1518 | InChI=1S/C16H22O4/c1-3-5-7-11-9-10-13(15(17)18)14(16(19)20)12(11)8-6-4-2/h9-10H,3-8H2,1-2H3,(H,17,18)(H,19,20)/p-2 | NUONTTMMRCYJMI-UHFFFAOYSA-L | 3026 | Plasticiser
CON00020 | Dibutyl Sebacate | C18H34O4 | 314.2457 | InChI=1S/C18H34O4/c1-3-5-15-21-17(19)13-11-9-7-8-10-12-14-18(20)22-16-6-4-2/h3-16H2,1-2H3 | PYGXAGIECVVIOZ-UHFFFAOYSA-N | 7986 | Plasticiser
CON00021 | Didodecyl 3,3'-thiodipropanoate sulfone product | C30H58O6S | 546.3954 |  |  |  | Antioxidant
CON00022 | Didodecyl 3,3'-thiodipropanoate sulfoxide product | C30H58O5S | 530.4005 |  |  |  | Antioxidant
CON00023 | Diethyl carbitol | C8H18O3 | 162.1256 | InChI=1S/C8H18O3/c1-3-9-5-7-11-8-6-10-4-2/h3-8H2,1-2H3 | RRQYJINTUHWNHW-UHFFFAOYSA-N | 8179 | Scintillation cocktail
CON00024 | Diisooctyl phthalate | C24H38O4 | 390.277 | InChI=1S/C24H38O4/c1-19(2)13-7-5-11-17-27-23(25)21-15-9-10-16-22(21)24(26)28-18-12-6-8-14-20(3)4/h9-10,15-16,19-20H,5-8,11-14,17-18H2,1-4H3 | IJFPVINAQGWBRJ-UHFFFAOYSA-N | 33934 | Plasticiser
CON00025 | Diisopropylethylamine | C8H19N | 129.1517 | InChI=1S/C8H19N/c1-6-9(7(2)3)8(4)5/h7-8H,6H2,1-5H3 |  | 81531 | Solvent
CON00026 | Dimethyl Formamide | C3H7NO | 73.0528 | InChI=1S/C3H7NO/c1-4(2)3-5/h3H,1-2H3 | ZMXDDKWLCZADIW-UHFFFAOYSA-N | 6228 | Solvent
CON00027 | Dimethyl Phthalate | C10H10O4 | 194.0579 | InChI=1S/C10H10O4/c1-13-9(11)7-5-3-4-6-8(7)10(12)14-2/h3-6H,1-2H3 |  | 8554 | Plasticiser
CON00028 | Dimethylaminopyridine | C7H10N2 | 122.0844 | InChI=1S/C7H10N2/c1-9(2)7-5-3-4-6-8-7/h3-6H,1-2H3 | PSHKMPUSSFXUIA-UHFFFAOYSA-N | 21885 | Solvent
CON00029 | Dimethylsulfoxide (DMSO) | C2H6OS | 78.0139 | InChI=1S/C2H6OS/c1-4(2)3/h1-2H3 | IAZDPXIOMUYVGZ-UHFFFAOYSA-N | 679 | Solvent
CON00030 | Dipalmityldimethylammonium chloride | C34H72NCl | 529.5353 |  |  |  | Personal care products
CON00031 | Diphenyl ketone | C13H10O | 182.0732 | InChI=1S/C13H10O/c14-13(11-7-3-1-4-8-11)12-9-5-2-6-10-12/h1-10H | RWCCWEUUXYIKHB-UHFFFAOYSA-N | 3102
CON00032 | Distearyldimethylammonium chloride | C38H80NCl | 585.5979 | InChI=1S/C38H80N.ClH/c1-5-7-9-11-13-15-17-19-21-23-25-27-29-31-33-35-37-39(3,4)38-36-34-32-30-28-26-24-22-20-18-16-14-12-10-8-6-2;/h5-38H2,1-4H3;1H/q+1;/p-1 | REZZEXDLIUJMMS-UHFFFAOYSA-M | 7879 | Personal care products
CON00033 | Dodecamethylcyclohexasiloxane | C12H36Si6O6 | 444.1128 | InChI=1S/C12H36O6Si6/c1-19(2)13-20(3,4)15-22(7,8)17-24(11,12)18-23(9,10)16-21(5,6)14-19/h1-12H3 | IUMSDRXLFWAGNT-UHFFFAOYSA-N | 109111 | Airborne Contaminant
CON00034 | Eicosamethylcyclodecasiloxane | C20H60Si10O10 | 740.1879 | InChI=1S/C20H60O10Si10/c1-31(2)21-32(3,4)23-34(7,8)25-36(11,12)27-38(15,16)29-40(19,20)30-39(17,18)28-37(13,14)26-35(9,10)24-33(5,6)22-31/h1-20H3 | CDNNKGWZSNSADW-UHFFFAOYSA-N | 519601 | Airborne Contaminant
CON00035 | Erucamide | C22H43NO | 337.3345 | InChI=1S/C22H43NO/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18-19-20-21-22(23)24/h9-10H,2-8,11-21H2,1H3,(H2,23,24)/b10-9- | UAUDZVJPLUQNMU-KTKRTIGZSA-N | 5365371 | Anti-static agent
CON00036 | Formic acid | H2CO2 | 46.0055 | InChI=1S/CH2O2/c2-1-3/h1H,(H,2,3) | BDAGIHXWWSANSR-UHFFFAOYSA-N | 284
CON00036 | Formic acid | H2CO2 | 46.0055 | InChI=1S/CH2O2/c2-1-3/h1H,(H,2,3) | BDAGIHXWWSANSR-UHFFFAOYSA-N
CON00037 | Hexadecamethylcyclooctasiloxane | C16H48Si8O8 | 592.1503 | InChI=1S/C16H48O8Si8/c1-25(2)17-26(3,4)19-28(7,8)21-30(11,12)23-32(15,16)24-31(13,14)22-29(9,10)20-27(5,6)18-25/h1-16H3 | XKJMJYZFAWYREL-UHFFFAOYSA-N | 11170 | Airborne Contaminant
CON00038 | Iodine anion (Iodide) | I | 126.9045
CON00039 | Irganox | C35H62O3 | 530.4699 | InChI=1S/C35H62O3/c1-8-9-10-11-12-13-14-15-16-17-18-19-20-21-22-23-26-38-32(36)25-24-29-27-30(34(2,3)4)33(37)31(28-29)35(5,6)7/h27-28,37H,8-26H2,1-7H3 | SSDSCDGVMJFTEQ-UHFFFAOYSA-N | 16386 | Antioxidant
CON00040 | Methanesulfonate | CH4O3S | 95.9881 | InChI=1S/CH4O3S/c1-5(2,3)4/h1H3,(H,2,3,4)/p-1 | AFVFQIVMOAPDHO-UHFFFAOYSA-M
CON00041 | Methanol | CH3OH | 32.0262 | InChI=1S/CH4O/c1-2/h2H,1H3 | OKKJLVBELUTLKV-UHFFFAOYSA-N | 887 | Solvent
CON00042 | Methanol.2.Water.2 | C2H12O4 | 100.0736 |  |  |  | Solvent
CON00043 | Methanol.3.Water.2 | C3H16O5 | 132.0998 |  |  |  | Solvent
CON00044 | Methanol.4.Water.3 | C4H22O7 | 182.1366 |  |  |  | Solvent
CON00045 | Monomethoxytrityl | C20H17O | 273.1279 | InChI=1S/C20H17O/c1-21-19-14-12-18(13-15-19)20(16-8-4-2-5-9-16)17-10-6-3-7-11-17/h2-15H,1H3 | NTSFISBKIKSZLT-UHFFFAOYSA-N
CON00046 | N,N-Bis(2-hydroxyethyl)dodecanamide | C16H33NO3 | 287.246 | InChI=1S/C16H33NO3/c1-2-3-4-5-6-7-8-9-10-11-16(20)17(12-14-18)13-15-19/h18-19H,2-15H2,1H3 | AOMUHOFOVNGZAN-UHFFFAOYSA-N |  | Anti-static agent
CON00047 | N-Butylbenzenesulfonamide | C10H15NO2S | 213.0824 | InChI=1S/C10H15NO2S/c1-2-3-9-11-14(12,13)10-7-5-4-6-8-10/h4-8,11H,2-3,9H2,1H3 | IPRJXAGUEGOFGG-UHFFFAOYSA-N | 19241 | Plasticiser
CON00048 | N-Methyl 2-pyrrolidone | C5H9NO | 99.0684 | InChI=1S/C5H9NO/c1-6-4-2-3-5(6)7/h2-4H2,1H3 | SECXISVLQFMRJM-UHFFFAOYSA-N | 13387 | Solvent
CON00051 | Nylon 66 polymer | C24H44N4O4 | 452.3363 |  |  |  | Filter
CON00051 | Nylon 66 polymer | C36H66N6O6 | 678.5044 |  |  |  | Filter
CON00051 | Nylon 66 polymer | C48H88N8O8 | 904.6725 |  |  |  | Filter
CON00052 | Octadecamethylcyclononasiloxane | C18H54Si9O9 | 666.1691 | InChI=1S/C18H54O9Si9/c1-28(2)19-29(3,4)21-31(7,8)23-33(11,12)25-35(15,16)27-36(17,18)26-34(13,14)24-32(9,10)22-30(5,6)20-28/h1-18H3 | ISXOGOLHEGHGQF-UHFFFAOYSA-N | 11172 | Airborne Contaminant
CON00053 | Oleamide | C18H35NO | 281.2719 | InChI=1S/C18H35NO/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18(19)20/h9-10H,2-8,11-17H2,1H3,(H2,19,20)/b10-9- | FATBGEAMYMYZAF-KTKRTIGZSA-N | 5283387 | Slip agent
CON00054 | Oleic acid | C18H34O2 | 282.2559 | InChI=1S/C18H34O2/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18(19)20/h9-10H,2-8,11-17H2,1H3,(H,19,20)/b10-9- | ZQPPMHVWECSIRJ-KTKRTIGZSA-N | 445639
CON00055 | Palmitic acid | C16H32O2 | 256.2402 | InChI=1S/C16H32O2/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16(17)18/h2-15H2,1H3,(H,17,18) | IPCSVZSSVZVIGE-UHFFFAOYSA-N | 985
CON00056 | Palmitamidopropyl-trimonium chloride | C22H47N2OCl | 390.3377 |  |  |  | Personal care products
CON00057 | PEG.1 | C2H6O2 | 62.0368 | InChI=1S/C2H6O2/c3-1-2-4/h3-4H,1-2H2 | LYCAIKOWRPUZTN-UHFFFAOYSA-N | 174 | Polyether
CON00058 | PEG.10 | C20H42O11 | 458.2727 | InChI=1S/C20H42O11/c21-1-3-23-5-7-25-9-11-27-13-15-29-17-19-31-20-18-30-16-14-28-12-10-26-8-6-24-4-2-22/h21-22H,1-20H2 | DTPCFIHYWYONMD-UHFFFAOYSA-N | 79689 | Polyether
CON00059 | PEG.11 | C22H46O12 | 502.2989 | InChI=1S/C22H46O12/c23-1-3-25-5-7-27-9-11-29-13-15-31-17-19-33-21-22-34-20-18-32-16-14-30-12-10-28-8-6-26-4-2-24/h23-24H,1-22H2 | PSVXZQVXSXSQRO-UHFFFAOYSA-N | 81260 | Polyether
CON00060 | PEG.12 | C24H50O13 | 546.3251 | InChI=1S/C24H50O13/c25-1-3-27-5-7-29-9-11-31-13-15-33-17-19-35-21-23-37-24-22-36-20-18-34-16-14-32-12-10-30-8-6-28-4-2-26/h25-26H,1-24H2 | WRZXKWFJEFFURH-UHFFFAOYSA-N | 81248 | Polyether
CON00061 | PEG.13 | C26H54O14 | 590.3514 | InChI=1S/C26H54O14/c27-1-3-29-5-7-31-9-11-33-13-15-35-17-19-37-21-23-39-25-26-40-24-22-38-20-18-36-16-14-34-12-10-32-8-6-30-4-2-28/h27-28H,1-26H2 | AKWFJQNBHYVIPY-UHFFFAOYSA-N | 87168 | Polyether
CON00062 | PEG.14 | C28H58O15 | 634.3776 | InChI=1S/C28H58O15/c29-1-3-31-5-7-33-9-11-35-13-15-37-17-19-39-21-23-41-25-27-43-28-26-42-24-22-40-20-18-38-16-14-36-12-10-34-8-6-32-4-2-30/h29-30H,1-28H2 | ILLKMACMBHTSHP-UHFFFAOYSA-N |  | Polyether
CON00063 | PEG.15 | C30H62O16 | 678.4038 | InChI=1S/C30H62O16/c31-1-3-33-5-7-35-9-11-37-13-15-39-17-19-41-21-23-43-25-27-45-29-30-46-28-26-44-24-22-42-20-18-40-16-14-38-12-10-36-8-6-34-4-2-32/h31-32H,1-30H2 | OWTQQPNDSWCHOV-UHFFFAOYSA-N | 120062 | Polyether
CON00064 | PEG.16 | C32H66O17 | 722.43 | InChI=1S/C32H66O17/c33-1-3-35-5-7-37-9-11-39-13-15-41-17-19-43-21-23-45-25-27-47-29-31-49-32-30-48-28-26-46-24-22-44-20-18-42-16-14-40-12-10-38-8-6-36-4-2-34/h33-34H,1-32H2 | DHORSBRLGKJPFC-UHFFFAOYSA-N | 81262 | Polyether
CON00065 | PEG.17 | C34H70O18 | 766.4562 |  |  |  | Polyether
CON00066 | PEG.18 | C36H74O19 | 810.4824 |  |  |  | Polyether
CON00067 | PEG.19 | C38H78O20 | 854.5086 |  |  |  | Polyether
CON00068 | PEG.2 | C4H10O3 | 106.063 | InChI=1S/C4H10O3/c5-1-3-7-4-2-6/h5-6H,1-4H2 | MTHSVFCYNBDYFN-UHFFFAOYSA-N | 8117 | Polyether
CON00069 | PEG.20 | C40H82O21 | 898.5349 | InChI=1S/C40H82O21/c41-1-3-43-5-7-45-9-11-47-13-15-49-17-19-51-21-23-53-25-27-55-29-31-57-33-35-59-37-39-61-40-38-60-36-34-58-32-30-56-28-26-54-24-22-52-20-18-50-16-14-48-12-10-46-8-6-44-4-2-42/h41-42H,1-40H2 | ZHMUMDLGQBRJIW-UHFFFAOYSA-N |  | Polyether
CON00070 | PEG.3 | C6H14O4 | 150.0892 | InChI=1S/C6H14O4/c7-1-3-9-5-6-10-4-2-8/h7-8H,1-6H2 | ZIBGPFATKBEMQZ-UHFFFAOYSA-N | 8172 | Polyether
CON00071 | PEG.4 | C8H18O5 | 194.1154 | InChI=1S/C8H18O5/c9-1-3-11-5-7-13-8-6-12-4-2-10/h9-10H,1-8H2 | UWHCKJMYHZGTIT-UHFFFAOYSA-N | 8200 | Polyether
CON00072 | PEG.5 | C10H22O6 | 238.1416 | InChI=1S/C10H22O6/c11-1-3-13-5-7-15-9-10-16-8-6-14-4-2-12/h11-12H,1-10H2 | JLFNLZLINWHATN-UHFFFAOYSA-N | 62551 | Polyether
CON00073 | PEG.6 | C12H26O7 | 282.1679 | InChI=1S/C12H26O7/c13-1-3-15-5-7-17-9-11-19-12-10-18-8-6-16-4-2-14/h13-14H,1-12H2 | IIRDTKBZINWQAW-UHFFFAOYSA-N |  | Polyether
CON00074 | PEG.7 | C14H30O8 | 326.1941 | InChI=1S/C14H30O8/c15-1-3-17-5-7-19-9-11-21-13-14-22-12-10-20-8-6-18-4-2-16/h15-16H,1-14H2 | XPJRQAIZZQMSCM-UHFFFAOYSA-N | 79718 | Polyether
CON00075 | PEG.8 | C16H34O9 | 370.2203 | InChI=1S/C16H34O9/c17-1-3-19-5-7-21-9-11-23-13-15-25-16-14-24-12-10-22-8-6-20-4-2-18/h17-18H,1-16H2 | GLZWNFNQMJAZGY-UHFFFAOYSA-N | 78798 | Polyether
CON00076 | PEG.9 | C18H38O10 | 414.2465 | InChI=1S/C18H38O10/c19-1-3-21-5-7-23-9-11-25-13-15-27-17-18-28-16-14-26-12-10-24-8-6-22-4-2-20/h19-20H,1-18H2 | YZUUTMGDONTGTN-UHFFFAOYSA-N | 4867 | Polyether
CON00077 | Pentafluoropropionic acid | C3HF5O2 | 163.9897 | InChI=1S/C3HF5O2/c4-2(5,1(9)10)3(6,7)8/h(H,9,10) | LRMSQVBRUNSOJL-UHFFFAOYSA-N | 62356
CON00078 | Phenyldiethylamine | C10H15N | 149.1204 | InChI=1S/C10H15N/c1-3-11(4-2)10-8-6-5-7-9-10/h5-9H,3-4H2,1-2H3 | GGSUCNLOZRCGPQ-UHFFFAOYSA-N | 7061 | Solvent
CON00079 | Phosphoric acid | H3PO4 | 97.9769 | InChI=1S/H3O4P/c1-5(2,3)4/h(H3,1,2,3,4) | NBIIXXVUZAFLBC-UHFFFAOYSA-N | 1004
CON00079 | Phosphoric acid | H3PO4 | 97.9769 | InChI=1S/H3O4P/c1-5(2,3)4/h(H3,1,2,3,4) | NBIIXXVUZAFLBC-UHFFFAOYSA-N
CON00080 | Phthalic acid | C8H6O4 | 166.0266 | InChI=1S/C8H6O4/c9-7(10)5-3-1-2-4-6(5)8(11)12/h1-4H,(H,9,10)(H,11,12) | XNGIFLGASWRNHJ-UHFFFAOYSA-N | 1017
CON00081 | Phthalic anhydride | C8H4O3 | 148.016 | InChI=1S/C8H4O3/c9-7-5-3-1-2-4-6(5)8(10)11-7/h1-4H | LGRFSURHDFAFJT-UHFFFAOYSA-N | 6811 | Plasticiser
CON00082 | PPG.1 | C3H8O2 | 76.0524 | InChI=1S/C3H8O2/c1-3(5)2-4/h3-5H,2H2,1H3 | DNIAPMSPPWPWGF-UHFFFAOYSA-N | 439846 | Polyether
CON00083 | PPG.10 | C30H62O11 | 598.4292 |  |  |  | Polyether
CON00084 | PPG.11 | C33H68O12 | 656.4711 |  |  |  | Polyether
CON00085 | PPG.12 | C36H74O13 | 714.5129 |  |  |  | Polyether
CON00086 | PPG.13 | C39H80O14 | 772.5548 |  |  |  | Polyether
CON00087 | PPG.14 | C42H86O15 | 830.5967 |  |  |  | Polyether
CON00088 | PPG.15 | C45H92O16 | 888.6385 |  |  |  | Polyether
CON00089 | PPG.16 | C48H98O17 | 946.6804 |  |  |  | Polyether
CON00090 | PPG.17 | C51H104O18 | 1004.7223 |  |  |  | Polyether
CON00091 | PPG.18 | C54H110O19 | 1062.7641 |  |  |  | Polyether
CON00092 | PPG.19 | C57H116O20 | 1120.806 |  |  |  | Polyether
CON00093 | PPG.20 | C60H122O21 | 1178.8479 |  |  |  | Polyether
CON00094 | PPG.2 | C6H14O3 | 134.0943 | InChI=1S/C6H14O3/c1-5(8)4-9-6(2)3-7/h5-8H,3-4H2,1-2H3 | DUFKCOQISQKSAV-UHFFFAOYSA-N | 32881 | Polyether
CON00095 | PPG.3 | C9H20O4 | 192.1362 | InChI=1S/C9H20O4/c1-7(11)5-12-9(3)6-13-8(2)4-10/h7-11H,4-6H2,1-3H3 | LCZVSXRMYJUNFX-UHFFFAOYSA-N |  | Polyether
CON00096 | PPG.4 | C12H26O5 | 250.178 | InChI=1S/C12H26O5/c1-9(14)6-15-11(3)8-17-12(4)7-16-10(2)5-13/h9-14H,5-8H2,1-4H3 | QVHMSMOUDQXMRS-UHFFFAOYSA-N | 91234 | Polyether
CON00097 | PPG.5 | C15H32O6 | 308.2199 | InChI=1S/C15H32O6/c1-11(17)7-18-13(3)9-20-15(5)10-21-14(4)8-19-12(2)6-16/h11-17H,6-10H2,1-5H3 | AQRQHYITOOVBTO-UHFFFAOYSA-N |  | Polyether
CON00098 | PPG.6 | C18H38O7 | 366.2618 |  |  |  | Polyether
CON00099 | PPG.7 | C21H44O8 | 424.3036 | InChI=1S/C21H44O8/c1-15(22)8-24-10-17(3)26-12-19(5)28-14-21(7)29-13-20(6)27-11-18(4)25-9-16(2)23/h15-23H,8-14H2,1-7H3/t15-,16+,17+,18-,19+,20-,21+/m1/s1 |  |  | Polyether
CON00100 | PPG.8 | C24H50O9 | 482.3455 |  |  |  | Polyether
CON00101 | PPG.9 | C27H56O10 | 540.3873 |  |  |  | Polyether
CON00102 | Propionic acid | C3H6O2 | 74.0368 | InChI=1S/C3H6O2/c1-2-3(4)5/h2H2,1H3,(H,4,5) | XBDQKXXYIPTUBI-UHFFFAOYSA-N | 1032
CON00103 | Sodium acetate | C2H3O2Na | 82.0031 | InChI=1S/C2H4O2.Na/c1-2(3)4;/h1H3,(H,3,4);/q;+1/p-1 | VMHLLURERBWHNL-UHFFFAOYSA-M | 517045 | Solvent
CON00104 | Sodium trifluoroacetate | C2F3O2Na | 135.9748 | InChI=1S/C2HF3O2.Na/c3-2(4,5)1(6)7;/h(H,6,7);/q;+1/p-1 | UYCAUPASBSROMS-UHFFFAOYSA-M | 51709 | Salt
CON00105 | Stearamide | C18H37NO | 283.2875 | InChI=1S/C18H37NO/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18(19)20/h2-17H2,1H3,(H2,19,20) | LYRFLYHAGKPMFH-UHFFFAOYSA-N | 31292 | Slip agent
CON00106 | Stearic acid | C18H36O2 | 284.2715 | InChI=1S/C18H36O2/c1-2-3-4-5-6-7-8-9-10-11-12-13-14-15-16-17-18(19)20/ | QIQXTHQIDYTFRH-UHFFFAOYSA-N | 5281
CON00107 | Stearyl-palmityldimethylammonium chloride | C36H76NCl | 557.5666 | h2-17H2,1H3,(H,19,20) | UKAFOQJMXIZSMD-UHFFFAOYSA-M | 161744 | Personal care products
CON00108 | Sulphuric acid | H2SO4 | 97.9674 | InChI=1S/H2O4S/c1-5(2,3)4/h(H2,1,2,3,4) | QAOWNCQODCNURD-UHFFFAOYSA-N | 1118
CON00109 | Tetrabutylammonium (TBA) | C16H36N | 242.2848 | InChI=1S/C16H36N/c1-5-9-13-17(14-10-6-2,15-11-7-3)16-12-8-4/h5-16H2,1-4H3/q+1 | DZLFLBLQUQXARW-UHFFFAOYSA-N | 70681 | Buffer
CON00110 | Tetradecamethylcycloheptasiloxane | C14H42Si7O7 | 518.1315 | InChI=1S/C14H42O7Si7/c1-22(2)15-23(3,4)17-25(7,8)19-27(11,12)21-28(13,14)20-26(9,10)18-24(5,6)16-22/h1-14H3 | GSANOGQCVHBHIF-UHFFFAOYSA-N | 7874 | Airborne Contaminant
CON00111 | Tributyl Tin Formate | C13H28O2Sn | 336.1111 | InChI=1S/3C4H9.CH2O2.Sn/c3*1-3-4-2;2-1-3;/h3*1,3-4H2,2H3;1H,(H,2,3);/q;;;;+1/p-1 | POUGAXCXKDLEMK-UHFFFAOYSA-M | 22167734 | Catalyst
CON00112 | Tributylamine | C12H27N | 185.2143 | InChI=1S/C12H27N/c1-4-7-10-13(11-8-5-2)12-9-6-3/h4-12H2,1-3H3 | IMFACGCPASFAPR-UHFFFAOYSA-N | 7622 | Solvent
CON00113 | Tributylphosphate | C12H27O4P | 266.1647 | InChI=1S/C12H27O4P/c1-4-7-10-14-17(13,15-11-8-5-2)16-12-9-6-3/h4-12H2,1-3H3 | STCOOQWBFONSKY-UHFFFAOYSA-N | 31357 | Plasticiser
CON00114 | Triethylamine | C6H15N | 101.1204 | InChI=1S/C6H15N/c1-4-7(5-2)6-3/h4-6H2,1-3H3 | ZMANZCXQSJIPKH-UHFFFAOYSA-N | 8471 | Buffer
CON00115 | Triethylamine Hydrochloride (35) | C6H15N | 101.1204 | InChI=1S/C6H15N/c1-4-7(5-2)6-3/h4-6H2,1-3H3 | ZMANZCXQSJIPKH-UHFFFAOYSA-N | 8471 | Buffer
CON00116 | Triethylamine Hydrochloride (37) | C6H15N | 101.1204 | InChI=1S/C6H15N/c1-4-7(5-2)6-3/h4-6H2,1-3H3 | ZMANZCXQSJIPKH-UHFFFAOYSA-N | 8471 | Buffer
CON00117 | Trifluoroacetic acid | C2HO2F3 | 113.9929 | InChI=1S/C2HF3O2/c3-2(4,5)1(6)7/h(H,6,7) | DTQVDTLACAAQTR-UHFFFAOYSA-N | 6422
CON00118 | Triphenyl phosphate | C18H15O4P | 326.0708 | InChI=1S/C18H15O4P/c19-23(20-16-10-4-1-5-11-16,21-17-12-6-2-7-13-17)22-18-14-8-3-9-15-18/h1-15H | XZZNDPSIHUTMOC-UHFFFAOYSA-N | 8289 | Flame retardant in plastics
CON00119 | Triphenylphosphine oxide | C18H15OP | 278.0861 | InChI=1S/C18H15OP/c19-20(16-10-4-1-5-11-16,17-12-6-2-7-13-17)18-14-8-3-9-15-18/h1-15H | FIQMHBFVRAXMOP-UHFFFAOYSA-N | 13097 | Crystallizing agent
CON00120 | Tripropylamine | C9H21N | 143.1674 | InChI=1S/C9H21N/c1-4-7-10(8-5-2)9-6-3/h4-9H2,1-3H3 | YFTHZRPMJXBUME-UHFFFAOYSA-N | 7616 | Solvent
CON00121 | TRIS | C4H11NO3 | 121.0739 | InChI=1S/C4H11NO3/c5-4(1-6,2-7)3-8/h6-8H,1-3,5H2 | LENZDBCJOHFCAS-UHFFFAOYSA-N | 6503 | Buffer
CON00122 | Triton X-100.1 | C16H26O2 | 250.1933 | InChI=1S/C16H26O2/c1-15(2,3)12-16(4,5)13-6-8-14(9-7-13)18-11-10-17/h6-9,17H,10-12H2,1-5H3 | JYCQQPHGFMYQCF-UHFFFAOYSA-N | 5590 | Buffer
CON00122 | Triton X-100.1 | C16H26O2 | 250.1933 | InChI=1S/C16H26O2/c1-15(2,3)12-16(4,5)13-6-8-14(9-7-13)18-11-10-17/h6-9,17H,10-12H2,1-5H3 | JYCQQPHGFMYQCF-UHFFFAOYSA-N |  | Buffer
CON00123 | Triton X-100.10 | C34H62O11 | 646.4292 | InChI=1S/C34H62O11/c1-33(2,3)30-34(4,5)31-6-8-32(9-7-31)45-29-28-44-27-26-43-25-24-42-23-22-41-21-20-40-19-18-39-17-16-38-15-14-37-13-12-36-11-10-35/h6-9,35H,10-30H2,1-5H3 | IVKNZCBNXPYYKL-UHFFFAOYSA-N |  | Buffer
CON00124 | Triton X-100.11 | C36H66O12 | 690.4554 |  |  |  | Buffer
CON00125 | Triton X-100.12 | C38H70O13 | 734.4816 |  |  |  | Buffer
CON00126 | Triton X-100.13 | C40H74O14 | 778.5079 |  |  |  | Buffer
CON00127 | Triton X-100.14 | C42H78O15 | 822.5341 |  |  |  | Buffer
CON00128 | Triton X-100.15 | C44H82O16 | 866.5603 |  |  |  | Buffer
CON00129 | Triton X-100.16 | C46H86O17 | 910.5865 |  |  |  | Buffer
CON00130 | Triton X-100.17 | C48H90O18 | 954.6127 |  |  |  | Buffer
CON00131 | Triton X-100.18 | C50H94O19 | 998.6389 |  |  |  | Buffer
CON00132 | Triton X-100.19 | C52H98O20 | 1042.6652 |  |  |  | Buffer
CON00133 | Triton X-100.20 | C54H102O21 | 1086.6914 |  |  |  | Buffer
CON00134 | Triton X-100.2 | C18H30O3 | 294.2195 | InChI=1S/C18H30O3/c1-17(2,3)14-18(4,5)15-6-8-16(9-7-15)21-13-12-20-11-10-19/h6-9,19H,10-14H2,1-5H3 | LBCZOTMMGHGTPH-UHFFFAOYSA-N | 24775 | Buffer
CON00134 | Triton X-100.2 | C18H30O3 | 294.2195 |  |  |  | Buffer
CON00135 | Triton X-100.3 | C20H34O4 | 338.2457 | InChI=1S/C20H34O4/c1-19(2,3)16-20(4,5)17-6-8-18(9-7-17)24-15-14-23-13-12-22-11-10-21/h6-9,21H,10-16H2,1-5H3 | OLOHNJXENGJYTB-UHFFFAOYSA-N | 448876 | Buffer
CON00135 | Triton X-100.3 | C20H34O4 | 338.2457 |  |  |  | Buffer
CON00136 | Triton X-100.4 | C22H38O5 | 382.2719 | InChI=1S/C22H38O5/c1-21(2,3)18-22(4,5)19-6-8-20(9-7-19)27-17-16-26-15-14-25-13-12-24-11-10-23/h6-9,23H,10-18H2,1-5H3 | UYDLBVPAAFVANX-UHFFFAOYSA-N | 628327 | Buffer
CON00136 | Triton X-100.4 | C22H38O5 | 382.2719 |  |  |  | Buffer
CON00137 | Triton X-100.5 | C24H42O6 | 426.2981 | InChI=1S/C24H42O6/c1-23(2,3)20-24(4,5)21-6-8-22(9-7-21)30-19-18-29-17-16-28-15-14-27-13-12-26-11-10-25/h6-9,25H,10-20H2,1-5H3 | WQBPSJIAFZCNBR-UHFFFAOYSA-N | 560774 | Buffer
CON00137 | Triton X-100.5 | C24H42O6 | 426.2981 |  |  |  | Buffer
CON00138 | Triton X-100.6 | C26H46O7 | 470.3244 | InChI=1S/C26H46O7/c1-25(2,3)22-26(4,5)23-6-8-24(9-7-23)33-21-20-32-19-18-31-17-16-30-15-14-29-13-12-28-11-10-27/h6-9,27H,10-22H2,1-5H3 | SQAUJCLQNBQZMT-UHFFFAOYSA-N | 17245 | Buffer
CON00138 | Triton X-100.6 | C26H46O7 | 470.3244 |  |  |  | Buffer
CON00139 | Triton X-100.7 | C28H50O8 | 514.3506 | InChI=1S/C28H50O8/c1-27(2,3)24-28(4,5)25-6-8-26(9-7-25)36-23-22-35-21-20-34-19-18-33-17-16-32-15-14-31-13-12-30-11-10-29/h6-9,29H,10-24H2,1-5H3 | HNLXNOZHXNSSPN-UHFFFAOYSA-N | 75622 | Buffer
CON00139 | Triton X-100.7 | C28H50O8 | 514.3506 |  |  |  | Buffer
CON00140 | Triton X-100.8 | C30H54O9 | 558.3768 |  |  |  | Buffer
CON00141 | Triton X-100.9 | C32H58O10 | 602.403 | InChI=1S/C32H58O10/c1-31(2,3)28-32(4,5)29-6-8-30(9-7-29)42-27-26-41-25-24-40-23-22-39-21-20-38-19-18-37-17-16-36-15-14-35-13-12-34-11-10-33/h6-9,33H,10-28H2,1-5H3 | VJYAJQFKKLYARJ-UHFFFAOYSA-N |  | Buffer
CON00141 | Triton X-100.9 | C32H58O10 | 602.403 |  |  |  | Buffer
CON00142 | Triton X-100R.1 | C16H32O2 | 256.2402 |  |  |  | Buffer
CON00143 | Triton X-100R.10 | C34H68O11 | 652.4762 | InChI=1S/C34H68O11/c1-33(2,3)30-34(4,5)31-6-8-32(9-7-31)45-29-28-44-27-26-43-25-24-42-23-22-41-21-20-40-19-18-39-17-16-38-15-14-37-13-12-36-11-10-35/h31-32,35H,6-30H2,1-5H3 | AXMJGXMRGDCRND-UHFFFAOYSA-N |  | Buffer
CON00144 | Triton X-100R.11 | C36H72O12 | 696.5024 |  |  |  | Buffer
CON00145 | Triton X-100R.12 | C38H76O13 | 740.5286 |  |  |  | Buffer
CON00146 | Triton X-100R.13 | C40H80O14 | 784.5548 |  |  |  | Buffer
CON00147 | Triton X-100R.14 | C42H84O15 | 828.581 |  |  |  | Buffer
CON00148 | Triton X-100R.15 | C44H88O16 | 872.6072 |  |  |  | Buffer
CON00149 | Triton X-100R.16 | C46H92O17 | 916.6335 |  |  |  | Buffer
CON00150 | Triton X-100R.17 | C48H96O18 | 960.6597 |  |  |  | Buffer
CON00151 | Triton X-100R.18 | C50H100O19 | 1004.6859 |  |  |  | Buffer
CON00152 | Triton X-100R.19 | C52H104O20 | 1048.7121 |  |  |  | Buffer
CON00153 | Triton X-100R.20 | C54H108O21 | 1092.7383 |  |  |  | Buffer
CON00154 | Triton X-100R.2 | C18H36O3 | 300.2664 | InChI=1S/C18H36O3/c1-17(2,3)14-18(4,5)15-6-8-16(9-7-15)21-13-12-20-11-10-19/h15-16,19H,6-14H2,1-5H3 | PNAKTNFERWFEIX-UHFFFAOYSA-N |  | Buffer
CON00155 | Triton X-100R.3 | C20H40O4 | 344.2927 |  |  |  | Buffer
CON00156 | Triton X-100R.4 | C22H44O5 | 388.3189 |  |  |  | Buffer
CON00157 | Triton X-100R.5 | C24H48O6 | 432.3451 |  |  |  | Buffer
CON00158 | Triton X-100R.6 | C26H52O7 | 476.3713 |  |  |  | Buffer
CON00159 | Triton X-100R.7 | C28H56O8 | 520.3975 |  |  |  | Buffer
CON00160 | Triton X-100R.8 | C30H60O9 | 564.4237 | InChI=1S/C30H60O9/c1-29(2,3)26-30(4,5)27-6-8-28(9-7-27)39-25-24-38-23-22-37-21-20-36-19-18-35-17-16-34-15-14-33-13-12-32-11-10-31/h27-28,31H,6-26H2,1-5H3 | LVHZSSVOODJLAY-UHFFFAOYSA-N |  | Buffer
CON00161 | Triton X-100R.9 | C32H64O10 | 608.4499 |  |  |  | Buffer
CON00162 | Triton X-101.1 | C17H28O2 | 264.2089 |  |  |  | Detergent
CON00163 | Triton X-101.10 | C35H64O11 | 660.4449 |  |  |  | Detergent
CON00164 | Triton X-101.11 | C37H68O12 | 704.4711 |  |  |  | Detergent
CON00165 | Triton X-101.12 | C39H72O13 | 748.4973 |  |  |  | Detergent
CON00166 | Triton X-101.13 | C41H76O14 | 792.5235 |  |  |  | Detergent
CON00167 | Triton X-101.14 | C43H80O15 | 836.5497 |  |  |  | Detergent
CON00168 | Triton X-101.15 | C45H84O16 | 880.5759 |  |  |  | Detergent
CON00169 | Triton X-101.16 | C47H88O17 | 924.6022 |  |  |  | Detergent
CON00170 | Triton X-101.17 | C49H92O18 | 968.6284 |  |  |  | Detergent
CON00171 | Triton X-101.18 | C51H96O19 | 1012.6546 |  |  |  | Detergent
CON00172 | Triton X-101.19 | C53H100O20 | 1056.6808 |  |  |  | Detergent
CON00173 | Triton X-101.20 | C55H104O21 | 1100.707 |  |  |  | Detergent
CON00174 | Triton X-101.2 | C19H32O3 | 308.2351 |  |  |  | Detergent
CON00175 | Triton X-101.3 | C21H36O4 | 352.2614 |  |  |  | Detergent
CON00176 | Triton X-101.4 | C23H40O5 | 396.2876 |  |  |  | Detergent
CON00177 | Triton X-101.5 | C25H44O6 | 440.3138 |  |  |  | Detergent
CON00178 | Triton X-101.6 | C27H48O7 | 484.34 |  |  |  | Detergent
CON00179 | Triton X-101.7 | C29H52O8 | 528.3662 |  |  |  | Detergent
CON00180 | Triton X-101.8 | C31H56O9 | 572.3924 |  |  |  | Detergent
CON00181 | Triton X-101.9 | C33H60O10 | 616.4186 |  |  |  | Detergent
CON00182 | Triton X-101R.1 | C17H34O2 | 270.2559 |  |  |  | Detergent
CON00183 | Triton X-101R.10 | C35H70O11 | 666.4918 |  |  |  | Detergent
CON00184 | Triton X-101R.11 | C37H74O12 | 710.518 |  |  |  | Detergent
CON00185 | Triton X-101R.12 | C39H78O13 | 754.5442 |  |  |  | Detergent
CON00186 | Triton X-101R.13 | C41H82O14 | 798.5705 |  |  |  | Detergent
CON00187 | Triton X-101R.14 | C43H86O15 | 842.5967 |  |  |  | Detergent
CON00188 | Triton X-101R.15 | C45H90O16 | 886.6229 |  |  |  | Detergent
CON00189 | Triton X-101R.16 | C47H94O17 | 930.6491 |  |  |  | Detergent
CON00190 | Triton X-101R.17 | C49H98O18 | 974.6753 |  |  |  | Detergent
CON00191 | Triton X-101R.18 | C51H102O19 | 1018.7015 |  |  |  | Detergent
CON00192 | Triton X-101R.19 | C53H106O20 | 1062.7278 |  |  |  | Detergent
CON00193 | Triton X-101R.20 | C55H110O21 | 1106.754 |  |  |  | Detergent
CON00194 | Triton X-101R.2 | C19H38O3 | 314.2821 |  |  |  | Detergent
CON00195 | Triton X-101R.3 | C21H42O4 | 358.3083 |  |  |  | Detergent
CON00196 | Triton X-101R.4 | C23H46O5 | 402.3345 |  |  |  | Detergent
CON00197 | Triton X-101R.5 | C25H50O6 | 446.3607 |  |  |  | Detergent
CON00198 | Triton X-101R.6 | C27H54O7 | 490.387 |  |  |  | Detergent
CON00199 | Triton X-101R.7 | C29H58O8 | 534.4132 |  |  |  | Detergent
CON00200 | Triton X-101R.8 | C31H62O9 | 578.4394 |  |  |  | Detergent
CON00201 | Triton X-101R.9 | C33H66O10 | 622.4656 |  |  |  | Detergent
CON00202 | Trityl cation | C19H15 | 243.1174 | InChI=1S/C19H15/c1-4-10-16(11-5-1)19(17-12-6-2-7-13-17)18-14-8-3-9-15-18/h1-15H/q+1 | OLFPYUPGPBITMH-UHFFFAOYSA-N | 84899 | Catalyst
CON00203 | Tween 20.10 | C38H74O16 | 786.4977 |  |  |  | Buffer
CON00204 | Tween 20.11 | C40H78O17 | 830.5239 |  |  |  | Buffer
CON00205 | Tween 20.12 | C42H82O18 | 874.5501 |  |  |  | Buffer
CON00206 | Tween 20.13 | C44H86O19 | 918.5763 |  |  |  | Buffer
CON00207 | Tween 20.14 | C46H90O20 | 962.6025 |  |  |  | Buffer
CON00208 | Tween 20.15 | C48H94O21 | 1006.6288 |  |  |  | Buffer
CON00209 | Tween 20.16 | C50H98O22 | 1050.655 |  |  |  | Buffer
CON00210 | Tween 20.17 | C52H102O23 | 1094.6812 |  |  |  | Buffer
CON00211 | Tween 20.18 | C54H106O24 | 1138.7074 |  |  |  | Buffer
CON00212 | Tween 20.19 | C56H110O25 | 1182.7336 |  |  |  | Buffer
CON00213 | Tween 20.20 | C58H114O26 | 1226.7598 |  |  |  | Buffer
CON00214 | Tween 20.21 | C60H118O27 | 1270.7861 |  |  |  | Buffer
CON00215 | Tween 20.22 | C62H122O28 | 1314.8123 |  |  |  | Buffer
CON00216 | Tween 20.23 | C64H126O29 | 1358.8385 |  |  |  | Buffer
CON00217 | Tween 20.24 | C66H130O30 | 1402.8647 |  |  |  | Buffer
CON00218 | Tween 20.25 | C68H134O31 | 1446.8909 |  |  |  | Buffer
CON00219 | Tween 20.26 | C70H138O32 | 1490.9171 |  |  |  | Buffer
CON00220 | Tween 20.27 | C72H142O33 | 1534.9433 |  |  |  | Buffer
CON00221 | Tween 20.28 | C74H146O34 | 1578.9696 |  |  |  | Buffer
CON00222 | Tween 20.29 | C76H150O35 | 1622.9958 |  |  |  | Buffer
CON00223 | Tween 20.30 | C78H154O36 | 1667.022 |  |  |  | Buffer
CON00224 | Tween 20.31 | C80H158O37 | 1711.0482 |  |  |  | Buffer
CON00225 | Tween 20.32 | C82H162O38 | 1755.0744 |  |  |  | Buffer
CON00226 | Tween 40.10 | C42H82O16 | 842.5603 |  |  |  | Buffer
CON00227 | Tween 40.11 | C44H86O17 | 886.5865 |  |  |  | Buffer
CON00228 | Tween 40.12 | C46H90O18 | 930.6127 |  |  |  | Buffer
CON00229 | Tween 40.13 | C48H94O19 | 974.6389 |  |  |  | Buffer
CON00230 | Tween 40.14 | C50H98O20 | 1018.6652 |  |  |  | Buffer
CON00231 | Tween 40.15 | C52H102O21 | 1062.6914 |  |  |  | Buffer
CON00232 | Tween 40.16 | C54H106O22 | 1106.7176 |  |  |  | Buffer
CON00233 | Tween 40.17 | C56H110O23 | 1150.7438 |  |  |  | Buffer
CON00234 | Tween 40.18 | C58H114O24 | 1194.77 |  |  |  | Buffer
CON00235 | Tween 40.19 | C60H118O25 | 1238.7962 |  |  |  | Buffer
CON00236 | Tween 40.20 | C62H122O26 | 1282.8224 |  |  |  | Buffer
CON00237 | Tween 40.21 | C64H126O27 | 1326.8487 |  |  |  | Buffer
CON00238 | Tween 40.22 | C66H130O28 | 1370.8749 |  |  |  | Buffer
CON00239 | Tween 40.23 | C68H134O29 | 1414.9011 |  |  |  | Buffer
CON00240 | Tween 40.24 | C70H138O30 | 1458.9273 |  |  |  | Buffer
CON00241 | Tween 40.25 | C72H142O31 | 1502.9535 |  |  |  | Buffer
CON00242 | Tween 40.26 | C74H146O32 | 1546.9797 |  |  |  | Buffer
CON00243 | Tween 40.27 | C76H150O33 | 1591.0059 |  |  |  | Buffer
CON00244 | Tween 40.28 | C78H154O34 | 1635.0322 |  |  |  | Buffer
CON00245 | Tween 40.29 | C80H158O35 | 1679.0584 |  |  |  | Buffer
CON00246 | Tween 40.30 | C82H162O36 | 1723.0846 |  |  |  | Buffer
CON00247 | Tween 40.31 | C84H166O37 | 1767.1108 |  |  |  | Buffer
CON00248 | Tween 40.32 | C86H170O38 | 1811.137 |  |  |  | Buffer
CON00249 | Tween 60.10 | C44H86O16 | 870.5916 |  |  |  | Buffer
CON00250 | Tween 60.11 | C46H90O17 | 914.6178 |  |  |  | Buffer
CON00251 | Tween 60.12 | C48H94O18 | 958.644 |  |  |  | Buffer
CON00252 | Tween 60.13 | C50H98O19 | 1002.6702 |  |  |  | Buffer
CON00253 | Tween 60.14 | C52H102O20 | 1046.6964 |  |  |  | Buffer
CON00254 | Tween 60.15 | C54H106O21 | 1090.7227 |  |  |  | Buffer
CON00255 | Tween 60.16 | C56H110O22 | 1134.7489 |  |  |  | Buffer
CON00256 | Tween 60.17 | C58H114O23 | 1178.7751 |  |  |  | Buffer
CON00257 | Tween 60.18 | C60H118O24 | 1222.8013 |  |  |  | Buffer
CON00258 | Tween 60.19 | C62H122O25 | 1266.8275 |  |  |  | Buffer
CON00259 | Tween 60.20 | C64H126O26 | 1310.8537 |  |  |  | Buffer
CON00260 | Tween 60.21 | C66H130O27 | 1354.88 |  |  |  | Buffer
CON00261 | Tween 60.22 | C68H134O28 | 1398.9062 |  |  |  | Buffer
CON00262 | Tween 60.23 | C70H138O29 | 1442.9324 |  |  |  | Buffer
CON00263 | Tween 60.24 | C72H142O30 | 1486.9586 |  |  |  | Buffer
CON00264 | Tween 60.25 | C74H146O31 | 1530.9848 |  |  |  | Buffer
CON00265 | Tween 60.26 | C76H150O32 | 1575.011 |  |  |  | Buffer
CON00266 | Tween 60.27 | C78H154O33 | 1619.0372 |  |  |  | Buffer
CON00267 | Tween 60.28 | C80H158O34 | 1663.0635 |  |  |  | Buffer
CON00268 | Tween 60.29 | C82H162O35 | 1707.0897 |  |  |  | Buffer
CON00269 | Tween 60.30 | C84H166O36 | 1751.1159 |  |  |  | Buffer
CON00270 | Tween 60.31 | C86H170O37 | 1795.1421 |  |  |  | Buffer
CON00271 | Tween 60.32 | C88H174O38 | 1839.1683 |  |  |  | Buffer
CON00272 | Tween 80.10 | C44H84O16 | 868.5759 |  |  |  | Buffer
CON00273 | Tween 80.11 | C46H88O17 | 912.6022 |  |  |  | Buffer
CON00274 | Tween 80.12 | C48H92O18 | 956.6284 |  |  |  | Buffer
CON00275 | Tween 80.13 | C50H96O19 | 1000.6546 |  |  |  | Buffer
CON00276 | Tween 80.14 | C52H100O20 | 1044.6808 |  |  |  | Buffer
CON00277 | Tween 80.15 | C54H104O21 | 1088.707 |  |  |  | Buffer
CON00278 | Tween 80.16 | C56H108O22 | 1132.7332 |  |  |  | Buffer
CON00279 | Tween 80.17 | C58H112O23 | 1176.7594 |  |  |  | Buffer
CON00280 | Tween 80.18 | C60H116O24 | 1220.7857 |  |  |  | Buffer
CON00281 | Tween 80.19 | C62H120O25 | 1264.8119 |  |  |  | Buffer
CON00282 | Tween 80.20 | C64H124O26 | 1308.8381 |  |  |  | Buffer
CON00283 | Tween 80.21 | C66H128O27 | 1352.8643 |  |  |  | Buffer
CON00284 | Tween 80.22 | C68H132O28 | 1396.8905 |  |  |  | Buffer
CON00285 | Tween 80.23 | C70H136O29 | 1440.9167 |  |  |  | Buffer
CON00286 | Tween 80.24 | C72H140O30 | 1484.943 |  |  |  | Buffer
CON00287 | Tween 80.25 | C74H144O31 | 1528.9692 |  |  |  | Buffer
CON00288 | Tween 80.26 | C76H148O32 | 1572.9954 |  |  |  | Buffer
CON00289 | Tween 80.27 | C78H152O33 | 1617.0216 |  |  |  | Buffer
CON00290 | Tween 80.28 | C80H156O34 | 1661.0478 |  |  |  | Buffer
CON00291 | Tween 80.29 | C82H160O35 | 1705.074 |  |  |  | Buffer
CON00292 | Tween 80.30 | C84H164O36 | 1749.1002 |  |  |  | Buffer
CON00293 | Tween 80.31 | C86H168O37 | 1793.1265 |  |  |  | Buffer
CON00294 | Tween 80.32 | C88H172O38 | 1837.1527 |  |  |  | Buffer
CON00295 | 1,3-Dicyclohexylurea | C13H24N2O | 224.1889 | InChI=1S/C13H24N2O/c16-13(14-11-7-3-1-4-8-11)15-12-9-5-2-6-10-12/h11-12H,1-10H2,(H2,14,15,16) | ADFXKUOMJKEIND-UHFFFAOYSA-N | 4277
CON00296 | 1,8-Diazabicycloundec-7-ene | C9H16N2 | 152.1313 | InChI=1S/C9H16N2/c1-2-5-9-10-6-4-8-11(9)7-3-1/h1-8H2 | GQHTUMJGOHRCHB-UHFFFAOYSA-N | 81184 | Catalyst
CON00297 | 3-((3-cholamidopropyl)dimethylammonium)-1-propanesulfonate (CHAPS) | C32H58N2O7S | 614.3965 | InChI=1S/C32H58N2O7S/c1-21(8-11-29(38)33-14-6-15-34(4,5)16-7-17-42(39,40)41)24-9-10-25-30-26(20-28(37)32(24,25)3)31(2)13-12-23(35)18-22(31)19-27(30)36/h21-28,30,35-37H,6-20H2,1-5H3,(H-,33,38,39,40,41)/t21-,22+,23-,24-,25+,26+,27-,28+,30+,31+,32-/m1/s1 | UMCMPZBLKLEWAF-BCTGSCMUSA-N | 107670 | Solvent
CON00298 | 4-HCCA | C10H7NO3 | 189.0426 | InChI=1S/C10H7NO3/c11-6-8(10(13)14)5-7-1-3-9(12)4-2-7/h1-5,12H,(H,13,14)/b8-5+ | AFVLVVWMAFSXCK-VMPITWQZSA-N | 5328791 | Matrix compound

## NewObjective
New Objective | common background ions [pdf] | http://www.newobjective.com/downloads/technotes/PV-3.pdf
I N N O V A T I O N I N H I G H - P E R F O R M A N C E L C - M S A N A L Y S I S
1 888 220 2998 www.newobjective.com |  | Technical Note PV-3
Common Background Ions for Electrospray
(Positive Ion)
Ion | Compound
101 | (M+Na)+ DMSO
102 | (M+H)+ triethylamine (TEA)
104/106 | (M+Cu)+ acetonitrile
105 | (2M+Na)+ acetonitrile
120 | (M+Na+CH3CN)+ DMSO
122 | (M+H)+ Tris
123 | (M+H)+ dimethylaminopyridine (DMAP)
130 | (M+H)+ diisopropylethylamine (DIPEA)
137 | (M+ACN+NH4)+ DMSO
144 | (M+H)+ tripropylamine (TPA)
145/147 | (2M+Cu)+ acetonitrile
146 | (3M+Na)+ acetonitrile
150 | (M+H)+ phenyldiethylamine
153 | (M+H)+ 1,8-diazabicyclo[5.4.0]undec-7-ene (DBU)
157 | (2M+H)+ DMSO
159 | (M+Na)+ sodium trifluoroacetate
179 | (2M+Na)+ DMSO
186 | (M+H)+ tributylamine
214 | n-butyl benzenesulfonamide (plastisizer)
217 | Unidentified contaminant
225 | (M+H)+ dicyclohexyl urea (DCU)
231 | (M+NH3)+ Unknown†
239/241 | [(M.HCl)2-Cl]+ triethylamine
242 | M+ tetrabutylammonium (C4H9)4N+
243 | M+ trityl cation
257 | (3M+H)+ DMSO
267 | (M+H)+ tributylphosphate
273 | M+ monomethoxytrityl cation (MMT)
279 | (M+H)+ dibutylphthalate (plasticiser)
282 | plasticizer in polyethylene
301 | (M+Na)+ dibutylphthalate (plasticiser)
317 | (M+K)+ dibutylphthalate (plasticiser)
336 | (M+H)+ tributyl tin formate
338 | (M+H)+ erucamide
360 | (M+Na)+ erucamide
371 | (M+H)+ polysiloxane, followed by 388
391 | (M+H)+ diisooctyl phthalate (plasticiser)
413 | (M+Na)+ diisooctyl phthalate (plasticiser)
425 | (M+Na)+ unidentified contaminant (plasticiser)
429 | (M+K)+ diisooctyl phthalate (plasticiser)
445 | (M+H)+ polysiloxane, followed by 462
449 | (2M+H)+ dicyclohexyl urea (DCU)
454 | (M+Na+CH3CN)+ diisooctyl phthalate (plasticiser)
522 | Unknown†
550 | Unknown†
599 | Unknown
798 | (2M+NH4)+ diisooctyl phthalate (plasticiser)
803 | (2M+Na)+ diisooctyl phthalate (plasticiser)
879 | Unknown
933 | Unknown
PEG polymers exhibit peaks 44 amu apart
Common Adducts:
(M+NH4)+ (M+18)
(M+Na)+ (M+23)
(M+K)+ (M+39)
The information contained in this document is believed reliable and accurate; however, nothing set forth herein constitutes a warranty or representation of any kind or
nature. ©2007 New Objective, Inc. All rights reserved.
* See Andreas Schlosser and Rudolf Volkmer-Engert “Volatile polydimethylcyclosiloxanes in the ambient laboratory air identified as source of extreme background signals in
nanoelectrospray mass spectrometry”, J. Mass. Spectrom. 2003; 38: 523-525 for more information on these common background ions.
† Derived from rubber tip from disposable syringe plunger
Revised 07/09/09

## Waters
Waters | http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf
Rev. A BACKGROUND ION LIST
ESI+ Common Background Ions
| m/z | Ion | Compound | Source of Unknown
ESI+ | 33 | (M+H)+ | methanol
ESI+ | 42 | (M+H)+ | acetonitrile
ESI+ | 59 | (M+NH4)+ | acetonitrile
ESI+ | 64 | (M+Na)+ | acetonitrile
ESI+ | 65 | (2M+H)+ | methanol
ESI+ | 74 | (M+H)+ | dimethyl formamide
ESI+ | 79 | (M+H)+ | DMSO
ESI+ | 83 | (2M+H)+ | Acetonitrile
ESI+ | 85 | (M+H)+ | d6-DMSO
ESI+ | 88 | (M+Formic Acid+H)+ | acetonitrile
ESI+ | 101 | (M+Na)+ | DMSO
ESI+ | 102 | (M+H)+ | triethylamine (TEA)
ESI+ | 104 | (M+Cu)+ | acetonitrile
ESI+ | 106 | (M+Cu)+ | acetonitrile
ESI+ | 105 | (2M+Na)+ | acetonitrile
ESI+ | 115 | (M+dimethylformamide+H)+ | acetonitrile
ESI+ | 120 | (M+Na+CH3CN)+ | DMSO
ESI+ | 122 | (M+H)+ | tris, tris(hydroxymethyl)aminomethane
ESI+ | 123 | (M+H)+ | dimethylaminopyridine (DMAP)
ESI+ | 130 | (M+H)+ | diisopropylethyl amine (DIPEA)
ESI+ | 137 | (M+CH3CN+NH4)+ | DMSO
ESI+ | 144 | (M+H)+ | tripropylamine (TPA)
ESI+ | 145/147 | (2M+Cu)+ | acetonitrile
ESI+ | 146 | (3M+Na)+ | acetonitrile
ESI+ | 149 | (M+H)+ | phthalic anhydride
ESI+ | 150 | (M+H)+ | phenyldiethylamine
ESI+ | 153 | (M+H)+ | 1,8-diazabicyclo[5.4.0]undec-7-ene (DBU)
ESI+ | 157 | (2M+H)+ DMSO
ESI+ | 159 | (M+Na)+ sodium trifluoroacetate
ESI+ | 163 |  | dimethyl phthalate
ESI+ | 167 |  | dioctyl phthalate
ESI+ | 169 | (2M+H)+ d6-DMSO
ESI+ | 171 | (M+Na)+ phthalic anhydride
ESI+ | 179 | (2M+Na)+ DMSO
ESI+ | 186 | (M+H)+ tributylamine
ESI+ | 195 | (M+H)+ dimethyl phthalate
ESI+ | 214 | (M+H)+ n-butyl benzenesulfonamide (plastisizer)
ESI+ | 225 | (M+H)+ dicyclohexyl urea (DCU)
ESI+ | 231 | (M+NH3)+ n-butyl benzenesulfonamide (plastisizer)
ESI+ | 236 | (M+Na)+ n-butyl benzenesulfonamide (plastisizer)
ESI+ | 239/241 | [(M.HCl)2-Cl]+ triethylamine
ESI+ | 242 | M+ tetrabutylammonium (C4H9)4N+
ESI+ | 243 | M+ trityl cation
ESI+ | 257 | (3M+H)+ DMSO
ESI+ | 267 | (M+H)+ tributylphosphate
ESI+ | 273 | M+ monomethoxytrityl cation (MMT)
ESI+ | 279 | M+ triphenylphosphine oxide
ESI+ | 279 | M+ dibutylphthalate (plasticiser)
ESI+ | 282 | (M+H)+ oleamide slip agent in polyethylene films
ESI+ | 284 | (M+H)+ stearamide slip agent in polyethylene films
ESI+ | 288 | (M+H)+ n,n-bis(2-hydroxyethyl) dodecanamide anti-static agent
ESI+ | 301 | (M+Na)+ dibutylphthalate (plasticiser)
ESI+ | 303 | (M+Na)+ oleamide slip agent in polyethylene films
ESI+ | 317 | (M+K)+ dibutylphthalate (plasticiser)
ESI+ | 336 | (M+H)+ tributyl tin formate
ESI+ | 338 | (M+H)+ erucamide slip agent in polyethylene films
ESI+ | 355 | siloxane from silica column
ESI+ | 360 | (M+Na)+ erucamide slip agent in polyethylene films
ESI+ | 371 | (M+H)+ decamethylcyclopentasiloxane
ESI+ | 388 | decamethylcyclopentasiloxane
ESI+ | 391 | (M+H)+ diisooctyl phthalate (plasticiser)
ESI+ | 413 | (M+Na)+ diisooctyl phthalate (plasticiser)
ESI+ | 425 | (M+Na)+ unidentified contaminant (plasticiser)
ESI+ | 429 | (M+K)+ diisooctyl phthalate (plasticiser) seen w/ silica capillary NanoSprayer
ESI+ | 445 | (M+H)+ dodecamethylcyclohexasiloxane seen w/ silica capillary NanoSprayer
ESI+ | 449 | (2M+H)+ dicyclohexyl urea (DCU)
ESI+ | 454 | (M+Na+CH3CN)+ diisooctyl phthalate (plasticiser) also observed w/ high conc. of formic acid
ESI+ | 462 | dodecamethylcyclohexasiloxane
ESI+ | 503 | seen w/ silica capillary NanoSprayer
ESI+ | 519 | (M+H)+ tetradecamethylcycloheptasiloxane seen w/ silica capillary NanoSprayer
ESI+ | 522 |  | Derived from rubber tip from disposable syringe plunger
ESI+ | 531 | (M+H)+ seen w/ Nunc polypropylene tubes
ESI+ | 536 | seen w/ silica capillary NanoSprayer
ESI+ | 538 | (acetate)n+(Fe)n+ observed w/ high conc. of acetic acid
ESI+ | 550 |  | Derived from rubber tip from disposable syringe plunger
ESI+ | 553 | (M+Na)+ seen w/ Nunc polypropylene tubes
ESI+ | 555 | (acetate)n+(Fe)n+ observed w/ high conc. of acetic acid
ESI+ | 569 | (M+K)+ seen w/ Nunc polypropylene tubes
ESI+ | 587 | (M+H)+ seen w/ Vanguard polypropylene tubes
ESI+ | 593 | (M+H)+ hexadecamethylcyclooctasiloxane
ESI+ | 609 | (M+Na)+ seen w/ Vanguard polypropylene tubes
ESI+ | 610 | seen w/ silica capillary NanoSprayer
ESI+ | 625 | (M+K)+ seen w/ Vanguard polypropylene tubes
ESI+ | 638 | seen w/ silica capillary NanoSprayer
ESI+ | 667 | (M+H)+ octadecamethylcyclononasiloxane
ESI+ | 675 | seen w/ new PEEK tubing
ESI+ | 679 | (M+H)+ from purified water and solvent filters
ESI+ | 684 | seen w/ silica capillary NanoSprayer
ESI+ | 696 | (M+NH4)+ from purified water and solvent filters
ESI+ | 701 | (M+Na)+ from purified water and solvent filters
ESI+ | 741 | (M+H)+ eicosamethylcyclodecasiloxane
ESI+ | 786 | (M+Ti)+ titanium adduct
ESI+ | 798 | (2M+NH4)+ diisooctyl phthalate (plasticiser)
ESI+ | 803 | (2M+Na)+ diisooctyl phthalate (plasticiser)
ESI- | 26 | CN- acetonitrile
ESI- | 45 | (HCOO)- formate
ESI- | 59 | (CH3CO2)- acetate
ESI- | 69 | (CF3CO2Na)- sodium trifluoracetate, TFA
ESI- | 79 | PO3- phosphate
ESI- | 80 | SO3- sulfate
ESI- | 95 | CH3SO3- methanesulfonate
ESI- | 97 | HSO4-
ESI- | 97 | H2PO4-
ESI- | 113 | sodium trifluoracetate, TFA
ESI- | 127 | (CF3CO2Na)- sodium trifluoracetate, TFA
ESI- | 227 | (CF3CO2Na)- sodium trifluoracetate, TFA
ESI- | 249 | (CF3CO2Na)- sodium trifluoracetate, TFA
ESI+ Clusters
m/z Compound
ESI+ Clusters | 18 | water
ESI+ Clusters | 32 | methanol
ESI+ Clusters | 41 | acetonitrile
ESI+ Clusters | 44 | PEG, HO-(CH2CH2On-H)+
ESI+ Clusters | 50 | perfluoro, CF2
ESI+ Clusters | 53 | NH4Cl
ESI+ Clusters | 58 | PPG
ESI+ Clusters | 58 | NaCl
ESI+ Clusters | 63 | NH4Cl salts
ESI+ Clusters | 68 | NaFA salts
ESI+ Clusters | 72 | Replacement of OH by OSi(CH3)3
ESI+ Clusters | 74 | KCl
ESI+ Clusters | 74 | Si(CH3)2O (from rubber)
ESI+ Clusters | 77 | NH4Ac salts
ESI+ Clusters | 78 | DMSO
ESI+ Clusters | 82 | NaAc
ESI+ Clusters | 136 | NaTFA
ESI+ Clusters | 162 | Polysaccharides
ESI+ Common Adducts | (M+1) | (M+H)+
ESI+ Common Adducts | (M+18) | (M+NH4)+
ESI+ Common Adducts | (M+23) | (M+Na)+
ESI+ Common Adducts | (M+39) | (M+K)+
ESI+ Common Adducts | (M+33) | (M+CH3OH+H)+
ESI+ Common Adducts | (M+42) | (M+ACN+H)+
ESI+ Common Adducts | (M+64) | (M+ACN+Na)+
ESI+ Common Adducts | (M+79) | (M+DMSO+H)+
ESI+ Common Adducts | (M+83) | (M+2ACN+H)+
ESI+ Common Ion Series | 79 |  | (nDMSO+H)+
ESI+ Common Ion Series | 157 |  | (nDMSO+H)+
ESI+ Common Ion Series | 235 |  | (nDMSO+H)+
ESI+ Common Ion Series | 101 |  | (nDMSO+Na)+
ESI+ Common Ion Series | 179 |  | (nDMSO+Na)+
ESI+ Common Ion Series | 257 |  | (nDMSO+Na)+
ESI+ Common Ion Series | 231 |  | plastisizer
ESI+ Common Ion Series | 522 |  | plastisizer
ESI+ Common Ion Series | 550 |  | plastisizer
ESI+ Common Ion Series | 267 |  | tributylphosphate
ESI+ Common Ion Series | 289 |  | tributylphosphate
ESI+ Common Ion Series | 330 |  | tributylphosphate
ESI+ Common Ion Series | 371 |  | cyclosiloxane
ESI+ Common Ion Series | 388 |  | cyclosiloxane
ESI+ Common Ion Series | 445 |  | cyclosiloxane
ESI+ Common Ion Series | 462 |  | cyclosiloxane
ESI+ Common Ion Series | 355 |  | plastisizer
ESI+ Common Ion Series | 429 |  | plastisizer
ESI+ Common Ion Series | 503 |  | plastisizer
ESI+ Common Ion Series | 593 |  | plastisizer
ESI+ Common Ion Series | 667 |  | plastisizer
ESI+ Common Ion Series | 741 |  | plastisizer
ESI+ Common Ion Series | 815 |  | plastisizer

## Glossary
Type of ions: | Explanation
f+ | fragment ion
[f+H]+ | protonated fragment ion, e.g. in-source fragmentation of peptide ions
[M+H]+ | protonated molecular ion (pseudomolecular ion)
[M+Na]+ | sodiated molecular ion
[M+K]+ | potassiated molecular ion
[M2+H]+, [M3+H]+ etc... | protonated dimeric, trimeric, etc... molecular ion
[AnBm+H]+ | protonated moleuclar ion consisting of n A and m B subunits
Abbreviations | Explanation
4-HCCA | α-cyano-4-hydroxycinnamic acid, common matrix substance for MALDI MS analysis
2,5-DHB | 2,5-Dihydroxy benzoic acid, common matrix substance for MALDI MS analysis
MeCN, ACN | acetonitrile, solvent
MeOH | methanol, solvent
MeNO2 | nitromethane, solvent
HABA | 2-(4-hydroxyphenyl-azo)-benzoic acid, matrix substance for MALDI MS analysis
SA | sinapic or sinapinic acid, common matrix substance for MALDI MS analysis
PEG | Polyethylene glycol, Repeat unit: -[O-CH2-CH2-]-; 44 Da
PPG | Polypropylene glycol, Repeat unit: -[O-C(CH3)H-CH2-]-; 58 Da
XaaCcamXaa | carbamidomethylated cysteine residue (+ICH2CONH2 - HI), +57Da
XaaMoxXaa | singly oxidized methionine residue (+O, 16 Da)

## Refs
Ref Label | Author(s) | Citation or Website/Download Information
A | Waters Corporation | Background Ion List online available at: www2.waters.com/CEConversion.nsf/files/3929E3EC20E43AAA8525710D004AB62E/$file/bkgrnd_ion_mstr_list.pdf | Waters | http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf | common background ions [pdf]
B | Applied Biosystems | Appendix D: Commonly Observed Background Ions: Mariner Biospectrometry's Workstation Users Guide
C | New Objective | Technical Note: Common Background Ions for Electrospray, online available at: http://www.newobjective.com/downloads/technotes/PV-3.pdf | New Objective | http://www.newobjective.com/downloads/technotes/PV-3.pdf | common background ions [pdf]
D | Sigma-Aldrich | Chemical Formulas for Tween, Triton, and reduced Triton were adopted from the Sigma-Aldrich catalogue available online at: http://www.sigmaaldrich.com
E | Thermo Corporation, Mahn, B. | List of LC/MS contaminants online available at: www.abrf.org/index.cfm/list.msg/66994
F | Tong, H.;Bell, D.; Tabei, K.; Siegel, M.M. | J. Am. Soc. Mass Spectrom., 10 (1999) 1174.
G | Andersen, J.S.; Kuester, B.; Podtelejnikov, A.; Mortz, E.; Mann, M. | Proc. 47th ASMS Conf. Mass Spectrom. Allied Topics, 1999, Dallas,TX, USA.,
H | Keller, B.O.; Li, L. | J. Am. Soc. Mass Spectrom., 11 (2000) 88.
I | Keller, B.O.; Li, L.; Keller, H. | Program to predict and confirm matrix clusters masses, online available at: www.chem.ualberta.ca/~liweb/links/MaClust.htm
J | Harris,W.A.; Janecki, D.J.; Reilly, J.P. | Rapid Commun. Mass Spectrom., 16 (2002) 1714.
K | Keller, B.O.; Sui., J; Young, A.B.; Whittal, R.M.. | unpublished results; ESI background ions, Tween, Triton, PEGs, PPGs at www.chem.ualberta.ca/~massspec/es_ions.pdf (R.M. Whittal)
L | Schlosser, A.; Volkmer-Engert, R. | J. Mass Spectrom., 38 (2003) 523.
M | Tran, J.C.; Doucette, A.A. | J. Am. Soc. Mass Spectrom., 17 (2006) 652.
N | Verge, K.M.; Agnes, G.R. | J. Am. Soc. Mass Spectrom., 13 (2002) 901.
O | Paez, A.; Howe, A. | Canadian Chemical News, 56 (2004) 14.
P | Purves, R.W.; Gabryelski, W.; Li, L. | Rev. Sci. Instrument, 68 (1997) 3252.
Q | Gibson, C.R.; Brown, C.M. | J. Am. Soc. Mass spectrom., 14 (2003) 14.
R | Beavis,R. C.; Chait,B. T. | Anal. Chem., 62 (1990) 1836.
S | Guzzetta, A. | www.ionsource.com; Carbohydrate marker ions | Ionsource | http://ionsource.com/ | pretty cool site with lots of information
T | Clauser, K.R.; Hall, S.C.; Smith, D.M.; Webb, J.W.; Andrews, L.E.; Tran, H.M.; Epstein, L.B; Burlingame, A.L. | Proc. Natl. Acad. Sci. U.S.A., 92 (1995) 5072., and http://prospector.ucsf.edu
U | Macha,S.F.; Limbach,P.A.; Hanton,S.D.; Owens,K.G. | J. Am. Soc. Mass Spectrom., 12 (2001), 732.
V | Pleasance,S.; Thibault,P.; Sim,P.G.; Boyd,R.K. | Rapid Commun. Mass Spectrom., 5 (1991) 307.
W | Xia,Y.-Q.; Patel,S.; Bakhtiar,R.; Franklin,R.B.; Doss,G.A. | J. Am. Soc. Mass Spectrom., 16 (2005) 417.
X | Guo,X.; Bruins,A.P.; Covey,T.R. | Rapid Commun. Mass Spectrom., 20 (2006) 3145.
Y | Ijames, C.F.; Dutky, R.C.; Fales, H.M. | J. Am. Soc. Mass Spectrom., 6 (1995) 1226.
Z | Hesse, M.; Meier, H.; Zeeh, B. | Spektroskopische Methoden in der organischen Chemie, Georg Thieme Verlag, Stuttgart, 3. Aufl. 1987, ISBN: 3-13-576103-7
| Ionsource | http://ionsource.com/ | pretty cool site with lots of information
| UNIMOD | http://www.unimod.org | interactive list of protein modifications for mass spectrometry
| New Objective | http://www.newobjective.com/downloads/technotes/PV-3.pdf | common background ions [pdf]
| Waters | http://www.waters.com/webassets/cms/support/docs/bkgrnd_ion_mstr_list.pdf | common background ions [pdf]
| Eppenodorf | http://www.eppendorf.com/uploads/media/Application_282_Safe-Lock-Tubes_DNA-LoBind-Tubes_Avoiding-interferenc_01.pdf | Avoiding interferences and contaminants using Eppendorf Safe-Lock Tubes in mass spectrometry studies of natural products
| Agilent | Common LC/MS Contaminants - CIGS
| Eppenodorf | http://www.eppendorf.com/uploads/media/Application_282_Safe-Lock-Tubes_DNA-LoBind-Tubes_Avoiding-interferenc_01.pdf | Avoiding interferences and contaminants using Eppendorf Safe-Lock Tubes in mass spectrometry studies of natural products